## All
TYPE | NOM | Method | Units | Result | Z-Score | Average | Std Dev | Count | Min | Max | Rdat | Rpub | VR | AD | User | Date | runsum (pas tout de suite svp) | Comment
D- | D-1191 | Cetane Number - ASTM D613 |  |  |  | 39.59 | 0.59 | 8 | 38.6 | 40.5 | 1.625 | 2.748 |  | 0.29
D- | D-1191 | Test 1 Derived Cetane Number - ASTM D6890 |  |  |  | 40.75 | 0.56 | 4 | 40.2 | 41.51 | 1.563 | 2.262 |  | 0.302
D- | D-1191 | Gravity, API at 60 °F - All Methods |  |  |  | 32.01 | 0.02 | 10 | 32 | 32.06 | 0.052 |  |  | 2.465
D- | D-1191 | Sulfur - All Methods | mg/kg |  |  | 7.6 | 1.3 | 11 | 4.91 | 9.4 | 3.683 |  |  | 0.821
D- | D-1191 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.8853 | 0.0113 | 11 | 2.866 | 2.912 | 0.031 | 0.065 |  | 0.635
D- | D-1191 | Aniline Point - ASTM D611 | °C |  |  | 54.135 | 0.206 | 4 | 53.89 | 54.35 | 0.569 | 0.5 |  | 0.263
D- | D-1191 | Total Aromatics - All Methods | Vol. % |  |  | 36.63 | 2.02 | 4 | 34.7 | 39 | 5.602 |  |  | 0.31
D- | D-1191 | Flash Point - ASTM D93 | °F |  |  | 145.8 | 2.3 | 10 | 143 | 150.8 | 6.502 |  |  | 0.406
D- | D-1191 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 336.59 | 7.93 | 12 | 319.1 | 349.5 | 21.973 |  |  | 0.387
D- | D-1191 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 416.1 | 2.64 | 12 | 411.8 | 420.7 | 7.323 |  |  | 0.166
D- | D-1191 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 512.6 | 1.53 | 12 | 509 | 514.7 | 4.227 |  |  | 0.435
D- | D-1191 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 623.94 | 2 | 12 | 621.7 | 628.9 | 5.552 |  |  | 0.571
D- | D-1191 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 674.95 | 2.27 | 11 | 672 | 679 | 6.277 |  |  | 0.23
D- | D-1191 | Recovery - ASTM D86 | Vol. % |  |  | 97.92 | 0.49 | 12 | 96.8 | 98.7 | 1.361 |  |  | 1.063
D- | D-1191 | Residue - ASTM D86 | Vol. % |  |  | 1.41 | 0.23 | 12 | 1.1 | 1.8 | 0.641 |  |  | 0.62
D- | D-1191 | Loss - ASTM D86 | Vol. % |  |  | 0.68 | 0.45 | 12 | 0 | 1.4 | 1.256 |  |  | 0.291
D- | D-1192 | Cetane Number - ASTM D613 |  |  |  | 43.53 | 0.91 | 10 | 42.1 | 45 | 2.529 | 3.241 |  | 0.288 | Eric Salvail | 2023-05-23
D- | D-1192 | Test 1 Derived Cetane Number - ASTM D6890 |  | 43 | 1.4062 | 43.9 | 0.64 | 4 | 43.21 | 44.5 | 1.785 | 2.383 |  | 0.468 | Eric Salvail | 2023-05-23
D- | D-1192 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 42.86 | 0.38 | 3 | 42.47 | 43.22 | 1.041 | 1.012 |  | 0.292 | Eric Salvail | 2023-05-23
D- | D-1192 | Gravity, API at 60 °F - All Methods |  |  |  | 33.77 | 0.04 | 10 | 33.7 | 33.8 | 0.107 |  |  | 1.213 | Eric Salvail | 2023-05-23
D- | D-1192 | Sulfur - All Methods | mg/kg |  |  | 12.1 | 1.3 | 10 | 10 | 14 | 3.73 |  |  | 0.343 | Eric Salvail | 2023-05-23
D- | D-1192 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.3191 | 0.0046 | 8 | 2.31 | 2.323 | 0.013 | 0.052 |  | 0.996 | Eric Salvail | 2023-05-23
D- | D-1192 | Aniline Point - ASTM D611 | °C |  |  | 37.05 | 26.024 | 3 | 7 | 52.1 | 72.087 | 0.5 |  | 0.732 | Eric Salvail | 2023-05-23
D- | D-1192 | Total Aromatics - All Methods | Vol. % |  |  | 36.53 | 3.3 | 3 | 33.8 | 40.2 | 9.142 |  |  | 0.384 | Eric Salvail | 2023-05-23
D- | D-1192 | Aromatics - ASTM D1319 | Vol. % |  |  | 35.7 | 1.71 | 4 | 33.8 | 37.2 | 4.75 |  |  | 0.501 | Eric Salvail | 2023-05-23
D- | D-1192 | Olefins - ASTM D1319 | Vol. % |  |  | 1.63 | 0.5 | 4 | 0.9 | 2 | 1.383 |  |  | 0.562 | Eric Salvail | 2023-05-23
D- | D-1192 | Saturates - ASTM D1319 | Vol. % |  |  | 62.68 | 1.88 | 4 | 60.9 | 64.4 | 5.214 |  |  | 0.616 | Eric Salvail | 2023-05-23
D- | D-1192 | Flash Point - ASTM D93 | °F |  |  | 161.7 | 2.1 | 8 | 160 | 166 | 5.779 |  |  | 0.667 | Eric Salvail | 2023-05-23
D- | D-1192 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 354.94 | 5.37 | 10 | 346.8 | 363.77 | 14.881 |  |  | 0.158 | Eric Salvail | 2023-05-23
D- | D-1192 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 415.37 | 2.88 | 10 | 411.7 | 420.5 | 7.978 |  |  | 0.324 | Eric Salvail | 2023-05-23
D- | D-1192 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 493.42 | 0.98 | 10 | 491.9 | 495 | 2.723 |  |  | 0.141 | Eric Salvail | 2023-05-23
D- | D-1192 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 563.1 | 0.95 | 10 | 562.1 | 564.5 | 2.631 |  |  | 0.529 | Eric Salvail | 2023-05-23
D- | D-1192 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 592.6 | 1.95 | 9 | 590.4 | 596.6 | 5.41 |  |  | 0.611 | Eric Salvail | 2023-05-23
D- | D-1192 | Recovery - ASTM D86 | Vol. % |  |  | 98.2 | 0.28 | 10 | 97.7 | 98.7 | 0.783 |  |  | 0.313 | Eric Salvail | 2023-05-23
D- | D-1192 | Residue - ASTM D86 | Vol. % |  |  | 1.3 | 0.07 | 9 | 1.2 | 1.4 | 0.196 |  |  | 0.858 | Eric Salvail | 2023-05-23
D- | D-1192 | Loss - ASTM D86 | Vol. % |  |  | 0.45 | 0.34 | 10 | 0 | 1 | 0.953 |  |  | 0.281 | Eric Salvail | 2023-05-23
D- | D-1193 | Cetane Number - ASTM D613 |  |  |  | 50.55 | 1.56 | 10 | 47.8 | 53.4 | 4.31 | 4.119 |  | 0.223 | Eric Salvail | 2023-07-14
D- | D-1193 | Test 1 Derived Cetane Number - ASTM D6890 |  | 49.29 | 1.9737 | 50.79 | 0.76 | 5 | 49.97 | 51.84 | 2.102 | 2.649 |  | 0.318 | Eric Salvail | 2023-07-14
D- | D-1193 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 50.32 | 0.62 | 3 | 49.84 | 51.02 | 1.717 | 1.358 |  | 0.444 | Eric Salvail | 2023-07-14
D- | D-1193 | Gravity, API at 60 °F - All Methods |  |  |  | 35.7 | 0.01 | 10 | 35.69 | 35.72 | 0.026 |  |  | 1.836 | Eric Salvail | 2023-07-14
D- | D-1193 | Sulfur - All Methods | mg/kg | 8.8 | 0.8 | 8 | 1 | 12 | 6 | 9.3 | 2.79 |  |  | 0.602 | Eric Salvail | 2023-07-14
D- | D-1193 | Viscosity at 40°C - ASTM D445 | cSt | 3.187 | 0.86 | 3.1785 | 0.0098 | 11 | 3.157 | 3.194 | 0.027 | 0.071 |  | 0.476 | Eric Salvail | 2023-07-14
D- | D-1193 | Aniline Point - ASTM D611 | °C | 68.65 | -0.46 | 87.62 | 41.134 | 5 | 68.65 | 161.2 | 113.942 | 0.5 |  | 1.466 | Eric Salvail | 2023-07-14
D- | D-1193 | Total Aromatics - All Methods | Vol. % |  |  | 23.4 | 3.5 | 3 | 19.4 | 25.9 | 9.695 |  |  | 0.525 | Eric Salvail | 2023-07-14
D- | D-1193 | Aromatics - ASTM D1319 | Vol. % |  |  | 25.73 | 1.46 | 3 | 24.2 | 27.1 | 4.036 |  |  | 0.299 | Eric Salvail | 2023-07-14
D- | D-1193 | Olefins - ASTM D1319 | Vol. % |  |  | 3.13 | 1.8 | 3 | 1.9 | 5.2 | 4.989 |  |  | 0.564 | Eric Salvail | 2023-07-14
D- | D-1193 | Saturates - ASTM D1319 | Vol. % |  |  | 71.13 | 0.61 | 3 | 70.6 | 71.8 | 1.692 |  |  | 0.345 | Eric Salvail | 2023-07-14
D- | D-1193 | Flash Point - ASTM D93 | °F | 152.9 | 0.28 | 152.4 | 1.6 | 11 | 149.9 | 156 | 4.562 |  |  | 0.355 | Eric Salvail | 2023-07-14
D- | D-1193 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 354.4 | 0.52 | 350.65 | 7.16 | 12 | 334.1 | 358.7 | 19.836 |  |  | 0.614 | Eric Salvail | 2023-07-14
D- | D-1193 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 437.7 | 0.79 | 435.07 | 3.33 | 12 | 429 | 440.2 | 9.221 |  |  | 0.222 | Eric Salvail | 2023-07-14
D- | D-1193 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 533.1 | 0.48 | 532.36 | 1.54 | 12 | 529.7 | 535.1 | 4.273 |  |  | 0.195 | Eric Salvail | 2023-07-14
D- | D-1193 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 622.2 | -0.64 | 623.72 | 2.37 | 12 | 620.52 | 629 | 6.571 |  |  | 0.255 | Eric Salvail | 2023-07-14
D- | D-1193 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 680 | 1.26 | 677.16 | 2.26 | 10 | 674.2 | 681.1 | 6.265 |  |  | 0.315 | Eric Salvail | 2023-07-14
D- | D-1193 | Recovery - ASTM D86 | Vol. % | 98.8 | 1.82 | 97.86 | 0.52 | 12 | 97.1 | 98.8 | 1.434 |  |  | 0.238 | Eric Salvail | 2023-07-14
D- | D-1193 | Residue - ASTM D86 | Vol. % | 1 | -0.98 | 1.29 | 0.3 | 12 | 0.6 | 1.7 | 0.822 |  |  | 0.825 | Eric Salvail | 2023-07-14
D- | D-1193 | Loss - ASTM D86 | Vol. % | 0.2 | -1.3 | 0.85 | 0.5 | 12 | 0.2 | 1.5 | 1.38 |  |  | 0.557 | Eric Salvail | 2023-07-14
D- | D-1194 | Cetane Number - ASTM D613 |  |  |  | 46 | 0.68 | 9 | 44.7 | 47.2 | 1.874 | 3.55 |  | 0.538
D- | D-1194 | Test 1 Derived Cetane Number - ASTM D6890 |  |  |  | 45.66 | 0.73 | 3 | 45.23 | 46.5 | 2.023 | 2.451 |  | 0.721
D- | D-1194 | Gravity, API at 60 °F - All Methods |  |  |  | 37.21 | 0.02 | 10 | 37.2 | 37.24 | 0.043 |  |  | 0.963
D- | D-1194 | Sulfur - All Methods | mg/kg |  |  | 7.8 | 1.1 | 10 | 5.5 | 9 | 3.013 |  |  | 0.385
D- | D-1194 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.3265 | 0.0073 | 8 | 2.315 | 2.34 | 0.02 | 0.052 |  | 0.266
D- | D-1194 | Aniline Point - ASTM D611 | °C |  |  | 74.003 | 18.669 | 3 | 63.05 | 95.56 | 51.714 | 0.5 |  | 0.715
D- | D-1194 | Total Aromatics - All Methods | Vol. % |  |  | 24.5 | 1.83 | 3 | 22.9 | 26.5 | 5.077 |  |  | 0.345
D- | D-1194 | Aromatics - ASTM D1319 | Vol. % |  |  | 24.32 | 1.14 | 5 | 22.9 | 25.8 | 3.155 |  |  | 0.249
D- | D-1194 | Olefins - ASTM D1319 | Vol. % |  |  | 0.92 | 0.23 | 5 | 0.7 | 1.2 | 0.632 |  |  | 0.396
D- | D-1194 | Saturates - ASTM D1319 | Vol. % |  |  | 73.02 | 4.43 | 6 | 64.2 | 76.2 | 12.268 |  |  | 1.032
D- | D-1194 | Flash Point - ASTM D93 | °F |  |  | 138.5 | 4 | 9 | 130 | 143 | 11.184 |  |  | 0.448
D- | D-1194 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 325.48 | 2.64 | 10 | 319.8 | 328.89 | 7.315 |  |  | 0.335
D- | D-1194 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 382.72 | 2.88 | 10 | 378.2 | 387.5 | 7.975 |  |  | 0.295
D- | D-1194 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 486.79 | 2.46 | 10 | 482 | 490 | 6.811 |  |  | 0.264
D- | D-1194 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 610.59 | 3.02 | 10 | 607 | 615 | 8.369 |  |  | 0.693
D- | D-1194 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 667.22 | 3.32 | 10 | 661.3 | 672.3 | 9.193 |  |  | 0.225
D- | D-1194 | Recovery - ASTM D86 | Vol. % |  |  | 98.24 | 0.42 | 10 | 97.6 | 98.8 | 1.155 |  |  | 0.353
D- | D-1194 | Residue - ASTM D86 | Vol. % |  |  | 1.26 | 0.25 | 10 | 0.8 | 1.6 | 0.706 |  |  | 0.412
D- | D-1194 | Loss - ASTM D86 | Vol. % |  |  | 0.5 | 0.32 | 10 | 0 | 1 | 0.895 |  |  | 0.226
D- | D-1195 | Cetane Number - ASTM D613 |  |  |  | 46.13 | 0.94 | 11 | 44.8 | 47.7 | 2.599 | 3.566 |  | 0.295
D- | D-1195 | Test 1 Derived Cetane Number - ASTM D6890 |  |  |  | 44.79 | 0.61 | 5 | 43.88 | 45.5 | 1.697 | 2.417 |  | 0.314
D- | D-1195 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 45.98 | 0.59 | 4 | 45.6 | 46.84 | 1.628 | 1.157 |  | 0.671
D- | D-1195 | Gravity, API at 60 °F - All Methods |  |  |  | 35.4 | 0.01 | 9 | 35.38 | 35.42 | 0.033 |  |  | 1.261
D- | D-1195 | Sulfur - All Methods | mg/kg |  |  | 6.7 | 0.5 | 9 | 5.55 | 7.2 | 1.522 |  |  | 0.682
D- | D-1195 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.5438 | 0.1266 | 9 | 2.314 | 2.69 | 0.351 | 0.057 |  | 1.396
D- | D-1195 | Aniline Point - ASTM D611 | °C |  |  | 62.1 | 0.05 | 3 | 62.05 | 62.15 | 0.139 | 0.5 |  | 0.284
D- | D-1195 | Total Aromatics - All Methods | Vol. % |  |  | 27.73 | 2.97 | 3 | 24.6 | 30.5 | 8.219 |  |  | 0.303
D- | D-1195 | Aromatics - ASTM D1319 | Vol. % |  |  | 29.68 | 3.71 | 5 | 24.6 | 34.7 | 10.265 |  |  | 0.243
D- | D-1195 | Olefins - ASTM D1319 | Vol. % |  |  | 1.62 | 0.67 | 5 | 0.9 | 2.3 | 1.852 |  |  | 0.677
D- | D-1195 | Saturates - ASTM D1319 | Vol. % |  |  | 68.72 | 3.33 | 5 | 64.4 | 73.1 | 9.227 |  |  | 0.169
D- | D-1195 | Flash Point - ASTM D93 | °F |  |  | 113.8 | 3.7 | 8 | 108 | 121 | 10.369 |  |  | 0.424
D- | D-1195 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 318.29 | 5.68 | 10 | 306.2 | 325.7 | 15.726 |  |  | 0.386
D- | D-1195 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 405.53 | 2.14 | 10 | 402 | 408.3 | 5.924 |  |  | 0.307
D- | D-1195 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 505.72 | 1.57 | 10 | 503 | 507.8 | 4.336 |  |  | 0.487
D- | D-1195 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 620.66 | 1.89 | 10 | 617.6 | 623.4 | 5.23 |  |  | 0.18
D- | D-1195 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 672.14 | 3.2 | 10 | 666 | 677.1 | 8.858 |  |  | 0.167
D- | D-1195 | Recovery - ASTM D86 | Vol. % |  |  | 98.06 | 0.26 | 11 | 97.5 | 98.4 | 0.715 |  |  | 0.352
D- | D-1195 | Residue - ASTM D86 | Vol. % |  |  | 1.29 | 0.39 | 11 | 0.3 | 1.8 | 1.076 |  |  | 0.839
D- | D-1195 | Loss - ASTM D86 | Vol. % |  |  | 0.64 | 0.36 | 11 | 0.3 | 1.4 | 0.985 |  |  | 0.68
D- | D-1196 | Cetane Number - ASTM D613 |  |  |  | 50.35 | 0.78 | 13 | 48.3 | 51.2 | 2.162 | 4.093 | 0.279 - Pass | 0.892 | Eric Salvail | 2023-11-10
D- | D-1196 | Test 1 Derived Cetane Number - ASTM D6890 |  | 51.12 | 0.64045 | 51.69 | 0.89 | 6 | 50.36 | 53.1 | 2.467 | 2.683 |  | 0.348 | Eric Salvail | 2023-11-10
D- | D-1196 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 50.27 | 0.28 | 4 | 50 | 50.65 | 0.775 | 1.355 |  | 0.31 | Eric Salvail | 2023-11-10
D- | D-1196 | Gravity, API at 60 °F - All Methods |  |  |  | 38.75 | 0.05 | 12 | 38.66 | 38.8 | 0.135 |  |  | 0.564 | Eric Salvail | 2023-11-10
D- | D-1196 | Sulfur - All Methods | mg/kg | 2.7 | -0.22 | 2.8 | 0.7 | 12 | 1.88 | 4 | 1.858 |  |  | 0.282 | Eric Salvail | 2023-11-10
D- | D-1196 | Viscosity at 40°C - ASTM D445 | cSt | 2.213 | 1.17 | 2.2004 | 0.0107 | 10 | 2.183 | 2.216 | 0.03 | 0.049 |  | 0.344 | Eric Salvail | 2023-11-10
D- | D-1196 | Aniline Point - ASTM D611 | °C | 66.25 | -0.07 | 66.31 | 0.893 | 7 | 65.22 | 67.78 | 2.475 | 0.5 |  | 0.357 | Eric Salvail | 2023-11-10
D- | D-1196 | Total Aromatics - All Methods | Vol. % |  |  | 25.7 | 2.55 | 6 | 21.1 | 28.7 | 7.076 |  |  | 0.527 | Eric Salvail | 2023-11-10
D- | D-1196 | Aromatics - ASTM D1319 | Vol. % |  |  | 27.28 | 7.43 | 5 | 21.1 | 39.6 | 20.593 |  |  | 0.549 | Eric Salvail | 2023-11-10
D- | D-1196 | Olefins - ASTM D1319 | Vol. % |  |  | 2.22 | 1.02 | 5 | 1.1 | 3.9 | 2.834 |  |  | 0.644 | Eric Salvail | 2023-11-10
D- | D-1196 | Saturates - ASTM D1319 | Vol. % |  |  | 70.5 | 7.76 | 5 | 58.4 | 77.8 | 21.497 |  |  | 0.443 | Eric Salvail | 2023-11-10
D- | D-1196 | Flash Point - ASTM D93 | °F | 148.6 | -0.94 | 152 | 3.6 | 12 | 147 | 161 | 9.931 |  |  | 0.557 | Eric Salvail | 2023-11-10
D- | D-1196 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 347.6 | 0.66 | 344.81 | 4.22 | 12 | 335 | 349.6 | 11.698 |  |  | 0.428 | Eric Salvail | 2023-11-10
D- | D-1196 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 392.9 | -0.34 | 393.54 | 1.87 | 13 | 391 | 397.2 | 5.18 |  |  | 0.19 | Eric Salvail | 2023-11-10
D- | D-1196 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 479.1 | 0.73 | 477.79 | 1.79 | 13 | 475 | 480.4 | 4.948 |  |  | 0.278 | Eric Salvail | 2023-11-10
D- | D-1196 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 608.9 | 1.21 | 605.82 | 2.54 | 13 | 600 | 608.9 | 7.031 |  |  | 0.362 | Eric Salvail | 2023-11-10
D- | D-1196 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 674.5 | 2.02 | 667.1 | 3.67 | 13 | 661.7 | 674.5 | 10.178 |  |  | 0.18 | Eric Salvail | 2023-11-10
D- | D-1196 | Recovery - ASTM D86 | Vol. % | 97.9 | -0.38 | 98.05 | 0.39 | 13 | 97.4 | 98.8 | 1.071 |  |  | 0.183 | Eric Salvail | 2023-11-10
D- | D-1196 | Residue - ASTM D86 | Vol. % | 0.5 | -2.16 | 1.29 | 0.37 | 13 | 0.5 | 1.9 | 1.014 |  |  | 1.7 | Eric Salvail | 2023-11-10
D- | D-1196 | Loss - ASTM D86 | Vol. % | 1.6 | 1.89 | 0.66 | 0.5 | 13 | 0 | 1.6 | 1.378 |  |  | 0.582 | Eric Salvail | 2023-11-10
D- | D-1197 | Cetane Number - ASTM D613 |  |  |  | 42.77 | 0.83 | 10 | 41.6 | 44.2 | 2.288 | 3.146 |  | 0.22
D- | D-1197 | Test 1 Derived Cetane Number - ASTM D6890 |  |  |  | 42.54 | 0.78 | 5 | 41.51 | 43.5 | 2.163 | 2.331 |  | 0.219
D- | D-1197 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 43.43 | 0.49 | 3 | 43.11 | 44 | 1.371 | 1.039 |  | 0.623
D- | D-1197 | Gravity, API at 60 °F - All Methods |  |  |  | 35.9 | 0.02 | 10 | 35.86 | 35.92 | 0.046 |  |  | 1.291
D- | D-1197 | Sulfur - All Methods | mg/kg |  |  | 5.8 | 1.7 | 11 | 1.7 | 7.6 | 4.583 |  |  | 0.931
D- | D-1197 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.2351 | 0.0121 | 9 | 2.214 | 2.253 | 0.034 | 0.05 |  | 0.162
D- | D-1197 | Aniline Point - ASTM D611 | °C |  |  | 72.337 | 32.539 | 6 | 58.15 | 138.7 | 90.133 | 0.5 |  | 1.749
D- | D-1197 | Total Aromatics - All Methods | Vol. % |  |  | 32.3 | 1.71 | 3 | 30.4 | 33.7 | 4.725 |  |  | 0.391
D- | D-1197 | Aromatics - ASTM D1319 | Vol. % |  |  | 29.38 | 7.5 | 4 | 18.2 | 33.7 | 20.77 |  |  | 0.864
D- | D-1197 | Olefins - ASTM D1319 | Vol. % |  |  | 5.15 | 7.97 | 4 | 0.7 | 17.1 | 22.087 |  |  | 0.975
D- | D-1197 | Saturates - ASTM D1319 | Vol. % |  |  | 65.48 | 0.9 | 4 | 64.7 | 66.7 | 2.502 |  |  | 0.38
D- | D-1197 | Flash Point - ASTM D93 | °F |  |  | 139.4 | 2.6 | 10 | 136 | 145 | 7.14 |  |  | 0.387
D- | D-1197 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 324.37 | 7.44 | 11 | 312.1 | 337 | 20.619 |  |  | 0.184
D- | D-1197 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 381.44 | 2.74 | 11 | 376.4 | 386 | 7.576 |  |  | 0.201
D- | D-1197 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 481.98 | 1.79 | 10 | 477.6 | 484 | 4.97 |  |  | 0.804
D- | D-1197 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 624.04 | 3.55 | 11 | 617.1 | 630 | 9.836 |  |  | 0.296
D- | D-1197 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 676.83 | 8.74 | 11 | 656.8 | 688 | 24.198 |  |  | 0.733
D- | D-1197 | Recovery - ASTM D86 | Vol. % |  |  | 98.03 | 0.28 | 10 | 97.5 | 98.4 | 0.784 |  |  | 0.311
D- | D-1197 | Residue - ASTM D86 | Vol. % |  |  | 1.3 | 0.17 | 10 | 1.1 | 1.6 | 0.471 |  |  | 0.62
D- | D-1197 | Loss - ASTM D86 | Vol. % |  |  | 0.73 | 0.39 | 11 | 0.2 | 1.3 | 1.081 |  |  | 0.58
D- | D-1198 | Cetane Number - ASTM D613 |  |  |  | 44.91 | 1.28 | 10 | 42.5 | 47 | 3.558 | 3.414 |  | 0.186 | Eric Salvail | 2023-09-21
D- | D-1198 | Test 1 Derived Cetane Number - ASTM D6890 |  | 44.93 | 0.98837 | 45.78 | 0.86 | 4 | 44.5 | 46.4 | 2.396 | 2.455 |  | 0.667 | Eric Salvail | 2023-09-21
D- | D-1198 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 47.17 | 3.41 | 3 | 44.81 | 51.08 | 9.453 | 1.212 |  | 0.56 | Eric Salvail | 2023-09-21
D- | D-1198 | Gravity, API at 60 °F - All Methods |  | 31.95 | -2.06 | 32.01 | 0.03 | 11 | 31.95 | 32.05 | 0.075 |  |  | 0.953 | Eric Salvail | 2023-09-21
D- | D-1198 | Sulfur - All Methods | mg/kg | 8.7 | 0.31 | 8.4 | 1.1 | 11 | 7.02 | 11 | 3.104 |  |  | 0.464 | Eric Salvail | 2023-09-21
D- | D-1198 | Viscosity at 40°C - ASTM D445 | cSt | 3.201 | 0.78 | 3.1935 | 0.0097 | 11 | 3.184 | 3.218 | 0.027 | 0.072 |  | 0.901 | Eric Salvail | 2023-09-21
D- | D-1198 | Aniline Point - ASTM D611 | °C | 53.55 | -0.47 | 69.77 | 34.344 | 5 | 53.55 | 131.2 | 95.133 | 0.5 |  | 1.458 | Eric Salvail | 2023-09-21
D- | D-1198 | Total Aromatics - All Methods | Vol. % |  |  | 31.65 | 4.67 | 4 | 28.3 | 38.5 | 12.948 |  |  | 0.624 | Eric Salvail | 2023-09-21
D- | D-1198 | Aromatics - ASTM D1319 | Vol. % |  |  | 40 | 2.77 | 3 | 38.3 | 43.2 | 7.681 |  |  | 0.673 | Eric Salvail | 2023-09-21
D- | D-1198 | Olefins - ASTM D1319 | Vol. % |  |  | 2.3 | 0.79 | 3 | 1.4 | 2.9 | 2.199 |  |  | 0.468 | Eric Salvail | 2023-09-21
D- | D-1198 | Saturates - ASTM D1319 | Vol. % |  |  | 57.7 | 3.33 | 3 | 53.9 | 60.1 | 9.22 |  |  | 0.513 | Eric Salvail | 2023-09-21
D- | D-1198 | Flash Point - ASTM D93 | °F | 156 | 0.31 | 155.5 | 1.6 | 11 | 153 | 158 | 4.448 |  |  | 0.428 | Eric Salvail | 2023-09-21
D- | D-1198 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 363.9 | 0.47 | 360.54 | 7.13 | 12 | 347.8 | 368.2 | 19.748 |  |  | 0.66 | Eric Salvail | 2023-09-21
D- | D-1198 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 439.7 | 0.21 | 439.22 | 2.34 | 12 | 435.2 | 443.5 | 6.484 |  |  | 0.188 | Eric Salvail | 2023-09-21
D- | D-1198 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 526.4 | 1.32 | 524.61 | 1.36 | 12 | 522 | 527.34 | 3.764 |  |  | 0.499 | Eric Salvail | 2023-09-21
D- | D-1198 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 626.6 | -0.37 | 627.46 | 2.33 | 12 | 623.3 | 632.17 | 6.442 |  |  | 0.199 | Eric Salvail | 2023-09-21
D- | D-1198 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 685.9 | 2.17 | 677.71 | 3.78 | 12 | 671.2 | 685.9 | 10.475 |  |  | 0.372 | Eric Salvail | 2023-09-21
D- | D-1198 | Recovery - ASTM D86 | Vol. % | 98.3 | 0.86 | 98.08 | 0.25 | 11 | 97.7 | 98.3 | 0.699 |  |  | 1.416 | Eric Salvail | 2023-09-21
D- | D-1198 | Residue - ASTM D86 | Vol. % | 1 | -0.69 | 1.27 | 0.39 | 12 | 0.5 | 1.96 | 1.091 |  |  | 0.791 | Eric Salvail | 2023-09-21
D- | D-1198 | Loss - ASTM D86 | Vol. % | 0.7 | 0.1 | 0.66 | 0.36 | 11 | 0.2 | 1.4 | 0.985 |  |  | 0.494 | Eric Salvail | 2023-09-21
D- | D-1199 | Cetane Number - ASTM D613 |  |  |  | 51.1 | 1.4 | 8 | 49.6 | 53.4 | 3.87 | 4.188 |  | 0.556 | Eric Salvail | 2023-10-16
D- | D-1199 | Test 1 Derived Cetane Number - ASTM D6890 |  | 49.46w |  | 50.07 | 0.75 | 5 | 48.96 | 50.9 | 2.089 | 2.621 |  | 0.364 | Eric Salvail | 2023-10-16
D- | D-1199 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 50.49 | 0.34 | 3 | 50.1 | 50.76 | 0.954 | 1.365 |  | 0.424 | Eric Salvail | 2023-10-16
D- | D-1199 | Gravity, API at 60 °F - All Methods |  |  |  | 37.56 | 0.05 | 9 | 37.5 | 37.6 | 0.131 |  |  | 1.323 | Eric Salvail | 2023-10-16
D- | D-1199 | Sulfur - All Methods | mg/kg | 8.1 | -0.64 | 8.4 | 0.4 | 9 | 7.8 | 9.09 | 1.151 |  |  | 0.187 | Eric Salvail | 2023-10-16
D- | D-1199 | Viscosity at 40°C - ASTM D445 | cSt | 2.727 | 0.56 | 2.7212 | 0.0104 | 10 | 2.703 | 2.742 | 0.029 | 0.061 |  | 0.329 | Eric Salvail | 2023-10-16
D- | D-1199 | Aniline Point - ASTM D611 | °C | 67.1 | -1.15 | 67.438 | 0.293 | 4 | 67.1 | 67.8 | 0.811 | 0.5 |  | 0.229 | Eric Salvail | 2023-10-16
D- | D-1199 | Total Aromatics - All Methods | Vol. % |  |  | 27.48 | 2.97 | 4 | 24.5 | 31 | 8.24 |  |  | 0.308 | Eric Salvail | 2023-10-16
D- | D-1199 | Aromatics - ASTM D1319 | Vol. % |  |  | 27.6 | 1.65 | 4 | 25.6 | 29.1 | 4.574 |  |  | 0.38 | Eric Salvail | 2023-10-16
D- | D-1199 | Olefins - ASTM D1319 | Vol. % |  |  | 2.25 | 0.56 | 4 | 1.5 | 2.8 | 1.542 |  |  | 0.294 | Eric Salvail | 2023-10-16
D- | D-1199 | Saturates - ASTM D1319 | Vol. % |  |  | 70.18 | 1.67 | 4 | 68.1 | 72 | 4.632 |  |  | 0.221 | Eric Salvail | 2023-10-16
D- | D-1199 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 353.9 | 0.42 | 352.42 | 3.51 | 12 | 344.5 | 358.32 | 9.717 |  |  | 0.361 | Eric Salvail | 2023-10-16
D- | D-1199 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 415.6 | 0.66 | 413.22 | 3.62 | 12 | 404.4 | 417.9 | 10.04 |  |  | 0.479 | Eric Salvail | 2023-10-16
D- | D-1199 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 524.1 | 0.96 | 522.1 | 2.09 | 12 | 518.2 | 524.2 | 5.783 |  |  | 0.616 | Eric Salvail | 2023-10-16
D- | D-1199 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 621.9 | 0.16 | 621.48 | 2.64 | 12 | 617.2 | 627.4 | 7.3 |  |  | 0.26 | Eric Salvail | 2023-10-16
D- | D-1199 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 660.9 | -0.22 | 661.71 | 3.72 | 11 | 656.1 | 668.5 | 10.31 |  |  | 0.437 | Eric Salvail | 2023-10-16
D- | D-1199 | Recovery - ASTM D86 | Vol. % | 98.1 | 0.25 | 97.93 | 0.71 | 12 | 96.8 | 99.4 | 1.977 |  |  | 0.19 | Eric Salvail | 2023-10-16
D- | D-1199 | Residue - ASTM D86 | Vol. % | 1.5 | 0.45 | 1.38 | 0.28 | 12 | 0.8 | 1.8 | 0.776 |  |  | 0.364 | Eric Salvail | 2023-10-16
D- | D-1199 | Loss - ASTM D86 | Vol. % | 0.4 | -0.69 | 0.82 | 0.61 | 12 | 0.1 | 2.4 | 1.678 |  |  | 0.846 | Eric Salvail | 2023-10-16
D- | D-1199 | Flash Point - ASTM D93 | °F | 113 | -22.32 | 158.7 | 2 | 9 | 156 | 162 | 5.673 |  |  | 0.664 | Eric Salvail | 2023-10-16
D- | D-1200 | Cetane Number - ASTM D613 |  |  |  | 52.51 | 0.98 | 16 | 50.2 | 54.2 | 2.702 | 4.363 | 0.383 - Pass | 0.501 | Eric Salvail | 2023-11-24
D- | D-1200 | Test 1 Derived Cetane Number - ASTM D6890 |  | 50.02w |  | 50.17 | 0.79 | 5 | 49.15 | 50.9 | 2.202 | 2.625 |  | 0.479 | Eric Salvail | 2023-11-24
D- | D-1200 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 51.91 | 0.32 | 4 | 51.44 | 52.15 | 0.877 | 1.431 |  | 0.561 | Eric Salvail | 2023-11-24
D- | D-1200 | Gravity, API at 60 °F - All Methods |  |  |  | 38.8 | 0.01 | 10 | 38.78 | 38.806 | 0.02 |  |  | 1.782 | Eric Salvail | 2023-11-24
D- | D-1200 | Sulfur - All Methods | mg/kg | 3.3 | 0.33 | 3.1 | 0.6 | 11 | 1.92 | 4.3 | 1.684 |  |  | 0.508 | Eric Salvail | 2023-11-24
D- | D-1200 | Viscosity at 40°C - ASTM D445 | cSt | 2.615 | 1.37 | 2.6081 | 0.005 | 11 | 2.597 | 2.615 | 0.014 | 0.058 |  | 0.375 | Eric Salvail | 2023-11-24
D- | D-1200 | Aniline Point - ASTM D611 | °C | 69.95 | 0.15 | 69.91 | 0.274 | 5 | 69.44 | 70.11 | 0.759 | 0.5 |  | 0.759 | Eric Salvail | 2023-11-24
D- | D-1200 | Total Aromatics - All Methods | Vol. % |  |  | 21.6 | 2.06 | 5 | 19.3 | 24.9 | 5.704 |  |  | 0.487 | Eric Salvail | 2023-11-24
D- | D-1200 | Aromatics - ASTM D1319 | Vol. % |  |  | 20.08 | 1.92 | 5 | 18 | 22.7 | 5.312 |  |  | 0.335 | Eric Salvail | 2023-11-24
D- | D-1200 | Olefins - ASTM D1319 | Vol. % |  |  | 2.54 | 0.81 | 5 | 1.7 | 3.8 | 2.238 |  |  | 0.28 | Eric Salvail | 2023-11-24
D- | D-1200 | Saturates - ASTM D1319 | Vol. % |  |  | 77.38 | 2.34 | 5 | 73.5 | 79.2 | 6.489 |  |  | 0.577 | Eric Salvail | 2023-11-24
D- | D-1200 | Flash Point - ASTM D93 | °F | 177.8 | 0.21 | 177.4 | 1.7 | 10 | 174.2 | 180 | 4.826 |  |  | 0.443 | Eric Salvail | 2023-11-24
D- | D-1200 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 389.4 | 0.66 | 385.92 | 5.25 | 13 | 378.2 | 394.7 | 14.54 |  |  | 0.207 | Eric Salvail | 2023-11-24
D- | D-1200 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 439.7 | 1.16 | 437.71 | 1.71 | 13 | 435 | 440.45 | 4.728 |  |  | 0.215 | Eric Salvail | 2023-11-24
D- | D-1200 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 498.7 | 0.94 | 497.44 | 1.34 | 13 | 494.2 | 499.1 | 3.716 |  |  | 0.419 | Eric Salvail | 2023-11-24
D- | D-1200 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 590.5 | 1.05 | 588.1 | 2.29 | 13 | 583.2 | 592.3 | 6.331 |  |  | 0.286 | Eric Salvail | 2023-11-24
D- | D-1200 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 646.7 | 0.48 | 644.67 | 4.25 | 13 | 638.2 | 650.2 | 11.773 |  |  | 0.865 | Eric Salvail | 2023-11-24
D- | D-1200 | Recovery - ASTM D86 | Vol. % | 98.2 | 0.06 | 98.18 | 0.36 | 13 | 97.6 | 98.8 | 0.987 |  |  | 0.194 | Eric Salvail | 2023-11-24
D- | D-1200 | Residue - ASTM D86 | Vol. % | 1.4 | 0.25 | 1.35 | 0.22 | 13 | 0.9 | 1.8 | 0.605 |  |  | 0.6 | Eric Salvail | 2023-11-24
D- | D-1200 | Loss - ASTM D86 | Vol. % | 0.4 | -0.27 | 0.48 | 0.28 | 13 | 0 | 1 | 0.777 |  |  | 0.325 | Eric Salvail | 2023-11-24
D- | D-1201 | Cetane Number - ASTM D613 |  |  |  | 19.84 | 0.93 | 10 | 18.6 | 21.6 | 2.576 | 0.28 |  | 0.341 | Martine Carrier | 2023-12-21
D- | D-1201 | Test 1 Derived Cetane Number - ASTM D6890 |  |  |  | 22.24 | 0.74 | 4 | 21.38 | 23 | 2.045 | 1.549 |  | 0.296 | Martine Carrier | 2023-12-21
D- | D-1201 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 19.25 | 0.26 | 4 | 18.89 | 19.5 | 0.708 | 0.481 |  | 0.48 | Martine Carrier | 2023-12-21
D- | D-1201 | Gravity, API at 60 °F - All Methods |  | 47.2 | 1.39 | 47.14 | 0.04 | 13 | 47.09 | 47.2 | 0.121 |  |  | 0.932 | Martine Carrier | 2023-12-21
D- | D-1201 | Sulfur - All Methods | mg/kg | 2.04 | 0.52 | 1.8 | 0.4 | 12 | 1 | 2.5 | 1.157 |  |  | 0.334 | Martine Carrier | 2023-12-21
D- | D-1201 | Viscosity at 40°C - ASTM D445 | cSt | 1.247 | 0.51 | 1.2401 | 0.0135 | 9 | 1.214 | 1.255 | 0.037 | 0.028 |  | 0.691 | Martine Carrier | 2023-12-21
D- | D-1201 | Aniline Point - ASTM D611 | °C | 63.75 | 1.14 | 62.943 | 0.708 | 6 | 62.22 | 63.89 | 1.96 | 0.5 |  | 0.51 | Martine Carrier | 2023-12-21
D- | D-1201 | Total Aromatics - All Methods | Vol. % |  |  | 22.29 | 4.25 | 6 | 14.6 | 25.5 | 11.778 |  |  | 0.734 | Martine Carrier | 2023-12-21
D- | D-1201 | Aromatics - ASTM D1319 | Vol. % |  |  | 22.06 | 1.54 | 5 | 20 | 24.2 | 4.254 |  |  | 0.262 | Martine Carrier | 2023-12-21
D- | D-1201 | Olefins - ASTM D1319 | Vol. % |  |  | 1.14 | 0.24 | 5 | 0.8 | 1.4 | 0.667 |  |  | 0.251 | Martine Carrier | 2023-12-21
D- | D-1201 | Saturates - ASTM D1319 | Vol. % |  |  | 76.8 | 1.67 | 5 | 74.6 | 79.2 | 4.623 |  |  | 0.338 | Martine Carrier | 2023-12-21
D- | D-1201 | Flash Point - ASTM D93 | °F | 116 | -2.3 | 120.9 | 2.1 | 10 | 116 | 123 | 5.905 |  |  | 0.611 | Martine Carrier | 2023-12-21 |  | recheck demandé
D- | D-1201 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 336.6 | 0.5 | 334.8 | 3.58 | 13 | 326.4 | 339.3 | 9.922 |  |  | 0.563 | Martine Carrier | 2023-12-21
D- | D-1201 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 347 | 0.38 | 346.67 | 0.88 | 13 | 345 | 348.21 | 2.425 |  |  | 0.12 | Martine Carrier | 2023-12-21
D- | D-1201 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 356.8 | 0.67 | 356.07 | 1.09 | 13 | 353.9 | 358.06 | 3.022 |  |  | 0.141 | Martine Carrier | 2023-12-21
D- | D-1201 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 419.7 | 0.05 | 419.26 | 8.28 | 13 | 409.6 | 435.2 | 22.949 |  |  | 0.63 | Martine Carrier | 2023-12-21
D- | D-1201 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 546.9 | 0.38 | 545.7 | 3.16 | 13 | 538.8 | 552 | 8.748 |  |  | 0.426 | Martine Carrier | 2023-12-21
D- | D-1201 | Recovery - ASTM D86 | Vol. % | 98.9 | 0.75 | 98.48 | 0.55 | 13 | 97.3 | 99.4 | 1.533 |  |  | 0.23 | Martine Carrier | 2023-12-21
D- | D-1201 | Residue - ASTM D86 | Vol. % | 0.4 | -2.49 | 1.09 | 0.28 | 12 | 0.4 | 1.4 | 0.77 |  |  | 1.401 | Martine Carrier | 2023-12-21
D- | D-1201 | Loss - ASTM D86 | Vol. % | 0.7 | 0.58 | 0.46 | 0.41 | 13 | 0 | 1.5 | 1.134 |  |  | 0.587 | Martine Carrier | 2023-12-21
D- | D-1202 | Cetane Number - ASTM D613 |  |  |  | 42.37 | 1.04 | 10 | 41 | 44.7 | 2.876 | 3.096 |  | 0.389 | Martine Carrier | 2023-12-29
D- | D-1202 | Test 1 Derived Cetane Number - ASTM D6890 |  | 41 | 0.22 | 40.87 | 0.61 | 7 | 40.06 | 41.6 | 1.701 | 2.266 |  | 0.331 | Martine Carrier | 2023-12-29
D- | D-1202 | Test 1 Derived Cetane Number - ASTM D7668 |  |  |  | 42.79 | 0.68 | 3 | 42.02 | 43.3 | 1.885 | 1.069 |  | 0.483 | Martine Carrier | 2023-12-29
D- | D-1202 | Sulfur - All Methods | mg/kg | 4.6 | -0.26 | 4.7 | 0.3 | 10 | 4.3 | 5.2 | 0.902 |  |  | 0.8 | Martine Carrier | 2023-12-29
D- | D-1202 | Viscosity at 40°C - ASTM D445 | cSt | 2.455 | 0.72 | 2.4312 | 0.0329 | 10 | 2.361 | 2.464 | 0.091 | 0.054 |  | 0.858 | Martine Carrier | 2023-12-29
D- | D-1202 | Aniline Point - ASTM D611 | °C | 59.85 | 0.17 | 59.695 | 0.906 | 6 | 58.55 | 61.22 | 2.511 | 0.5 |  | 0.306 | Martine Carrier | 2023-12-29
D- | D-1202 | Total Aromatics - All Methods | Vol. % |  |  | 29.48 | 1.65 | 4 | 27.3 | 30.9 | 4.57 |  |  | 0.375 | Martine Carrier | 2023-12-29
D- | D-1202 | Aromatics - ASTM D1319 | Vol. % |  |  | 29.42 | 2.79 | 5 | 27 | 32.7 | 7.73 |  |  | 0.61 | Martine Carrier | 2023-12-29
D- | D-1202 | Olefins - ASTM D1319 | Vol. % |  |  | 1.28 | 0.79 | 5 | 0.6 | 2.5 | 2.176 |  |  | 0.413 | Martine Carrier | 2023-12-29
D- | D-1202 | Saturates - ASTM D1319 | Vol. % |  |  | 69.3 | 3.35 | 5 | 64.8 | 72.4 | 9.274 |  |  | 0.414 | Martine Carrier | 2023-12-29
D- | D-1202 | Flash Point - ASTM D93 | °F | 144 | -0.15 | 144.4 | 2.6 | 12 | 141 | 150.8 | 7.185 |  |  | 0.599 | Martine Carrier | 2023-12-29
D- | D-1202 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 328.6 | 0.3 | 326.34 | 7.45 | 12 | 308 | 335.42 | 20.641 |  |  | 0.759 | Martine Carrier | 2023-12-29
D- | D-1202 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 397.6 | 1.35 | 394.98 | 1.95 | 10 | 391.2 | 397.63 | 5.388 |  |  | 0.247 | Martine Carrier | 2023-12-29
D- | D-1202 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 497.6 | 1.47 | 495.97 | 1.11 | 11 | 493.9 | 497.6 | 3.075 |  |  | 0.232 | Martine Carrier | 2023-12-29
D- | D-1202 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 606.7 | 0.7 | 604.49 | 3.16 | 12 | 597.3 | 609 | 8.766 |  |  | 0.31 | Martine Carrier | 2023-12-29
D- | D-1202 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 672 | 0.38 | 669.54 | 6.44 | 12 | 655 | 677 | 17.85 |  |  | 0.451 | Martine Carrier | 2023-12-29
D- | D-1202 | Recovery - ASTM D86 | Vol. % | 98 | -0.31 | 98.15 | 0.48 | 12 | 97.3 | 98.9 | 1.334 |  |  | 0.25 | Martine Carrier | 2023-12-29
D- | D-1202 | Residue - ASTM D86 | Vol. % | 0.7 | -1.97 | 1.22 | 0.26 | 12 | 0.7 | 1.5 | 0.727 |  |  | 0.982 | Martine Carrier | 2023-12-29
D- | D-1202 | Loss - ASTM D86 | Vol. % | 1.3 | 1.47 | 0.63 | 0.45 | 12 | 0 | 1.3 | 1.257 |  |  | 0.413 | Martine Carrier | 2023-12-29
D | D328 | Residue - ASTM D86 (Automatic) | Vol. % | 1 | -10.79 | 1.4 | 0.04 | 23 | 1.3 | 1.5 | 0.102 |  |  | 6.19 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 117.6 | 0.27 | 117.2 | 1.47 | 9 | 115.2 | 119.7 | 4.083 | 7.735 |  | 0.173 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 175.8 | 0.04 | 175.76 | 1.19 | 9 | 173.3 | 177.6 | 3.287 | 4.136 |  | 0.403 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 197.6 | -0.38 | 198.23 | 1.66 | 10 | 194.9 | 201.2 | 4.585 | 4.473 |  | 0.322 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 220.2 | -0.5 | 221.01 | 1.63 | 10 | 217.1 | 222.8 | 4.517 | 4.815 |  | 0.673 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 238.8 | -0.17 | 239.1 | 1.79 | 10 | 234.7 | 241.1 | 4.957 | 4.408 |  | 0.763 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 255.8 | -0.27 | 256.31 | 1.86 | 10 | 252 | 258.7 | 5.163 | 4.3 |  | 0.513 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 272 | -0.28 | 272.59 | 2.14 | 10 | 267.5 | 275.1 | 5.927 | 4.3 |  | 0.517 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 288.6 | -0.38 | 289.37 | 2.03 | 10 | 284.8 | 291.8 | 5.635 | 4.3 |  | 0.426 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 306.2 | -0.33 | 306.91 | 2.18 | 10 | 301.9 | 309.2 | 6.03 | 4.3 |  | 0.543 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 327.4 | -0.32 | 328.17 | 2.39 | 10 | 322.9 | 330.2 | 6.613 | 4.3 |  | 0.688 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 353.4 | -0.26 | 353.79 | 1.52 | 10 | 351.5 | 356 | 4.224 | 4.3 |  | 0.232 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 368.2 | -0.35 | 369.36 | 3.36 | 10 | 363.8 | 376.6 | 9.296 | 5 |  | 0.535 | Martine Carrier | 2022-10-21
D | D328 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 425.6 | -1.04 | 429.1 | 3.38 | 8 | 425.6 | 436.8 | 9.367 | 11.8 |  | 0.859 | Martine Carrier | 2022-10-21
D | D328 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 175.3 | 0.32 | 174.06 | 3.87 | 28 | 168.2 | 184.8 | 10.733 | 9.574 | 1.257 - Pass | 0.78 | Martine Carrier | 2022-10-21
D | D328 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 218.6 | 0.67 | 217.8 | 1.18 | 28 | 214.9 | 220.5 | 3.271 | 4.792 | 0.466 - Pass | 0.302 | Martine Carrier | 2022-10-21
D | D328 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 272.3 | 1.46 | 270.78 | 1.05 | 28 | 267.5 | 272.3 | 2.901 | 3 | 0.935 - Pass | 0.549 | Martine Carrier | 2022-10-21
D | D328 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 333.2 | 0.95 | 331.89 | 1.38 | 27 | 328.5 | 334 | 3.821 | 4.978 | 0.589 - Pass | 0.46 | Martine Carrier | 2022-10-21
D | D328 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 362 | 1.25 | 359.63 | 1.9 | 28 | 353.8 | 364.1 | 5.263 | 7.1 | 0.550 - Pass | 0.452 | Martine Carrier | 2022-10-21
D | D328 | Recovery - ASTM D86 (Automatic) | Vol. % | 97.9 | -0.39 | 98.07 | 0.43 | 28 | 97.2 | 99.4 | 1.187 |  |  | 2.01 | Martine Carrier | 2022-10-21
D | D328 | Loss - ASTM D86 (Automatic) | Vol. % | 1.1 | 1.63 | 0.58 | 0.32 | 28 | 0.1 | 1.4 | 0.879 |  |  | 1.291 | Martine Carrier | 2022-10-21
D | D328 | Aniline Point - ASTM D611 | °C | 56.5 | -0.44 | 56.665 | 0.378 | 11 | 56.1 | 57.2 | 1.047 | 0.5 |  | 0.371 | Martine Carrier | 2022-10-21
D | D328 | Cetane Index, Calculated - ASTM D976 |  | 45 | 1.53 | 44.72 | 0.18 | 22 | 44.3 | 45 | 0.51 |  |  | 0.595 | Martine Carrier | 2022-10-21
D | D328 | Cetane Index, Calculated - ASTM D4737 B |  | 43.4 | 0.99 | 43.11 | 0.29 | 15 | 42.6 | 43.6 | 0.803 |  |  | 0.282 | Martine Carrier | 2022-10-21
D | D328 | Cetane Index, Calculated - ASTM D4737 A |  | 44.2 | 0.9 | 43.99 | 0.24 | 22 | 43.5 | 44.4 | 0.658 |  |  | 0.551 | Martine Carrier | 2022-10-21
D | D328 | Cetane Number - ASTM D613 |  |  |  | 45.5 | 1.57 | 5 | 43.5 | 47.8 | 4.353 | 3.488 |  | 0.263 | Martine Carrier | 2022-10-21
D | D328 | Cloud Point - ASTM D2500 | °C |  |  | -14.7 | 1.2 | 9 | -16 | -12 | 3.393 | 4 |  | 0.685 | Martine Carrier | 2022-10-21
D | D328 | Cloud Point - ASTM D5773 | °C | -14.8 | 0.91 | -15.43 | 0.7 | 26 | -17 | -14.4 | 1.94 | 2.5 | 0.602 - Pass | 0.257 | Martine Carrier | 2022-10-21
D | D328 | Density, at 15°C - ASTM D4052 | kg/m³ | 860.5 | 0.21 | 860.48 | 0.09 | 27 | 860.4 | 860.7 | 0.243 | 0.5 | 0.237 - Pass | 2.354 | Martine Carrier | 2022-10-21
D | D328 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 451 | 0.08 | 448.8 | 29 | 23 | 389 | 502 | 80.347 | 48.667 | 2.726 - Fail | 0.195 | Martine Carrier | 2022-10-21
D | D328 | Flash Point - ASTM D93 | °C | 68.4 | 0.15 | 68.15 | 1.72 | 28 | 63.3 | 72 | 4.755 | 4.838 | 0.966 - Pass | 0.597 | Martine Carrier | 2022-10-21
D | D328 | HFRR, Major Axis - ASTM D6079 | µm | 222 | -1.46 | 282.3 | 41.5 | 9 | 220 | 327 | 114.821 |  |  | 0.507 | Martine Carrier | 2022-10-21
D | D328 | HFRR, Minor Axis - ASTM D6079 | µm | 197 | -0.9 | 230.9 | 37.9 | 9 | 190 | 300 | 104.872 |  |  | 0.295 | Martine Carrier | 2022-10-21
D | D328 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 209 | -1.26 | 255.6 | 36.9 | 9 | 209 | 310 | 102.229 | 80 |  | 0.402 | Martine Carrier | 2022-10-21
D | D328 | Pour Point - ASTM D97 | °C |  |  | -26.8 | 2.7 | 9 | -30 | -24 | 7.43 | 9 |  | 0.732 | Martine Carrier | 2022-10-21
D | D328 | Pour Point - ASTM D5949 | °C | -27 | -0.42 | -26.1 | 2.1 | 24 | -30 | -21 | 5.736 | 6.3 | 0.829 - Pass | 1.771 | Martine Carrier | 2022-10-21
D | D328 | Sulfur - ASTM D5453 | mg/kg | 5.5 | 0.25 | 5.38 | 0.49 | 25 | 4.3 | 6.3 | 1.345 | 2.048 | 0.432 - Pass | 0.216 | Martine Carrier | 2022-10-21
D | D328 | Sulfur - ASTM D7039 | mg/kg |  |  | 4.94 | 0.17 | 9 | 4.7 | 5.3 | 0.462 | 1.75 |  | 0.551 | Martine Carrier | 2022-10-21
D | D328 | Viscosity at 40°C - ASTM D445 | cSt | 3.017 | 0.92 | 3.0076 | 0.0103 | 20 | 2.989 | 3.026 | 0.028 | 0.067 | 0.178 - Pass | 0.435 | Martine Carrier | 2022-10-21
D | D328 | Water Content - ASTM D6304, Procedure A | mg/kg | 44 | -0.18 | 44.9 | 4.7 | 20 | 33 | 52 | 13.032 | 4.157 | 9.830 - Fail | 0.263 | Martine Carrier | 2022-10-21
D | D328 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.7258 | 0.2666 | 5 | 2.51 | 3.019 | 0.739 | 0.079 |  | 0.886 | Martine Carrier | 2022-10-21
D | D328 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.9984 | 0.0346 | 9 | 2.922 | 3.027 | 0.096 | 0.079 |  | 0.859 | Martine Carrier | 2022-10-21
D | D328 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 842.7 | 0.2 | 4 | 842.5 | 842.9 | 0.571 | -5891.865 |  | 0.546 | Martine Carrier | 2022-10-21
D | D328 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 859.5 | 1.4 | 3 | 858 | 860.6 | 3.771 | 1.77 |  | 0.435 | Martine Carrier | 2022-10-21
D | D328 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 8.939 | 0.0469 | 3 | 8.886 | 8.975 | 0.13 |  |  | 0.451 | Martine Carrier | 2022-10-21
D | D328 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 4.68 | 1.243 | 5 | 2.5 | 5.54 | 3.443 | 0.934 |  | 0.895 | Martine Carrier | 2022-10-21
D | D329 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 145.6 | -0.82 | 146.72 | 1.37 | 13 | 144.6 | 149.5 | 3.789 | 9.683 | 0.153 - Pass | 0.255 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 166.6 | -1.34 | 168.32 | 1.28 | 13 | 166.2 | 170 | 3.549 | 4.025 | 0.778 - Pass | 0.342 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 177.6 | -1.19 | 179.03 | 1.2 | 13 | 177.1 | 181 | 3.319 | 4.185 | 0.629 - Pass | 0.257 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 188 | -1.03 | 189.22 | 1.18 | 13 | 187.1 | 191 | 3.258 | 4.338 | 0.564 - Pass | 0.243 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 197.8 | -0.78 | 198.85 | 1.35 | 13 | 196.1 | 201 | 3.736 | 3.885 | 0.925 - Pass | 0.292 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 205.4 | -0.81 | 206.46 | 1.31 | 13 | 203.8 | 208.1 | 3.625 | 4.3 | 0.711 - Pass | 0.408 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 214.6 | -0.84 | 215.68 | 1.29 | 13 | 212.9 | 217.4 | 3.574 | 4.3 | 0.691 - Pass | 0.396 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 225 | -0.76 | 226.05 | 1.39 | 13 | 223.2 | 228 | 3.862 | 4.3 | 0.807 - Pass | 0.333 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 238.2 | -0.85 | 239.32 | 1.32 | 13 | 236.7 | 241.1 | 3.655 | 4.3 | 0.723 - Pass | 0.206 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 256.2 | -0.63 | 256.98 | 1.24 | 13 | 254.6 | 259 | 3.434 | 4.3 | 0.638 - Pass | 0.356 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 281.2 | -0.99 | 282.42 | 1.23 | 13 | 280.1 | 284.2 | 3.409 | 4.3 | 0.629 - Pass | 0.158 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 300 | -1.12 | 301.25 | 1.12 | 13 | 299.1 | 303 | 3.109 | 5 | 0.387 - Pass | 0.421 | Martine Carrier | 2022-12-12
D | D329 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 337.4 | -0.97 | 339.58 | 2.26 | 13 | 336.9 | 344.3 | 6.255 | 11.8 | 0.281 - Pass | 1.004 | Martine Carrier | 2022-12-12
D | D329 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 187.1 | 1.57 | 182.41 | 2.99 | 20 | 176.5 | 187.1 | 8.271 | 10.032 | 0.680 - Pass | 0.269 | Martine Carrier | 2022-12-12
D | D329 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 196.7 | 0.36 | 196.49 | 0.58 | 20 | 195.4 | 197.3 | 1.617 | 4.323 | 0.140 - Pass | 0.493 | Martine Carrier | 2022-12-12
D | D329 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 214.8 | -0.77 | 215.24 | 0.57 | 20 | 214.2 | 216.2 | 1.583 | 3 | 0.279 - Pass | 0.249 | Martine Carrier | 2022-12-12
D | D329 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 264.9 | -0.5 | 265.79 | 1.77 | 20 | 261.8 | 268.8 | 4.912 | 3.987 | 1.518 - Pass | 0.431 | Martine Carrier | 2022-12-12
D | D329 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 301.4 | 1.98 | 297.28 | 2.08 | 20 | 291.4 | 301.4 | 5.774 | 7.1 | 0.661 - Pass | 0.718 | Martine Carrier | 2022-12-12
D | D329 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.8 | 1.05 | 98.28 | 0.49 | 20 | 96.9 | 99.1 | 1.371 |  |  | 0.518 | Martine Carrier | 2022-12-12
D | D329 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -0.84 | 1.32 | 0.26 | 20 | 0.8 | 2.1 | 0.722 |  |  | 1.378 | Martine Carrier | 2022-12-12
D | D329 | Loss - ASTM D86 (Automatic) | Vol. % | 0.1 | -0.92 | 0.29 | 0.21 | 18 | 0 | 0.7 | 0.569 |  |  | 0.445 | Martine Carrier | 2022-12-12
D | D329 | Aniline Point - ASTM D611 | °C | 62.15 | 1.81 | 60.924 | 0.677 | 7 | 59.83 | 62.15 | 1.875 | 0.5 |  | 0.7 | Martine Carrier | 2022-12-12
D | D329 | Cetane Index, Calculated - ASTM D976 |  | 40.8 | -0.22 | 40.85 | 0.21 | 13 | 40.6 | 41.2 | 0.572 |  |  | 0.566 | Martine Carrier | 2022-12-12
D | D329 | Cetane Index, Calculated - ASTM D4737 B |  | 44 | -0.36 | 44.07 | 0.19 | 9 | 43.7 | 44.3 | 0.518 |  |  | 0.389 | Martine Carrier | 2022-12-12
D | D329 | Cetane Index, Calculated - ASTM D4737 A |  | 42.9 | 0 | 42.9 | 0.22 | 13 | 42.5 | 43.2 | 0.609 |  |  | 0.369 | Martine Carrier | 2022-12-12
D | D329 | Cetane Number - ASTM D613 |  |  |  | 46.56 | 1.75 | 5 | 44.5 | 49.1 | 4.844 | 3.62 |  | 0.203 | Martine Carrier | 2022-12-12
D | D329 | Cloud Point - ASTM D2500 | °C |  |  | -47.2 | 2.1 | 6 | -51 | -45 | 5.919 | 4 |  | 0.456 | Martine Carrier | 2022-12-12
D | D329 | Cloud Point - ASTM D5773 | °C | -49.6 | -1.11 | -47.63 | 1.77 | 19 | -51 | -44.8 | 4.902 | 2.5 | 3.844 - Fail | 0.313 | Martine Carrier | 2022-12-12
D | D329 | Density, at 15°C - ASTM D4052 | kg/m³ | 824.8 | -1.29 | 824.93 | 0.1 | 24 | 824.7 | 825.2 | 0.277 | 0.5 | 0.306 - Pass | 1.773 | Martine Carrier | 2022-12-12
D | D329 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 1660 | -0.99 | 1759.1 | 99.8 | 16 | 1627 | 1930 | 276.519 | 157.294 | 3.090 - Fail | 0.344 | Martine Carrier | 2022-12-12
D | D329 | Flash Point - ASTM D93 | °C | 65.5 | -1.32 | 67.24 | 1.32 | 21 | 65.4 | 70 | 3.651 | 4.774 | 0.585 - Pass | 0.334 | Martine Carrier | 2022-12-12
D | D329 | HFRR, Major Axis - ASTM D6079 | µm | 487 | -0.37 | 495.9 | 23.8 | 8 | 455 | 540 | 66.007 |  |  | 0.491 | Martine Carrier | 2022-12-12
D | D329 | HFRR, Minor Axis - ASTM D6079 | µm | 428 | 1.26 | 400.6 | 21.8 | 8 | 368 | 428 | 60.342 |  |  | 0.283 | Martine Carrier | 2022-12-12
D | D329 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 457 | 0.46 | 447.5 | 20.7 | 8 | 412 | 480 | 57.383 | 80 |  | 0.214 | Martine Carrier | 2022-12-12
D | D329 | Pour Point - ASTM D97 | °C |  |  | -63.8 | 5.1 | 4 | -69 | -57 | 14.192 | 9 |  | 0.262 | Martine Carrier | 2022-12-12
D | D329 | Pour Point - ASTM D5949 | °C | -78 | -1.87 | -65.4 | 6.7 | 16 | -78 | -54 | 18.65 | 6.3 | 8.763 - Fail | 0.158 | Martine Carrier | 2022-12-12
D | D329 | Sulfur - ASTM D5453 | mg/kg | 1.7 | 0.44 | 1.57 | 0.3 | 22 | 0.8 | 2 | 0.822 | 0.812 | 1.023 - Pass | 0.633 | Martine Carrier | 2022-12-12
D | D329 | Sulfur - ASTM D7039 | mg/kg |  |  | 1.43 | 0.12 | 6 | 1.3 | 1.6 | 0.335 | 0.897 |  | 0.375 | Martine Carrier | 2022-12-12
D | D329 | Viscosity at 40°C - ASTM D445 | cSt | 1.657 | 0.02 | 1.6568 | 0.0109 | 16 | 1.642 | 1.677 | 0.03 | 0.037 | 0.659 - Pass | 0.536 | Martine Carrier | 2022-12-12
D | D329 | Water Content - ASTM D6304, Procedure A | mg/kg | 34 | 1.58 | 23.8 | 6.4 | 15 | 10 | 34 | 17.835 | 2.842 | 39.384 - Fail | 0.287 | Martine Carrier | 2022-12-12
D | D329 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 1.317 | 0.0245 | 6 | 1.271 | 1.335 | 0.068 | 0.038 |  | 0.705 | Martine Carrier | 2022-12-12
D | D329 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 1.6305 | 0.0342 | 10 | 1.574 | 1.679 | 0.095 | 0.049 |  | 0.499 | Martine Carrier | 2022-12-12
D | D329 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 675.1 | 330.4 | 6 | 0.806 | 823 | 915.142 | -4718.147 |  | 1.626 | Martine Carrier | 2022-12-12
D | D329 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 824.2 | 1.9 | 3 | 822 | 825.6 | 5.343 | 1.77 |  | 0.502 | Martine Carrier | 2022-12-12
D | D329 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 0.033 | 0.058 | 3 | 0 | 0.1 | 0.16 | 0.713 |  | 0.732 | Martine Carrier | 2022-12-12
D | D330 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 101.2 | -0.02 | 101.23 | 1.5 | 12 | 98.8 | 104 | 4.158 | 6.681 |  | 0.183 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 151.6 | 0.59 | 151.05 | 0.93 | 12 | 149.4 | 152.4 | 2.589 | 3.766 |  | 0.334 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 168.2 | -0.13 | 168.42 | 1.61 | 12 | 164.7 | 170.4 | 4.455 | 4.026 |  | 0.427 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 196.4 | -0.02 | 196.42 | 0.94 | 11 | 194.5 | 197.7 | 2.595 | 4.446 |  | 0.264 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 218.4 | -0.47 | 218.89 | 1.05 | 11 | 217 | 220.6 | 2.909 | 4.146 |  | 0.172 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 236.8 | -0.54 | 237.5 | 1.29 | 11 | 235.5 | 240 | 3.567 | 4.3 |  | 0.246 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 255.6 | -0.11 | 255.86 | 2.39 | 12 | 249.6 | 259.4 | 6.631 | 4.3 |  | 0.722 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 274.2 | -0.21 | 274.75 | 2.56 | 12 | 268 | 278.6 | 7.103 | 4.3 |  | 0.843 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 293.4 | -1.1 | 294.66 | 1.15 | 11 | 293 | 296.9 | 3.186 | 4.3 |  | 0.162 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 312.4 | -1.25 | 313.95 | 1.23 | 11 | 312.3 | 316.4 | 3.414 | 4.3 |  | 0.24 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 339.4 | -1.25 | 341.02 | 1.29 | 11 | 339 | 343.7 | 3.586 | 4.3 |  | 0.329 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 359 | -0.55 | 360.12 | 2.04 | 12 | 354.8 | 363.2 | 5.638 | 5 |  | 0.779 | Eric Salvail | 2022-12-30
D | D330 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 399.2 | -0.2 | 399.73 | 2.67 | 12 | 394.5 | 403 | 7.395 | 11.8 |  | 0.498 | Eric Salvail | 2022-12-30
D | D330 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 152.3 | 0.5 | 151.01 | 2.6 | 26 | 144.7 | 155.4 | 7.215 | 8.305 | 0.755 - Pass | 0.245 | Eric Salvail | 2022-12-30
D | D330 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 189.3 | 1.7 | 187.21 | 1.23 | 27 | 184.6 | 189.3 | 3.395 | 4.119 | 0.680 - Pass | 0.395 | Eric Salvail | 2022-12-30
D | D330 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 255 | 1.7 | 253.51 | 0.88 | 27 | 251.6 | 255 | 2.432 | 3 | 0.657 - Pass | 0.376 | Eric Salvail | 2022-12-30
D | D330 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 322.6 | 1.2 | 321.02 | 1.32 | 27 | 318.4 | 323.8 | 3.647 | 4.815 | 0.574 - Pass | 0.226 | Eric Salvail | 2022-12-30
D | D330 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 351.7 | 1.06 | 349.8 | 1.8 | 27 | 344.4 | 353.5 | 4.992 | 7.1 | 0.494 - Pass | 0.526 | Eric Salvail | 2022-12-30
D | D330 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.1 | -0.33 | 98.19 | 0.28 | 27 | 97.5 | 98.8 | 0.783 |  |  | 0.506 | Eric Salvail | 2022-12-30
D | D330 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0 | 1.4 | 0 | 22 | 1.4 | 1.4 | 0 |  |  |  | Eric Salvail | 2022-12-30
D | D330 | Loss - ASTM D86 (Automatic) | Vol. % | 0.5 | 0.37 | 0.41 | 0.24 | 27 | 0 | 1.1 | 0.667 |  |  | 0.519 | Eric Salvail | 2022-12-30
D | D330 | Aniline Point - ASTM D611 | °C | 55.1 | 1.93 | 53.609 | 0.774 | 10 | 52.56 | 55.1 | 2.144 | 0.5 |  | 0.347 | Eric Salvail | 2022-12-30
D | D330 | Cetane Index, Calculated - ASTM D976 |  |  |  | 43.15 | 0.27 | 19 | 42.7 | 43.8 | 0.741 |  |  | 1.039 | Eric Salvail | 2022-12-30
D | D330 | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 39.94 | 0.34 | 12 | 39.5 | 40.6 | 0.941 |  |  | 0.445 | Eric Salvail | 2022-12-30
D | D330 | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 41.85 | 0.23 | 18 | 41.5 | 42.3 | 0.639 |  |  | 0.454 | Eric Salvail | 2022-12-30
D | D330 | Cetane Number - ASTM D613 |  |  |  | 40.54 | 1.16 | 7 | 39.2 | 42.6 | 3.214 | 2.868 |  | 0.274 | Eric Salvail | 2022-12-30
D | D330 | Cloud Point - ASTM D2500 | °C |  |  | -36.7 | 1.6 | 11 | -39 | -35 | 4.307 | 4 |  | 0.58 | Eric Salvail | 2022-12-30
D | D330 | Cloud Point - ASTM D5773 | °C | -38 | -0.47 | -37.35 | 1.37 | 26 | -39.2 | -33.4 | 3.796 | 2.5 | 2.305 - Fail | 0.728 | Eric Salvail | 2022-12-30
D | D330 | Density, at 15°C - ASTM D4052 | kg/m³ | 852.6 | -1.53 | 852.74 | 0.09 | 29 | 852.5 | 852.9 | 0.25 | 0.5 | 0.250 - Pass | 1.83 | Eric Salvail | 2022-12-30
D | D330 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 804 | -1.85 | 966.3 | 87.8 | 24 | 804 | 1131 | 243.241 | 91.569 | 7.056 - Fail | 0.174 | Eric Salvail | 2022-12-30
D | D330 | Flash Point - ASTM D93 | °C | 50 | -1.14 | 51.36 | 1.19 | 30 | 49.1 | 54.4 | 3.304 | 3.647 | 0.821 - Pass | 0.403 | Eric Salvail | 2022-12-30
D | D330 | HFRR, Major Axis - ASTM D6079 | µm | 680 | 0.19 | 675.4 | 24.7 | 10 | 640 | 720 | 68.367 |  |  | 0.206 | Eric Salvail | 2022-12-30
D | D330 | HFRR, Minor Axis - ASTM D6079 | µm | 679 | 0.52 | 665 | 26.7 | 10 | 630 | 700 | 74.074 |  |  | 0.591 | Eric Salvail | 2022-12-30
D | D330 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 680 | 0.38 | 670.3 | 25.4 | 10 | 640 | 710 | 70.368 | 80 |  | 0.376 | Eric Salvail | 2022-12-30
D | D330 | Pour Point - ASTM D97 | °C |  |  | -45.9 | 2.5 | 10 | -51 | -42 | 6.841 | 9 |  | 1.041 | Eric Salvail | 2022-12-30
D | D330 | Pour Point - ASTM D5949 | °C | -47 | 0.25 | -47.7 | 2.9 | 24 | -54 | -45 | 7.982 | 6.3 | 1.605 - Pass | 1.426 | Eric Salvail | 2022-12-30
D | D330 | Sulfur - ASTM D5453 | mg/kg | 7.5 | -0.05 | 7.53 | 0.69 | 25 | 6.1 | 9.1 | 1.914 | 2.636 | 0.527 - Pass | 0.847 | Eric Salvail | 2022-12-30
D | D330 | Sulfur - ASTM D7039 | mg/kg |  |  | 7.29 | 0.9 | 9 | 6.5 | 9.3 | 2.502 | 2.158 |  | 0.94 | Eric Salvail | 2022-12-30
D | D330 | Viscosity at 40°C - ASTM D445 | cSt | 2.273 | 1.17 | 2.265 | 0.0068 | 21 | 2.254 | 2.282 | 0.019 | 0.051 | 0.137 - Pass | 0.363 | Eric Salvail | 2022-12-30
D | D330 | Water Content - ASTM D6304, Procedure A | mg/kg | 22 | -0.03 | 22.2 | 5.8 | 18 | 13 | 36 | 16.159 | 2.723 | 35.207 - Fail | 0.395 | Eric Salvail | 2022-12-30
D | D330 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 1.999 | 0.1837 | 10 | 1.834 | 2.277 | 0.509 | 0.058 |  | 1.628 | Eric Salvail | 2022-12-30
D | D330 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.2718 | 0.0138 | 10 | 2.242 | 2.296 | 0.038 | 0.064 |  | 0.553 | Eric Salvail | 2022-12-30
D | D330 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 834.8 | 0.4 | 7 | 834.4 | 835.4 | 1.081 | -5836.59 |  | 0.433 | Eric Salvail | 2022-12-30
D | D330 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 852.2 | 1.2 | 5 | 850 | 853 | 3.452 | 1.77 |  | 0.797 | Eric Salvail | 2022-12-30
D | D330 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 5.9283 | 0.0299 | 3 | 5.896 | 5.955 | 0.083 |  |  | 0.33 | Eric Salvail | 2022-12-30
D | D330 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 0.125 | 0.25 | 4 | 0 | 0.5 | 0.693 | 0.717 |  | 1.098 | Eric Salvail | 2022-12-30
D | D331 | Cetane Index, Calculated - ASTM D4737 A |  | 41.9 | -5.84 | 43.17 | 0.22 | 19 | 42.6 | 43.5 | 0.604 |  |  | 0.876 | Eric Salvail | 2023-03-15
D | D331 | Residue - ASTM D86 (Automatic) | Vol. % | 0.7 | -32.62 | 1.4 | 0.02 | 22 | 1.3 | 1.4 | 0.059 |  |  | 6.531 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 124.7 | -0.71 | 125.38 | 0.97 | 12 | 123.7 | 126.4 | 2.675 | 8.275 |  | 0.707 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 157.3 | -1.18 | 158.18 | 0.75 | 12 | 156.7 | 159 | 2.066 | 3.873 |  | 0.639 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 174.2 | -0.49 | 174.62 | 0.85 | 12 | 172.9 | 176 | 2.341 | 4.119 |  | 0.366 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 201.5 | -0.72 | 202.7 | 1.67 | 12 | 198.7 | 205 | 4.617 | 4.541 |  | 0.858 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 222.2 | -0.74 | 223.6 | 1.9 | 12 | 219.1 | 226.6 | 5.265 | 4.207 |  | 0.539 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 240.5 | -0.58 | 241.8 | 2.23 | 12 | 236.4 | 245 | 6.163 | 4.3 |  | 0.595 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 257.7 | -0.6 | 258.99 | 2.15 | 12 | 254 | 262.3 | 5.946 | 4.3 |  | 0.518 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 274.1 | -0.66 | 275.59 | 2.27 | 12 | 270.3 | 279 | 6.279 | 4.3 |  | 0.672 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 291.1 | -0.92 | 293.18 | 2.26 | 12 | 288.1 | 297 | 6.253 | 4.3 |  | 0.494 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 311.3 | -0.8 | 313.13 | 2.29 | 12 | 308.1 | 317 | 6.342 | 4.3 |  | 0.511 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 342.5 | -0.72 | 344.22 | 2.38 | 12 | 339 | 349 | 6.597 | 4.3 |  | 0.618 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 354.4 | -0.8 | 355.26 | 1.08 | 12 | 353.7 | 357 | 2.984 | 5 |  | 0.42 | Eric Salvail | 2023-03-15
D | D331 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 394.5 | -0.73 | 398.03 | 4.84 | 12 | 389.8 | 407.6 | 13.402 | 11.8 |  | 0.297 | Eric Salvail | 2023-03-15
D | D331 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 165.4 | 0.31 | 164.42 | 3.15 | 26 | 155.6 | 170.6 | 8.732 | 9.043 | 0.932 - Pass | 0.626 | Eric Salvail | 2023-03-15
D | D331 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 196.3 | 0.3 | 196.02 | 0.92 | 25 | 194.7 | 197.9 | 2.553 | 4.312 | 0.350 - Pass | 0.423 | Eric Salvail | 2023-03-15
D | D331 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 257.9 | 0.69 | 257.45 | 0.65 | 25 | 256.4 | 258.5 | 1.81 | 3 | 0.364 - Pass | 0.702 | Eric Salvail | 2023-03-15
D | D331 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 320.9 | 0.36 | 320.38 | 1.47 | 25 | 317.6 | 324.4 | 4.07 | 4.806 | 0.717 - Pass | 0.257 | Eric Salvail | 2023-03-15
D | D331 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 348 | 1.85 | 343.41 | 2.48 | 27 | 335.6 | 348 | 6.862 | 7.1 | 0.934 - Pass | 0.723 | Eric Salvail | 2023-03-15
D | D331 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.3 | 0.42 | 98.16 | 0.34 | 27 | 97.4 | 98.8 | 0.939 |  |  | 0.367 | Eric Salvail | 2023-03-15
D | D331 | Loss - ASTM D86 (Automatic) | Vol. % | 1 | 1.59 | 0.5 | 0.31 | 27 | 0.1 | 1.2 | 0.864 |  |  | 0.632 | Eric Salvail | 2023-03-15
D | D331 | Aniline Point - ASTM D611 | °C | 57.2 | 2.31 | 56.178 | 0.443 | 9 | 55.6 | 57.2 | 1.227 | 0.5 |  | 0.706 | Eric Salvail | 2023-03-15
D | D331 | Cetane Index, Calculated - ASTM D976 |  | 44.6 | 0.86 | 44.43 | 0.2 | 20 | 44 | 44.7 | 0.547 |  |  | 0.678 | Eric Salvail | 2023-03-15
D | D331 | Cetane Index, Calculated - ASTM D4737 B |  | 43.3 | 1.75 | 41.99 | 0.75 | 14 | 40.7 | 43.8 | 2.069 |  |  | 1.47 | Eric Salvail | 2023-03-15
D | D331 | Cetane Number - ASTM D613 |  |  |  | 42.36 | 0.63 | 7 | 41.1 | 42.9 | 1.736 | 3.095 |  | 0.592 | Eric Salvail | 2023-03-15
D | D331 | Cloud Point - ASTM D2500 | °C |  |  | -33.8 | 0.9 | 10 | -35 | -32 | 2.545 | 4 |  | 0.691 | Eric Salvail | 2023-03-15
D | D331 | Cloud Point - ASTM D5773 | °C | -33 | 0.99 | -34.17 | 1.18 | 26 | -36.5 | -32.5 | 3.279 | 2.5 | 1.720 - Pass | 0.791 | Eric Salvail | 2023-03-15
D | D331 | Density, at 15°C - ASTM D4052 | kg/m³ | 851.8 | 0.43 | 851.77 | 0.06 | 29 | 851.7 | 851.9 | 0.18 | 0.5 | 0.129 - Pass | 3.049 | Eric Salvail | 2023-03-15
D | D331 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 212 | -0.74 | 228.2 | 22 | 25 | 195 | 281 | 60.888 | 30.381 | 4.017 - Fail | 0.206 | Eric Salvail | 2023-03-15
D | D331 | Flash Point - ASTM D93 | °C | 61 | -1.47 | 62.56 | 1.06 | 30 | 60.6 | 64.5 | 2.94 | 4.442 | 0.438 - Pass | 0.397 | Eric Salvail | 2023-03-15
D | D331 | HFRR, Major Axis - ASTM D6079 | µm |  |  | 262.2 | 45.2 | 9 | 219 | 360 | 125.285 |  |  | 1.024 | Eric Salvail | 2023-03-15
D | D331 | HFRR, Minor Axis - ASTM D6079 | µm |  |  | 209 | 52.3 | 9 | 150 | 320 | 144.863 |  |  | 0.395 | Eric Salvail | 2023-03-15
D | D331 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 225 | -0.23 | 235.1 | 44.7 | 10 | 190 | 340 | 123.895 | 80 |  | 0.839 | Eric Salvail | 2023-03-15
D | D331 | Pour Point - ASTM D97 | °C |  |  | -58 | 1.5 | 6 | -60 | -57 | 4.291 | 9 |  | 1.301 | Eric Salvail | 2023-03-15
D | D331 | Pour Point - ASTM D5949 | °C | -59 | -1.98 | -53.6 | 2.7 | 24 | -60 | -51 | 7.571 | 6.3 | 1.444 - Pass | 1.336 | Eric Salvail | 2023-03-15
D | D331 | Sulfur - ASTM D5453 | mg/kg | 1.6 | 0.09 | 1.57 | 0.31 | 24 | 0.9 | 2.3 | 0.859 | 0.813 | 1.114 - Pass | 0.376 | Eric Salvail | 2023-03-15
D | D331 | Sulfur - ASTM D7039 | mg/kg |  |  | 1.39 | 0.28 | 9 | 0.9 | 1.8 | 0.777 | 0.882 |  | 0.221 | Eric Salvail | 2023-03-15
D | D331 | Viscosity at 40°C - ASTM D445 | cSt | 2.469 | 0.07 | 2.4684 | 0.0082 | 19 | 2.45 | 2.485 | 0.023 | 0.055 | 0.169 - Pass | 0.334 | Eric Salvail | 2023-03-15
D | D331 | Water Content - ASTM D6304, Procedure A | mg/kg | 34 | 2.66 | 19.4 | 5.5 | 16 | 9 | 34 | 15.235 | 2.512 | 36.783 - Fail | 0.799 | Eric Salvail | 2023-03-15 |  | recheck demandé
D | D331 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.0584 | 0.0115 | 7 | 2.036 | 2.069 | 0.032 | 0.06 |  | 0.506 | Eric Salvail | 2023-03-15
D | D331 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.4718 | 0.0116 | 13 | 2.444 | 2.486 | 0.032 | 0.068 | 0.221 - Pass | 0.522 | Eric Salvail | 2023-03-15
D | D331 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 833.7 | 0.3 | 8 | 833.1 | 834 | 0.79 | -5828.778 |  | 0.622 | Eric Salvail | 2023-03-15
D | D331 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 850.9 | 1.2 | 5 | 849 | 851.8 | 3.197 | 1.77 |  | 0.564 | Eric Salvail | 2023-03-15
D | D331 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 6.6627 | 0.039 | 3 | 6.623 | 6.701 | 0.108 |  |  | 0.286 | Eric Salvail | 2023-03-15
D | D331 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 4.557 | 0.387 | 7 | 4.2 | 5.14 | 1.071 | 0.928 |  | 0.517 | Eric Salvail | 2023-03-15
D | D332 | Residue - ASTM D86 (Automatic) | Vol. % | 1 | -5.37 | 1.33 | 0.06 | 27 | 1.2 | 1.5 | 0.172 |  |  | 3.598 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 96.6 | 0.15 | 96.28 | 2.11 | 12 | 91.5 | 98.5 | 5.844 | 6.354 |  | 0.737 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 146.7 | -0.74 | 147.43 | 0.98 | 12 | 146 | 149.4 | 2.709 | 3.711 |  | 0.919 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 161.3 | -1.4 | 162.32 | 0.73 | 12 | 161.3 | 163.7 | 2.018 | 3.935 |  | 0.202 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 181.8 | -0.89 | 182.62 | 0.92 | 12 | 180.9 | 184 | 2.541 | 4.239 |  | 0.25 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 200.4 | -0.64 | 201.05 | 1.02 | 12 | 199 | 203 | 2.821 | 3.914 |  | 0.343 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 217.4 | -0.36 | 217.72 | 0.88 | 12 | 215.7 | 219 | 2.443 | 4.3 |  | 0.436 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 233.2 | -0.79 | 234.25 | 1.33 | 12 | 230.9 | 236 | 3.672 | 4.3 |  | 0.678 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 250.4 | -0.55 | 251.11 | 1.29 | 12 | 248 | 253 | 3.585 | 4.3 |  | 0.449 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 268.2 | -0.54 | 269.03 | 1.56 | 12 | 265.1 | 271 | 4.308 | 4.3 |  | 0.622 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 288.2 | -0.46 | 288.68 | 1.05 | 12 | 286.3 | 290.1 | 2.922 | 4.3 |  | 0.252 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 313.9 | -0.76 | 315.45 | 2.06 | 13 | 311.7 | 321 | 5.697 | 4.3 | 1.755 - Pass | 1.045 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 337.9 | -0.96 | 339.65 | 1.81 | 13 | 335.9 | 343.6 | 5.017 | 5 | 1.007 - Pass | 0.394 | Eric Salvail | 2023-05-19
D | D332 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 386.6 | -1.57 | 388.24 | 1.04 | 12 | 386.3 | 389.3 | 2.889 | 11.8 |  | 0.742 | Eric Salvail | 2023-05-19
D | D332 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 151.6 | 1.16 | 147.68 | 3.39 | 28 | 141.5 | 153.2 | 9.377 | 8.123 | 1.333 - Pass | 0.359 | Eric Salvail | 2023-05-19
D | D332 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 180.7 | -0.6 | 181.25 | 0.92 | 30 | 179.5 | 183.6 | 2.555 | 3.988 | 0.411 - Pass | 0.313 | Eric Salvail | 2023-05-19
D | D332 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 230.3 | -0.86 | 230.99 | 0.8 | 30 | 229.9 | 233.1 | 2.218 | 3 | 0.547 - Pass | 0.607 | Eric Salvail | 2023-05-19
D | D332 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 296.5 | -1.5 | 299.12 | 1.75 | 30 | 296.5 | 303.6 | 4.842 | 4.487 | 1.165 - Pass | 0.749 | Eric Salvail | 2023-05-19
D | D332 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 339.8 | 2.88 | 335.03 | 1.65 | 29 | 331.6 | 339.8 | 4.583 | 7.1 | 0.417 - Pass | 0.383 | Eric Salvail | 2023-05-19
D | D332 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.8 | 2.24 | 98.1 | 0.31 | 30 | 97.1 | 98.8 | 0.864 |  |  | 0.734 | Eric Salvail | 2023-05-19
D | D332 | Loss - ASTM D86 (Automatic) | Vol. % | 0.2 | -1.58 | 0.55 | 0.22 | 29 | 0.2 | 1.1 | 0.609 |  |  | 0.831 | Eric Salvail | 2023-05-19
D | D332 | Aniline Point - ASTM D611 | °C | 61.25 | -1.85 | 62.481 | 0.664 | 9 | 61.25 | 63.75 | 1.839 | 0.5 |  | 0.468 | Eric Salvail | 2023-05-19
D | D332 | Cetane Index, Calculated - ASTM D976 |  | 46.2 | -0.63 | 46.36 | 0.25 | 21 | 46.1 | 47 | 0.687 |  |  | 1.151 | Eric Salvail | 2023-05-19
D | D332 | Cetane Index, Calculated - ASTM D4737 B |  | 46.1 | -1.24 | 46.35 | 0.2 | 13 | 46.1 | 46.7 | 0.55 |  |  | 0.727 | Eric Salvail | 2023-05-19
D | D332 | Cetane Index, Calculated - ASTM D4737 A |  | 46.5 | -0.9 | 46.75 | 0.28 | 22 | 46.4 | 47.4 | 0.768 |  |  | 1.058 | Eric Salvail | 2023-05-19
D | D332 | Cetane Number - ASTM D613 |  |  |  | 45.13 | 1.58 | 10 | 41.7 | 47.2 | 4.366 | 3.441 |  | 0.412 | Eric Salvail | 2023-05-19
D | D332 | Cloud Point - ASTM D2500 | °C |  |  | -35.7 | 2.5 | 10 | -38 | -30 | 7.038 | 4 |  | 0.846 | Eric Salvail | 2023-05-19
D | D332 | Cloud Point - ASTM D5773 | °C | -36.7 | -0.46 | -36.21 | 1.06 | 28 | -38.1 | -33.6 | 2.928 | 2.5 | 1.372 - Pass | 0.349 | Eric Salvail | 2023-05-19
D | D332 | Density, at 15°C - ASTM D4052 | kg/m³ | 824.7 | -0.84 | 824.76 | 0.07 | 34 | 824.6 | 824.9 | 0.205 | 0.5 | 0.168 - Pass | 2.839 | Eric Salvail | 2023-05-19
D | D332 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 903 | 0.22 | 882.9 | 91.6 | 27 | 740 | 1110 | 253.859 | 84.654 | 8.993 - Fail | 0.389 | Eric Salvail | 2023-05-19
D | D332 | Flash Point - ASTM D93 | °C | 40.1 | -0.81 | 41.79 | 2.08 | 30 | 39 | 48 | 5.749 | 2.967 | 3.755 - Fail | 1.469 | Eric Salvail | 2023-05-19
D | D332 | HFRR, Major Axis - ASTM D6079 | µm | 381 | -0.31 | 387 | 19.3 | 12 | 340 | 410 | 53.374 |  |  | 0.771 | Eric Salvail | 2023-05-19
D | D332 | HFRR, Minor Axis - ASTM D6079 | µm | 331 | -0.45 | 342.5 | 25.3 | 12 | 300 | 390 | 70.21 |  |  | 0.155 | Eric Salvail | 2023-05-19
D | D332 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 356 | -0.47 | 365.8 | 20.8 | 11 | 320 | 390 | 57.638 | 80 |  | 0.53 | Eric Salvail | 2023-05-19
D | D332 | Pour Point - ASTM D97 | °C |  |  | -41.7 | 2.6 | 10 | -45 | -39 | 7.276 | 9 |  | 0.876 | Eric Salvail | 2023-05-19
D | D332 | Pour Point - ASTM D5949 | °C | -42 | 0.6 | -43.4 | 2.3 | 25 | -48 | -38 | 6.239 | 6.3 | 0.981 - Pass | 0.88 | Eric Salvail | 2023-05-19
D | D332 | Sulfur - ASTM D5453 | mg/kg | 7.1 | -0.25 | 7.24 | 0.56 | 29 | 6.2 | 8.5 | 1.564 | 2.559 | 0.374 - Pass | 0.572 | Eric Salvail | 2023-05-19
D | D332 | Sulfur - ASTM D7039 | mg/kg |  |  | 6.98 | 0.79 | 9 | 5.2 | 8.2 | 2.193 | 2.108 |  | 0.88 | Eric Salvail | 2023-05-19
D | D332 | Viscosity at 40°C - ASTM D445 | cSt | 1.816 | 0.02 | 1.8158 | 0.0108 | 20 | 1.801 | 1.841 | 0.03 | 0.041 | 0.537 - Pass | 0.651 | Eric Salvail | 2023-05-19
D | D332 | Water Content - ASTM D6304, Procedure A | mg/kg | 68 | 2.47 | 36.9 | 12.6 | 21 | 3 | 68 | 34.894 | 3.698 | 89.056 - Fail | 0.646 | Eric Salvail | 2023-05-19 |  | recheck demandé
D | D332 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 1.4504 | 0.0118 | 9 | 1.427 | 1.465 | 0.033 | 0.042 |  | 0.526 | Eric Salvail | 2023-05-19
D | D332 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 1.7936 | 0.0338 | 14 | 1.719 | 1.855 | 0.094 | 0.053 | 3.111 - Fail | 0.464 | Eric Salvail | 2023-05-19
D | D332 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 806.5 | 0.3 | 6 | 806.2 | 806.8 | 0.701 | -5638.29 |  | 0.371 | Eric Salvail | 2023-05-19
D | D332 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 824.1 | 1.2 | 5 | 822 | 824.9 | 3.287 | 1.77 |  | 0.783 | Eric Salvail | 2023-05-19
D | D332 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 4.1658 | 0.0224 | 4 | 4.136 | 4.185 | 0.062 |  |  | 0.38 | Eric Salvail | 2023-05-19
D | D332 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 0.14 | 0.218 | 4 | 0 | 0.46 | 0.605 | 0.718 |  | 0.688 | Eric Salvail | 2023-05-19
D | D333 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 111.2 | 0.36 | 109.9 | 3.62 | 12 | 100.5 | 114.1 | 10.02 | 7.253 |  | 0.672 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 160.8 | -0.17 | 161.01 | 1.21 | 12 | 159.3 | 163.1 | 3.34 | 3.915 |  | 0.297 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 179.5 | 0.26 | 179.08 | 1.58 | 12 | 175 | 181 | 4.389 | 4.186 |  | 0.621 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 216.6 | 0.27 | 216.21 | 1.43 | 12 | 213 | 218 | 3.975 | 4.743 |  | 0.436 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 245.5 | 0 | 245.5 | 2.2 | 12 | 241.3 | 248.7 | 6.086 | 4.492 |  | 0.279 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 268.8 | -0.03 | 268.85 | 1.78 | 12 | 265.5 | 271.5 | 4.943 | 4.3 |  | 0.261 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 287.4 | 0 | 287.4 | 1.38 | 12 | 285.2 | 289.9 | 3.829 | 4.3 |  | 0.142 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 303.2 | -0.26 | 303.56 | 1.38 | 12 | 301.1 | 305.6 | 3.83 | 4.3 |  | 0.252 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 319.4 | -0.32 | 319.99 | 1.82 | 12 | 316.8 | 323.1 | 5.046 | 4.3 |  | 0.221 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 340.1 | -0.39 | 340.88 | 1.98 | 12 | 336.9 | 343.9 | 5.472 | 4.3 |  | 0.275 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 356.9 | -0.17 | 357.18 | 1.62 | 12 | 354.1 | 359.7 | 4.485 | 4.3 |  | 0.232 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 372.6 | -0.4 | 373.31 | 1.78 | 12 | 370 | 376.8 | 4.936 | 5 |  | 0.174 | Eric Salvail | 2023-07-14
D | D333 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 405.8 | 0.12 | 405.43 | 2.94 | 12 | 401.3 | 412.6 | 8.157 | 11.8 |  | 0.454 | Eric Salvail | 2023-07-14
D | D333 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 158.2 | -0.14 | 158.6 | 2.89 | 29 | 153.7 | 166.8 | 8.006 | 8.723 | 0.842 - Pass | 0.679 | Eric Salvail | 2023-07-14
D | D333 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 202 | 0.68 | 201.05 | 1.39 | 29 | 198.5 | 204.1 | 3.862 | 4.423 | 0.762 - Pass | 0.174 | Eric Salvail | 2023-07-14
D | D333 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 284.3 | 1.25 | 283.03 | 1.01 | 29 | 280.3 | 284.6 | 2.802 | 3 | 0.872 - Pass | 0.204 | Eric Salvail | 2023-07-14
D | D333 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 336.1 | 0.1 | 336 | 0.96 | 28 | 334.6 | 338.5 | 2.664 | 5.04 | 0.279 - Pass | 0.393 | Eric Salvail | 2023-07-14
D | D333 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 357.3 | 0.47 | 356.32 | 2.09 | 27 | 351 | 361 | 5.78 | 7.1 | 0.663 - Pass | 0.255 | Eric Salvail | 2023-07-14
D | D333 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.3 | 0.69 | 98.13 | 0.25 | 28 | 97.8 | 98.8 | 0.706 |  |  | 0.627 | Eric Salvail | 2023-07-14
D | D333 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0.28 | 1.39 | 0.05 | 23 | 1.2 | 1.4 | 0.127 |  |  | 6.051 | Eric Salvail | 2023-07-14
D | D333 | Loss - ASTM D86 (Automatic) | Vol. % | 0.3 | -0.9 | 0.53 | 0.25 | 28 | 0 | 1 | 0.703 |  |  | 0.362 | Eric Salvail | 2023-07-14
D | D333 | Aniline Point - ASTM D611 | °C | 61.8 | 0.34 | 61.644 | 0.463 | 9 | 61.25 | 62.4 | 1.281 | 0.5 |  | 0.851 | Eric Salvail | 2023-07-14
D | D333 | Cetane Index, Calculated - ASTM D976 |  |  |  | 50.95 | 0.18 | 18 | 50.5 | 51.2 | 0.505 |  |  | 0.447 | Eric Salvail | 2023-07-14
D | D333 | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 47.77 | 0.55 | 12 | 46.7 | 49 | 1.528 |  |  | 0.413 | Eric Salvail | 2023-07-14
D | D333 | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 48.78 | 0.36 | 19 | 47.9 | 49.6 | 0.997 |  |  | 0.497 | Eric Salvail | 2023-07-14
D | D333 | Cetane Number - ASTM D613 |  |  |  | 47.53 | 0.65 | 8 | 46.6 | 48.2 | 1.794 | 3.741 |  | 0.439 | Eric Salvail | 2023-07-14
D | D333 | Cloud Point - ASTM D2500 | °C |  |  | -12.6 | 1.3 | 9 | -14 | -10 | 3.693 | 4 |  | 0.629 | Eric Salvail | 2023-07-14
D | D333 | Cloud Point - ASTM D5773 | °C | -13.2 | 0 | -13.2 | 1.02 | 25 | -15.2 | -11.3 | 2.816 | 2.5 | 1.269 - Pass | 0.208 | Eric Salvail | 2023-07-14
D | D333 | Density, at 15°C - ASTM D4052 | kg/m³ | 848.3 | 0.74 | 848.24 | 0.09 | 31 | 848 | 848.4 | 0.243 | 0.5 | 0.236 - Pass | 3.221 | Eric Salvail | 2023-07-14
D | D333 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 1004 | 1.47 | 860.4 | 97.7 | 24 | 728 | 1073 | 270.63 | 82.788 | 10.686 - Fail | 1.107 | Eric Salvail | 2023-07-14
D | D333 | Flash Point - ASTM D93 | °C | 59.2 | -0.13 | 59.32 | 0.89 | 29 | 56.5 | 61 | 2.464 | 4.212 | 0.342 - Pass | 0.624 | Eric Salvail | 2023-07-14
D | D333 | HFRR, Major Axis - ASTM D6079 | µm | 230 | -0.9 | 292.9 | 70.2 | 10 | 210 | 425 | 194.574 |  |  | 0.327 | Eric Salvail | 2023-07-14
D | D333 | HFRR, Minor Axis - ASTM D6079 | µm | 252 | 0.21 | 239.8 | 59 | 10 | 160 | 313 | 163.541 |  |  | 0.532 | Eric Salvail | 2023-07-14
D | D333 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 208 | -0.84 | 263.1 | 65.7 | 10 | 180 | 350 | 181.864 | 80 |  | 0.517 | Eric Salvail | 2023-07-14
D | D333 | Pour Point - ASTM D97 | °C |  |  | -22 | 3 | 9 | -27 | -18 | 8.31 | 9 |  | 0.449 | Eric Salvail | 2023-07-14
D | D333 | Pour Point - ASTM D5949 | °C | -27 | -1.76 | -20.9 | 3.5 | 23 | -28 | -18 | 9.592 | 6.3 | 2.318 - Fail | 1.781 | Eric Salvail | 2023-07-14
D | D333 | Sulfur - ASTM D5453 | mg/kg | 7.1 | 0.45 | 6.87 | 0.51 | 26 | 5.7 | 8 | 1.418 | 2.46 | 0.333 - Pass | 0.204 | Eric Salvail | 2023-07-14
D | D333 | Sulfur - ASTM D7039 | mg/kg |  |  | 6.57 | 0.73 | 10 | 4.9 | 7.5 | 2.01 | 2.041 |  | 0.598 | Eric Salvail | 2023-07-14
D | D333 | Viscosity at 40°C - ASTM D445 | cSt | 3.002 | -0.16 | 3.0044 | 0.0156 | 18 | 2.981 | 3.036 | 0.043 | 0.067 | 0.414 - Pass | 0.508 | Eric Salvail | 2023-07-14
D | D333 | Water Content - ASTM D6304, Procedure A | mg/kg | 44 | 0.57 | 40.2 | 6.7 | 18 | 30 | 55 | 18.6 | 3.89 | 22.858 - Fail | 0.37 | Eric Salvail | 2023-07-14
D | D333 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.6208 | 0.2402 | 8 | 2.478 | 3.026 | 0.665 | 0.076 |  | 1.805 | Eric Salvail | 2023-07-14
D | D333 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.0026 | 0.0096 | 13 | 2.983 | 3.016 | 0.027 | 0.079 | 0.113 - Pass | 0.425 | Eric Salvail | 2023-07-14
D | D333 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 830.5 | 0.2 | 7 | 830.2 | 830.8 | 0.573 | -5806.59 |  | 0.307 | Eric Salvail | 2023-07-14
D | D333 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 847.8 | 1.2 | 4 | 846 | 848.4 | 3.28 | 1.77 |  | 0.976 | Eric Salvail | 2023-07-14
D | D333 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 8.8623 | 0.0348 | 4 | 8.816 | 8.889 | 0.096 |  |  | 0.472 | Eric Salvail | 2023-07-14
D | D333 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 4.816 | 0.465 | 5 | 4.34 | 5.47 | 1.287 | 0.941 |  | 0.3 | Eric Salvail | 2023-07-14
D | D334 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 102.5 | -0.51 | 103.37 | 1.71 | 13 | 100.5 | 106.3 | 4.739 | 6.822 | 0.483 - Pass | 0.353 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 156.3 | -0.32 | 156.78 | 1.49 | 13 | 154.1 | 160.6 | 4.138 | 3.852 | 1.154 - Pass | 0.711 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 176.5 | -0.12 | 176.72 | 1.78 | 13 | 173.2 | 181.1 | 4.928 | 4.151 | 1.409 - Pass | 0.787 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 203.1 | -0.41 | 203.88 | 1.92 | 13 | 199.9 | 208 | 5.31 | 4.558 | 1.357 - Pass | 0.469 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 225.2 | -0.35 | 225.98 | 2.23 | 13 | 221.1 | 230.8 | 6.169 | 4.238 | 2.119 - Fail | 0.732 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 247.1 | -0.19 | 247.53 | 2.32 | 13 | 242.5 | 252.6 | 6.414 | 4.3 | 2.225 - Fail | 0.773 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 268.2 | -0.15 | 268.55 | 2.31 | 13 | 263.2 | 273.2 | 6.392 | 4.3 | 2.209 - Fail | 0.685 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 288.4 | -0.24 | 288.89 | 2.06 | 13 | 284.4 | 292.8 | 5.695 | 4.3 | 1.754 - Pass | 0.55 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 308.2 | -0.32 | 308.86 | 2.07 | 13 | 304 | 312.3 | 5.723 | 4.3 | 1.772 - Pass | 0.386 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 330.2 | -0.28 | 330.78 | 2.05 | 13 | 326.1 | 334.1 | 5.687 | 4.3 | 1.749 - Pass | 0.417 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 360.4 | -0.3 | 361.09 | 2.34 | 13 | 356.7 | 364.7 | 6.468 | 4.3 | 2.263 - Fail | 0.447 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 385.2 | -0.1 | 385.5 | 2.93 | 13 | 378.9 | 389.6 | 8.11 | 5 | 2.631 - Fail | 0.277 | Eric Salvail | 2023-09-25
D | D334 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 480.3 | 0.61 | 476.65 | 5.98 | 11 | 460.5 | 482 | 16.56 | 11.8 |  | 1.118 | Eric Salvail | 2023-09-25
D | D334 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 153 | 0.62 | 151.53 | 2.39 | 29 | 146.8 | 157.5 | 6.611 | 8.334 | 0.629 - Pass | 0.444 | Eric Salvail | 2023-09-25
D | D334 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 195.9 | 1.03 | 194.86 | 1.01 | 29 | 193.5 | 197.6 | 2.802 | 4.287 | 0.427 - Pass | 0.448 | Eric Salvail | 2023-09-25
D | D334 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 264.4 | 1.05 | 263.37 | 0.98 | 29 | 260.8 | 265.2 | 2.713 | 3 | 0.818 - Pass | 1.163 | Eric Salvail | 2023-09-25
D | D334 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 342 | 0.63 | 340.9 | 1.76 | 29 | 336.5 | 343.7 | 4.874 | 5.114 | 0.909 - Pass | 0.719 | Eric Salvail | 2023-09-25
D | D334 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 374.8 | -0.83 | 377.27 | 2.98 | 29 | 371.1 | 384.5 | 8.265 | 7.1 | 1.355 - Pass | 0.187 | Eric Salvail | 2023-09-25
D | D334 | Recovery - ASTM D86 (Automatic) | Vol. % | 97.2 | -1.74 | 97.79 | 0.34 | 29 | 97.2 | 98.8 | 0.94 |  |  | 0.405 | Eric Salvail | 2023-09-25
D | D334 | Residue - ASTM D86 (Automatic) | Vol. % | 1.5 | 0.38 | 1.46 | 0.1 | 26 | 1.2 | 1.7 | 0.283 |  |  | 2.673 | Eric Salvail | 2023-09-25
D | D334 | Loss - ASTM D86 (Automatic) | Vol. % | 1.3 | 1.57 | 0.76 | 0.34 | 29 | 0 | 1.5 | 0.955 |  |  | 0.454 | Eric Salvail | 2023-09-25
D | D334 | Aniline Point - ASTM D611 | °C | 59.55 | -2.28 | 60.731 | 0.519 | 9 | 59.55 | 61.5 | 1.437 | 0.5 |  | 0.817 | Eric Salvail | 2023-09-25 |  | recheck demandé
D | D334 | Cetane Index, Calculated - ASTM D976 |  | 45.8 | 1.16 | 45.6 | 0.17 | 20 | 45.2 | 45.8 | 0.476 |  |  | 1.069 | Eric Salvail | 2023-09-25
D | D334 | Cetane Index, Calculated - ASTM D4737 B |  | 43.7 | 1.26 | 43.44 | 0.21 | 14 | 43 | 43.7 | 0.581 |  |  | 0.734 | Eric Salvail | 2023-09-25
D | D334 | Cetane Index, Calculated - ASTM D4737 A |  | 44.6 | 0.93 | 44.41 | 0.2 | 21 | 44.1 | 44.8 | 0.556 |  |  | 0.628 | Eric Salvail | 2023-09-25
D | D334 | Cetane Number - ASTM D613 |  |  |  | 41.48 | 1.41 | 9 | 39.7 | 43.7 | 3.917 | 2.985 |  | 0.345 | Eric Salvail | 2023-09-25
D | D334 | Cloud Point - ASTM D2500 | °C |  |  | -21.3 | 0.8 | 10 | -23 | -20 | 2.28 | 4 |  | 1.041 | Eric Salvail | 2023-09-25
D | D334 | Cloud Point - ASTM D5773 | °C | -21.1 | 0.63 | -21.69 | 0.93 | 27 | -23.2 | -20 | 2.581 | 2.5 | 1.066 - Pass | 0.621 | Eric Salvail | 2023-09-25
D | D334 | Density, at 15°C - ASTM D4052 | kg/m³ | 852.5 | 0.07 | 852.49 | 0.13 | 34 | 852.3 | 852.8 | 0.35 | 0.5 | 0.490 - Pass | 2.578 | Eric Salvail | 2023-09-25
D | D334 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 590 | 0.44 | 571.5 | 41.9 | 27 | 464 | 640 | 116.145 | 58.839 | 3.897 - Fail | 0.356 | Eric Salvail | 2023-09-25
D | D334 | Flash Point - ASTM D93 | °C | 53.1 | -0.94 | 53.97 | 0.92 | 32 | 52 | 56.5 | 2.557 | 3.832 | 0.445 - Pass | 0.556 | Eric Salvail | 2023-09-25
D | D334 | HFRR, Major Axis - ASTM D6079 | µm | 560 | -1.19 | 575.5 | 13 | 12 | 550 | 600 | 36.049 |  |  | 0.304 | Eric Salvail | 2023-09-25
D | D334 | HFRR, Minor Axis - ASTM D6079 | µm | 564 | 0.23 | 559.8 | 18.9 | 12 | 530 | 600 | 52.279 |  |  | 0.3 | Eric Salvail | 2023-09-25
D | D334 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 562 | -0.43 | 568.4 | 14.9 | 12 | 540 | 600 | 41.144 | 80 |  | 0.371 | Eric Salvail | 2023-09-25
D | D334 | Pour Point - ASTM D97 | °C |  |  | -41 | 2.6 | 9 | -45 | -36 | 7.197 | 9 |  | 0.792 | Eric Salvail | 2023-09-25
D | D334 | Pour Point - ASTM D5949 | °C | -55 | -2.85 | -40.1 | 5.2 | 26 | -55 | -33 | 14.53 | 6.3 | 5.319 - Fail | 0.804 | Eric Salvail | 2023-09-25 |  | recheck demandé
D | D334 | Sulfur - ASTM D5453 | mg/kg | 8 | 0.97 | 7.42 | 0.6 | 29 | 6.2 | 8.7 | 1.65 | 2.607 | 0.400 - Pass | 0.275 | Eric Salvail | 2023-09-25
D | D334 | Sulfur - ASTM D7039 | mg/kg |  |  | 7.28 | 0.8 | 9 | 5.8 | 8.2 | 2.215 | 2.157 |  | 0.852 | Eric Salvail | 2023-09-25
D | D334 | Viscosity at 40°C - ASTM D445 | cSt | 2.726 | 0.69 | 2.7197 | 0.0092 | 21 | 2.707 | 2.739 | 0.026 | 0.061 | 0.176 - Pass | 0.461 | Eric Salvail | 2023-09-25
D | D334 | Water Content - ASTM D6304, Procedure A | mg/kg | 36 | 0.48 | 32.8 | 6.8 | 21 | 20 | 48 | 18.867 | 3.443 | 30.033 - Fail | 0.261 | Eric Salvail | 2023-09-25
D | D334 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.4593 | 0.2339 | 7 | 2.264 | 2.727 | 0.648 | 0.071 |  | 1.131 | Eric Salvail | 2023-09-25
D | D334 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.7203 | 0.0167 | 12 | 2.682 | 2.746 | 0.046 | 0.073 |  | 0.698 | Eric Salvail | 2023-09-25
D | D334 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 667.8 | 372.8 | 5 | 0.844 | 834.8 | 1032.78 | -4667.452 |  | 1.498 | Eric Salvail | 2023-09-25
D | D334 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 7.777 | 0.0413 | 3 | 7.739 | 7.821 | 0.114 |  |  | 0.311 | Eric Salvail | 2023-09-25
D | D334 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 0.067 | 0.115 | 3 | 0 | 0.2 | 0.32 | 0.714 |  | 0.732 | Eric Salvail | 2023-09-25
D | D335 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 120.8 | 1.13 | 118.94 | 1.64 | 9 | 115.8 | 120.8 | 4.531 | 7.85 |  | 0.307 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 188.8 | -0.25 | 189.13 | 1.34 | 9 | 186.7 | 191 | 3.701 | 4.337 |  | 0.388 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 207.4 | -0.37 | 207.93 | 1.43 | 9 | 205.6 | 210.1 | 3.949 | 4.619 |  | 0.395 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 228.6 | -0.49 | 229.36 | 1.54 | 9 | 226.9 | 232.1 | 4.276 | 4.94 |  | 0.199 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 246.7 | -0.42 | 247.38 | 1.6 | 9 | 244.8 | 250.3 | 4.44 | 4.516 |  | 0.226 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 262.6 | -0.55 | 263.48 | 1.6 | 9 | 261 | 266.4 | 4.434 | 4.3 |  | 0.39 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 278.6 | -0.35 | 279.21 | 1.73 | 9 | 276.7 | 282.5 | 4.805 | 4.3 |  | 0.303 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 294.6 | -0.44 | 295.33 | 1.67 | 9 | 293 | 298.4 | 4.625 | 4.3 |  | 0.376 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 311.2 | -0.52 | 312.13 | 1.78 | 9 | 309.8 | 315.5 | 4.936 | 4.3 |  | 0.282 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 332.3 | -0.41 | 333.11 | 1.98 | 9 | 330.1 | 336.6 | 5.496 | 4.3 |  | 0.202 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 355.1 | -0.21 | 355.38 | 1.31 | 9 | 353.1 | 357.5 | 3.63 | 4.3 |  | 0.167 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 375.9 | -0.61 | 377.38 | 2.42 | 9 | 372.9 | 380.7 | 6.701 | 5 |  | 0.195 | Eric Salvail | 2023-11-10
D | D335 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 437.3 | -0.6 | 443.19 | 9.85 | 9 | 421.3 | 455.5 | 27.28 | 11.8 |  | 0.602 | Eric Salvail | 2023-11-10
D | D335 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 182.4 | 0.55 | 179.89 | 4.56 | 22 | 170.3 | 187 | 12.622 | 9.894 | 1.628 - Pass | 0.282 | Eric Salvail | 2023-11-10
D | D335 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 228.9 | 1.07 | 227.71 | 1.12 | 22 | 225.1 | 229.3 | 3.09 | 5.01 | 0.381 - Pass | 0.314 | Eric Salvail | 2023-11-10
D | D335 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 276 | 1.06 | 275.33 | 0.63 | 22 | 274.2 | 276.7 | 1.742 | 3 | 0.337 - Pass | 0.211 | Eric Salvail | 2023-11-10
D | D335 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 336.7 | 0.09 | 336.56 | 1.58 | 22 | 332 | 339.1 | 4.369 | 5.048 | 0.749 - Pass | 0.469 | Eric Salvail | 2023-11-10
D | D335 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 370.9 | 0.98 | 366.99 | 3.98 | 22 | 356.6 | 375.7 | 11.021 | 7.1 | 2.409 - Fail | 0.459 | Eric Salvail | 2023-11-10
D | D335 | Recovery - ASTM D86 (Automatic) | Vol. % | 97.8 | -0.12 | 97.85 | 0.37 | 22 | 96.9 | 98.4 | 1.014 |  |  | 1.005 | Eric Salvail | 2023-11-10
D | D335 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0.07 | 1.39 | 0.13 | 22 | 1 | 1.7 | 0.372 |  |  | 3.368 | Eric Salvail | 2023-11-10
D | D335 | Loss - ASTM D86 (Automatic) | Vol. % | 0.8 | 0.09 | 0.76 | 0.39 | 22 | 0.2 | 1.7 | 1.083 |  |  | 1.381 | Eric Salvail | 2023-11-10
D | D335 | Aniline Point - ASTM D611 | °C | 54.15 | 0.39 | 53.969 | 0.464 | 8 | 53.45 | 54.8 | 1.284 | 0.5 |  | 0.249 | Eric Salvail | 2023-11-10
D | D335 | Cetane Index, Calculated - ASTM D976 |  |  |  | 43.66 | 0.1 | 13 | 43.5 | 43.9 | 0.289 |  |  | 1.06 | Eric Salvail | 2023-11-10
D | D335 | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 42.37 | 0.25 | 9 | 41.9 | 42.7 | 0.693 |  |  | 0.258 | Eric Salvail | 2023-11-10
D | D335 | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 43.3 | 0.19 | 15 | 42.9 | 43.6 | 0.523 |  |  | 0.378 | Eric Salvail | 2023-11-10
D | D335 | Cetane Number - ASTM D613 |  |  |  | 44.99 | 0.95 | 7 | 43.3 | 46.2 | 2.644 | 3.423 |  | 0.22 | Eric Salvail | 2023-11-10
D | D335 | Cloud Point - ASTM D2500 | °C |  |  | -16.5 | 1.6 | 8 | -20 | -15 | 4.488 | 4 |  | 0.646 | Eric Salvail | 2023-11-10
D | D335 | Cloud Point - ASTM D5773 | °C | -16.2 | 0.53 | -16.79 | 1.09 | 20 | -18.7 | -14.9 | 3.029 | 2.5 | 1.468 - Pass | 0.331 | Eric Salvail | 2023-11-10
D | D335 | Density, at 15°C - ASTM D4052 | kg/m³ | 867.1 | 0.09 | 867.09 | 0.14 | 24 | 866.8 | 867.5 | 0.377 | 0.5 | 0.569 - Pass | 1.278 | Eric Salvail | 2023-11-10
D | D335 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 350 | 0.34 | 342.2 | 23.2 | 19 | 305 | 396 | 64.22 | 39.832 | 2.599 - Fail | 0.27 | Eric Salvail | 2023-11-10
D | D335 | Flash Point - ASTM D93 | °C | 70.3 | 0.97 | 68.87 | 1.48 | 21 | 65.4 | 72 | 4.099 | 4.89 | 0.703 - Pass | 0.242 | Eric Salvail | 2023-11-10
D | D335 | HFRR, Major Axis - ASTM D6079 | µm | 325 | 0.72 | 296.4 | 39.8 | 8 | 240 | 346 | 110.274 |  |  | 0.436 | Eric Salvail | 2023-11-10
D | D335 | HFRR, Minor Axis - ASTM D6079 | µm | 247 | -0.48 | 270.6 | 49.7 | 8 | 220 | 370 | 137.769 |  |  | 0.449 | Eric Salvail | 2023-11-10
D | D335 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 286 | 0.35 | 273.4 | 36.4 | 8 | 220 | 330 | 100.91 | 80 |  | 0.165 | Eric Salvail | 2023-11-10
D | D335 | Pour Point - ASTM D97 | °C |  |  | -26.1 | 2.9 | 7 | -30 | -24 | 7.904 | 9 |  | 1.004 | Eric Salvail | 2023-11-10
D | D335 | Pour Point - ASTM D5949 | °C | -27 | -0.09 | -26.8 | 2.9 | 20 | -31 | -19 | 8.132 | 6.3 | 1.666 - Pass | 0.954 | Eric Salvail | 2023-11-10
D | D335 | Sulfur - ASTM D5453 | mg/kg | 9.2 | 0.31 | 9.02 | 0.6 | 20 | 8.3 | 10.7 | 1.65 | 3.016 | 0.299 - Pass | 0.627 | Eric Salvail | 2023-11-10
D | D335 | Sulfur - ASTM D7039 | mg/kg |  |  | 8.17 | 0.47 | 9 | 7.4 | 8.8 | 1.304 | 2.295 |  | 0.407 | Eric Salvail | 2023-11-10
D | D335 | Viscosity at 40°C - ASTM D445 | cSt | 3.32 | 0.63 | 3.3125 | 0.0118 | 13 | 3.286 | 3.335 | 0.033 | 0.074 | 0.195 - Pass | 0.419 | Eric Salvail | 2023-11-10
D | D335 | Water Content - ASTM D6304, Procedure A | mg/kg | 39 | -1.05 | 43.4 | 4.1 | 14 | 35 | 52 | 11.487 | 4.073 | 7.952 - Fail | 0.41 | Eric Salvail | 2023-11-10
D | D335 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.8137 | 0.0248 | 3 | 2.787 | 2.836 | 0.069 | 0.081 |  | 0.322 | Eric Salvail | 2023-11-10
D | D335 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.3135 | 0.0182 | 11 | 3.282 | 3.338 | 0.05 | 0.086 |  | 0.394 | Eric Salvail | 2023-11-10
D | D335 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 849.4 | 0.1 | 4 | 849.2 | 849.5 | 0.358 | -5938.24 |  | 0.212 | Eric Salvail | 2023-11-10
D | D335 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 10.5467 | 0.0586 | 3 | 10.48 | 10.59 | 0.162 |  |  | 0.424 | Eric Salvail | 2023-11-10
D | D335 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 4.77 | 0.207 | 3 | 4.55 | 4.96 | 0.572 | 0.938 |  | 0.311 | Eric Salvail | 2023-11-10
D | D336 | Density, at 15°C - ASTM D4052 | kg/m³ | 849.4 | 0.53846 | 849.33 | 0.13 | 18 | 849 | 849.5 | 0.368 | 0.5 | 0.542 - Pass | 0.909 | Eric Salvail | 2023-11-24 |  | erreur typo rapport à Innotech, correction post rapport
D | D336 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 470 | -5.08 | 580.2 | 21.7 | 12 | 532 | 611 | 60.097 | 59.559 |  | 0.263 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 101.7 | -0.12 | 101.86 | 1.33 | 9 | 100 | 103.8 | 3.675 | 6.722 |  | 0.239 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 167.1 | -0.23 | 167.32 | 0.98 | 9 | 165.7 | 169.4 | 2.713 | 4.01 |  | 0.709 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 188 | -0.05 | 188.06 | 1.13 | 9 | 185.9 | 190 | 3.134 | 4.321 |  | 0.511 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 211.1 | -0.05 | 211.17 | 1.44 | 9 | 208 | 213.6 | 3.997 | 4.668 |  | 0.978 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 229.7 | -0.23 | 230.23 | 2.34 | 10 | 226 | 235.2 | 6.494 | 4.293 |  | 0.613 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 248.1 | 0.14 | 247.86 | 1.73 | 10 | 243.8 | 250.7 | 4.78 | 4.3 |  | 1.156 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 265.2 | 0.32 | 264.48 | 2.25 | 10 | 260.3 | 267.8 | 6.22 | 4.3 |  | 0.991 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 282.5 | 0.33 | 281.43 | 3.23 | 10 | 274.3 | 285.5 | 8.934 | 4.3 |  | 1.126 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 300.2 | -0.03 | 300.27 | 2.15 | 9 | 295.2 | 303.1 | 5.964 | 4.3 |  | 0.81 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 318.4 | 0 | 318.4 | 1.97 | 9 | 314 | 321 | 5.455 | 4.3 |  | 0.492 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 343.5 | 0.17 | 343.12 | 2.16 | 9 | 338.3 | 345.5 | 5.989 | 4.3 |  | 0.584 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 361.3 | 0.21 | 360.88 | 2.05 | 9 | 356.5 | 363.3 | 5.68 | 5 |  | 0.425 | Eric Salvail | 2023-11-24
D | D336 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 402.1 | 1.15 | 398.41 | 3.2 | 9 | 394 | 402.1 | 8.876 | 11.8 |  | 0.76 | Eric Salvail | 2023-11-24
D | D336 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 163 | 1.29 | 159.17 | 2.96 | 18 | 153.9 | 163 | 8.191 | 8.754 | 0.875 - Pass | 0.627 | Eric Salvail | 2023-11-24
D | D336 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 208.7 | 1.49 | 207.16 | 1.04 | 18 | 205 | 208.7 | 2.881 | 4.557 | 0.400 - Pass | 0.681 | Eric Salvail | 2023-11-24
D | D336 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 261.6 | 0.59 | 261.22 | 0.64 | 18 | 259.8 | 262.3 | 1.776 | 3 | 0.351 - Pass | 0.219 | Eric Salvail | 2023-11-24
D | D336 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 324.1 | -0.37 | 324.35 | 0.68 | 18 | 323.3 | 325.8 | 1.896 | 4.865 | 0.152 - Pass | 0.378 | Eric Salvail | 2023-11-24
D | D336 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 351.7 | 1.84 | 349.57 | 1.16 | 18 | 346.5 | 351.7 | 3.209 | 7.1 | 0.204 - Pass | 0.581 | Eric Salvail | 2023-11-24
D | D336 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.5 | 2.17 | 98.13 | 0.17 | 16 | 97.9 | 98.5 | 0.471 |  |  | 0.385 | Eric Salvail | 2023-11-24
D | D336 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0 | 1.4 | 0 | 15 | 1.4 | 1.4 | 0 |  |  |  | Eric Salvail | 2023-11-24
D | D336 | Loss - ASTM D86 (Automatic) | Vol. % | 0.1 | -1.31 | 0.49 | 0.3 | 18 | 0 | 1.3 | 0.836 |  |  | 0.4 | Eric Salvail | 2023-11-24
D | D336 | Aniline Point - ASTM D611 | °C | 63.4 | -0.82 | 63.51 | 0.134 | 5 | 63.4 | 63.7 | 0.372 | 0.5 |  | 0.505 | Eric Salvail | 2023-11-24
D | D336 | Cetane Index, Calculated - ASTM D976 |  |  |  | 46.06 | 0.2 | 10 | 45.8 | 46.5 | 0.542 |  |  | 0.463 | Eric Salvail | 2023-11-24
D | D336 | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 44.83 | 0.45 | 7 | 44.4 | 45.7 | 1.236 |  |  | 0.504 | Eric Salvail | 2023-11-24
D | D336 | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 45.54 | 0.18 | 10 | 45.3 | 45.9 | 0.509 |  |  | 0.396 | Eric Salvail | 2023-11-24
D | D336 | Cetane Number - ASTM D613 |  |  |  | 42.88 | 1.12 | 4 | 41.4 | 43.8 | 3.112 | 3.159 |  | 0.416 | Eric Salvail | 2023-11-24
D | D336 | Cloud Point - ASTM D2500 | °C |  |  | -23.8 | 0.8 | 6 | -25 | -23 | 2.085 | 4 |  | 0.567 | Eric Salvail | 2023-11-24
D | D336 | Cloud Point - ASTM D5773 | °C | -24.3 | 0.17 | -24.49 | 1.09 | 16 | -26.2 | -22.3 | 3.014 | 2.5 | 1.453 - Pass | 0.241 | Eric Salvail | 2023-11-24
D | D336 | Flash Point - ASTM D93 | °C | 55.2 | -1.23 | 56.76 | 1.27 | 18 | 55 | 59 | 3.504 | 4.03 | 0.756 - Pass | 0.36 | Eric Salvail | 2023-11-24
D | D336 | HFRR, Major Axis - ASTM D6079 | µm |  |  | 479.3 | 31.8 | 7 | 431 | 520 | 87.984 |  |  | 0.468 | Eric Salvail | 2023-11-24
D | D336 | HFRR, Minor Axis - ASTM D6079 | µm |  |  | 408 | 15.9 | 7 | 390 | 430 | 44.089 |  |  | 0.372 | Eric Salvail | 2023-11-24
D | D336 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 459 | 0.53 | 447.4 | 21.8 | 8 | 414 | 480 | 60.306 | 80 |  | 0.166 | Eric Salvail | 2023-11-24
D | D336 | Pour Point - ASTM D97 | °C |  |  | -38.5 | 2.3 | 6 | -42 | -36 | 6.256 | 9 |  | 0.567 | Eric Salvail | 2023-11-24
D | D336 | Pour Point - ASTM D5949 | °C | -34 | 1.1 | -38.4 | 4 | 15 | -49 | -32 | 11.12 | 6.3 | 3.115 - Fail | 0.692 | Eric Salvail | 2023-11-24
D | D336 | Sulfur - ASTM D5453 | mg/kg | 5.2 | 0.03 | 5.19 | 0.38 | 19 | 4.6 | 5.8 | 1.052 | 1.993 | 0.279 - Pass | 0.261 | Eric Salvail | 2023-11-24
D | D336 | Sulfur - ASTM D7039 | mg/kg |  |  | 4.88 | 0.69 | 5 | 3.8 | 5.4 | 1.903 | 1.738 |  | 0.546 | Eric Salvail | 2023-11-24
D | D336 | Viscosity at 40°C - ASTM D445 | cSt | 2.724 | 0.48 | 2.7192 | 0.0099 | 13 | 2.7 | 2.737 | 0.027 | 0.061 | 0.203 - Pass | 0.233 | Eric Salvail | 2023-11-24
D | D336 | Water Content - ASTM D6304, Procedure A | mg/kg | 25.6 | -0.23 | 26.9 | 5.7 | 12 | 14 | 36 | 15.665 | 3.057 |  | 0.597 | Eric Salvail | 2023-11-24
D | D336 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.2647 | 0.0103 | 6 | 2.253 | 2.279 | 0.029 | 0.065 |  | 0.282 | Eric Salvail | 2023-11-24
D | D336 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.7249 | 0.0113 | 8 | 2.708 | 2.742 | 0.031 | 0.074 |  | 0.334 | Eric Salvail | 2023-11-24
D | D336 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 831.6 | 0.2 | 6 | 831.3 | 831.9 | 0.572 | -5813.757 |  | 0.45 | Eric Salvail | 2023-11-24
D | D336 | Density at 15°C - ASTM D7042 | kg/m³ |  |  | 848.6 | 1.4 | 3 | 847 | 849.7 | 3.979 | 1.77 |  | 0.483 | Eric Salvail | 2023-11-24
D | D337 | Residue - ASTM D86 (Automatic) | Vol. % | 0.6 | -3.23 | 1.28 | 0.21 | 30 | 0.6 | 1.5 | 0.587 |  |  | 3.296 | Martine Carrier | 2023-12-21 |  | recheck demandé
D | D337 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 100.9 | 0.25 | 100.28 | 2.5 | 11 | 96.8 | 104.5 | 6.914 | 6.619 |  | 0.177 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 151.8 | 0.69 | 151.05 | 1.09 | 11 | 148.7 | 152.4 | 3.018 | 3.766 |  | 0.285 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 166.2 | 0.19 | 165.99 | 1.08 | 11 | 163.7 | 167.6 | 2.979 | 3.99 |  | 0.382 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 188.3 | 0.07 | 188.2 | 1.38 | 11 | 185.8 | 190.1 | 3.822 | 4.323 |  | 0.378 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 208.2 | 0.17 | 207.94 | 1.51 | 11 | 205 | 210 | 4.178 | 4.003 |  | 0.333 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 225.9 | 0.07 | 225.77 | 1.8 | 11 | 222 | 228.2 | 4.989 | 4.3 |  | 0.344 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 242.4 | 0.09 | 242.25 | 1.76 | 11 | 239 | 244.8 | 4.886 | 4.3 |  | 0.242 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 259.3 | 0.08 | 259.15 | 1.78 | 11 | 256 | 261.8 | 4.934 | 4.3 |  | 0.246 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 277.8 | -0.03 | 277.85 | 1.74 | 11 | 275 | 280.6 | 4.832 | 4.3 |  | 0.175 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 300.1 | -0.09 | 300.25 | 1.69 | 11 | 298 | 302.8 | 4.688 | 4.3 |  | 0.326 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 329.3 | -0.05 | 329.38 | 1.8 | 11 | 325.9 | 331.2 | 4.998 | 4.3 |  | 0.9 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 352.2 | 0.06 | 352.09 | 1.8 | 11 | 348.1 | 354.1 | 4.987 | 5 |  | 0.597 | Martine Carrier | 2023-12-21
D | D337 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 391.8 | -0.02 | 391.88 | 4.58 | 11 | 382.7 | 401 | 12.691 | 11.8 |  | 0.395 | Martine Carrier | 2023-12-21
D | D337 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 155.3 | 2 | 151.49 | 1.9 | 30 | 148.1 | 155.3 | 5.268 | 8.332 | 0.400 - Pass | 0.285 | Martine Carrier | 2023-12-21
D | D337 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 185.5 | 1.23 | 184.39 | 0.9 | 30 | 182.7 | 186.7 | 2.502 | 4.057 | 0.380 - Pass | 0.226 | Martine Carrier | 2023-12-21
D | D337 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 240.5 | 0.92 | 239.91 | 0.64 | 30 | 238.8 | 241.6 | 1.785 | 3 | 0.354 - Pass | 0.276 | Martine Carrier | 2023-12-21
D | D337 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 312.2 | -0.29 | 312.58 | 1.3 | 29 | 310.7 | 315.4 | 3.6 | 4.689 | 0.590 - Pass | 0.379 | Martine Carrier | 2023-12-21
D | D337 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 346.3 | 0.79 | 344.96 | 1.7 | 30 | 341.2 | 349.8 | 4.7 | 7.1 | 0.438 - Pass | 0.683 | Martine Carrier | 2023-12-21
D | D337 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.1 | -0.07 | 98.12 | 0.28 | 30 | 97.4 | 98.6 | 0.775 |  |  | 0.363 | Martine Carrier | 2023-12-21
D | D337 | Loss - ASTM D86 (Automatic) | Vol. % | 0.5 | -0.22 | 0.57 | 0.32 | 30 | 0 | 1.3 | 0.888 |  |  | 0.651 | Martine Carrier | 2023-12-21
D | D337 | Aniline Point - ASTM D611 | °C | 52.65 | 1.47 | 51.917 | 0.5 | 9 | 51 | 52.65 | 1.385 | 0.5 |  | 0.497 | Martine Carrier | 2023-12-21
D | D337 | Cetane Index, Calculated - ASTM D976 |  | 41.8 | 0.91 | 41.65 | 0.17 | 20 | 41.3 | 42 | 0.471 |  |  | 0.337 | Martine Carrier | 2023-12-21
D | D337 | Cetane Index, Calculated - ASTM D4737 B |  | 40.3 | 0.44 | 40.19 | 0.26 | 14 | 39.7 | 40.5 | 0.72 |  |  | 0.337 | Martine Carrier | 2023-12-21
D | D337 | Cetane Index, Calculated - ASTM D4737 A |  | 41.7 | 0.29 | 41.64 | 0.2 | 19 | 41.2 | 42 | 0.549 |  |  | 0.303 | Martine Carrier | 2023-12-21
D | D337 | Cetane Number - ASTM D613 |  |  |  | 41.79 | 0.95 | 9 | 40.8 | 43.1 | 2.629 | 3.024 |  | 0.804 | Martine Carrier | 2023-12-21
D | D337 | Cloud Point - ASTM D2500 | °C | -39 | 0.65 | -40 | 1.5 | 11 | -43 | -37 | 4.291 | 4 |  | 0.417 | Martine Carrier | 2023-12-21
D | D337 | Cloud Point - ASTM D5773 | °C |  |  | -40.13 | 1.21 | 26 | -41.8 | -36.9 | 3.357 | 2.5 | 1.803 - Pass | 0.545 | Martine Carrier | 2023-12-21
D | D337 | Density, at 15°C - ASTM D4052 | kg/m³ | 846.1 | 0.18 | 846.09 | 0.07 | 32 | 845.9 | 846.2 | 0.196 | 0.5 | 0.153 - Pass | 3.017 | Martine Carrier | 2023-12-21
D | D337 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 1266 | 0.09 | 1258.4 | 82.1 | 25 | 1120 | 1475 | 227.323 | 115.781 | 3.855 - Fail | 0.502 | Martine Carrier | 2023-12-21
D | D337 | Flash Point - ASTM D93 | °C | 50 | -0.67 | 50.75 | 1.11 | 30 | 49 | 53.5 | 3.083 | 3.603 | 0.732 - Pass | 0.485 | Martine Carrier | 2023-12-21
D | D337 | HFRR, Major Axis - ASTM D6079 | µm | 670 | 1.09 | 637.1 | 30.1 | 12 | 570 | 670 | 83.305 |  |  | 0.515 | Martine Carrier | 2023-12-21
D | D337 | HFRR, Minor Axis - ASTM D6079 | µm | 643 | 1.43 | 604.5 | 26.8 | 12 | 550 | 643 | 74.332 |  |  | 0.508 | Martine Carrier | 2023-12-21
D | D337 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 656 | 1.24 | 622.1 | 27.3 | 12 | 560 | 656 | 75.624 | 80 |  | 0.502 | Martine Carrier | 2023-12-21
D | D337 | Pour Point - ASTM D97 | °C |  |  | -48.3 | 1.7 | 10 | -51 | -45 | 4.717 | 9 |  | 1.561 | Martine Carrier | 2023-12-21
D | D337 | Pour Point - ASTM D5949 | °C | -51 | -0.32 | -50.2 | 2.5 | 25 | -57 | -46 | 6.925 | 6.3 | 1.208 - Pass | 1.411 | Martine Carrier | 2023-12-21
D | D337 | Sulfur - ASTM D7039 | mg/kg |  |  | 5.98 | 0.57 | 11 | 4.9 | 7 | 1.571 | 1.94 |  | 0.283 | Martine Carrier | 2023-12-21
D | D337 | Viscosity at 40°C - ASTM D445 | cSt | 1.979 | 0.29 | 1.9749 | 0.0141 | 19 | 1.944 | 2.008 | 0.039 | 0.044 | 0.774 - Pass | 0.817 | Martine Carrier | 2023-12-21
D | D337 | Water Content - ASTM D6304, Procedure A | mg/kg | 27 | 0.75 | 24.1 | 3.9 | 21 | 13 | 33 | 10.691 | 2.865 | 13.928 - Fail | 0.752 | Martine Carrier | 2023-12-21
D | D337 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s | 1.97 | 1.71 | 1.7066 | 0.1539 | 9 | 1.614 | 1.985 | 0.426 | 0.049 |  | 1.992 | Martine Carrier | 2023-12-21
D | D337 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 1.9693 | 0.0135 | 12 | 1.949 | 1.989 | 0.037 | 0.057 |  | 0.334 | Martine Carrier | 2023-12-21
D | D337 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 827.9 | 0.3 | 7 | 827.5 | 828.4 | 0.79 | -5788.19 |  | 0.342 | Martine Carrier | 2023-12-21
D | D337 | Density at 15°C - ASTM D7042 | kg/m³ | 846 | 0.46 | 845.4 | 1.3 | 5 | 843 | 846.2 | 3.718 | 1.77 |  | 1.083 | Martine Carrier | 2023-12-21
D | D337 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 4.807 | 0.0303 | 3 | 4.772 | 4.826 | 0.084 |  |  | 0.653 | Martine Carrier | 2023-12-21
D | D337 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 0.76 | 0.393 | 3 | 0.33 | 1.1 | 1.088 | 0.747 |  | 0.35 | Martine Carrier | 2023-12-21
D | D337 | Sulfur - ASTM D5453 | mg/kg | 2.4 | -7.4 | 6.7 | 0.58 | 26 | 5.6 | 8.3 | 1.61 | 2.415 | 0.445 - Pass | 0.263 | Martine Carrier | 2023-12-21 |  | recheck demandé
D | D338 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | *221.8 | *2.46 | 213.25 | 3.47 | 11 | 209 | 221.8 | 9.615 | 4.699 |  | 0.575 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | *235.2 | *1.39 | 231.61 | 2.58 | 11 | 227 | 235.2 | 7.152 | 4.311 |  | 0.281 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | *246.4 | *-0.43 | 247.43 | 2.39 | 11 | 243 | 251.6 | 6.616 | 4.3 |  | 0.259 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | *257.4 | *-1.45 | 261.65 | 2.93 | 11 | 257 | 266.1 | 8.118 | 4.3 |  | 0.197 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | *275.1 | *-0.47 | 276.37 | 2.73 | 11 | 271 | 280.8 | 7.55 | 4.3 |  | 0.241 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | *280 | *-2.46 | 290.85 | 4.4 | 11 | 280 | 296.3 | 12.193 | 4.3 |  | 0.625 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | *349.2 | *-2.46 | 401.7 | 21.32 | 11 | 349.2 | 432.5 | 59.046 | 11.8 |  | 0.947 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | *170.1 | *1 | 167.34 | 2.77 | 31 | 162.9 | 173.7 | 7.661 | 9.204 | 0.693 - Pass | 0.481 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | *206.2 | *1.17 | 204.96 | 1.06 | 32 | 203 | 207.2 | 2.946 | 4.509 | 0.427 - Pass | 0.772 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | *260.7 | *1.74 | 259.61 | 0.63 | 32 | 258.6 | 260.9 | 1.738 | 3 | 0.335 - Pass | 0.64 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | *316.5 | *0.03 | 316.47 | 1 | 32 | 314.9 | 319.6 | 2.774 | 4.747 | 0.341 - Pass | 0.45 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | *341.8 | *1.13 | 339.98 | 1.61 | 32 | 336.8 | 343.6 | 4.452 | 7.1 | 0.393 - Pass | 0.349 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Recovery - ASTM D86 (Automatic) | Vol. % | *98.2 | *0.22 | 98.12 | 0.36 | 32 | 97.4 | 98.8 | 0.997 |  |  | 0.435 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | *-0.39 | 1.42 | 0.05 | 26 | 1.4 | 1.6 | 0.136 |  |  | 5.807 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Loss - ASTM D86 (Automatic) | Vol. % | 0.4 | *-0.33 | 0.52 | 0.35 | 32 | 0 | 1.5 | 0.975 |  |  | 1.066 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Aniline Point - ASTM D611 | °C | 57.65 | *0.58 | 57.364 | 0.494 | 10 | 56.8 | 58.4 | 1.368 | 0.5 |  | 0.427 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cetane Index, Calculated - ASTM D976 |  |  |  | 44.2 | 0.12 | 21 | 44 | 44.4 | 0.333 |  |  | 0.682 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 41.93 | 0.17 | 12 | 41.6 | 42.2 | 0.459 |  |  | 0.3 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 43.2 | 0.15 | 19 | 42.8 | 43.4 | 0.402 |  |  | 0.753 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cetane Number - ASTM D613 |  |  |  | 46.78 | 1.89 | 9 | 43.8 | 49.6 | 5.244 | 3.647 |  | 0.355 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cloud Point - ASTM D2500 | °C |  |  | -15 | 18.8 | 8 | -34 | 10 | 52.117 | 4 |  | 0.59 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Cloud Point - ASTM D5773 | °C | 3.3 | *0.52 | -4.43 | 14.79 | 22 | -35.1 | 17.2 | 40.962 | 2.5 | 268.459 - Fail | 0.634 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Density, at 15°C - ASTM D4052 | kg/m³ | 854 | *0.46 | 853.94 | 0.14 | 34 | 853.6 | 854.4 | 0.391 | 0.5 | 0.612 - Pass | 2.187 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 107 | *-0.83 | 127.2 | 24.2 | 30 | 80 | 169 | 67.132 | 22.011 | 9.302 - Fail | 0.618 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Flash Point - ASTM D93 | °C | 61.7 | *-1.48 | 63.38 | 1.13 | 33 | 61.5 | 66 | 3.13 | 4.5 | 0.484 - Pass | 0.698 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | HFRR, Major Axis - ASTM D6079 | µm |  |  | 386.4 | 43.9 | 11 | 316 | 452 | 121.548 |  |  | 0.273 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | HFRR, Minor Axis - ASTM D6079 | µm |  |  | 298.4 | 62.5 | 11 | 229 | 457 | 173.085 |  |  | 0.732 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 225 | *-2.13 | 326 | 47.4 | 12 | 225 | 390 | 131.292 | 80 |  | 0.298 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Pour Point - ASTM D97 | °C |  |  | -52.5 | 4.2 | 8 | -60 | -48 | 11.752 | 9 |  | 0.461 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Pour Point - ASTM D5949 | °C | -52 | *0.35 | -53.3 | 3.7 | 24 | -63 | -44 | 10.185 | 6.3 | 2.614 - Fail | 1.095 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Sulfur - ASTM D5453 | mg/kg | 1.6 | *-0.04 | 1.62 | 0.38 | 32 | 1 | 2.7 | 1.051 | 0.831 | 1.600 - Pass | 0.645 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Sulfur - ASTM D7039 | mg/kg |  |  | 1.42 | 0.25 | 9 | 1 | 1.8 | 0.703 | 0.893 |  | 0.156 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Viscosity at 40°C - ASTM D445 | cSt | 2.619 | *0.25 | 2.6126 | 0.0252 | 23 | 2.55 | 2.675 | 0.07 | 0.059 | 1.428 - Pass | 2.018 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Water Content - ASTM D6304, Procedure A | mg/kg |  |  | 74.7 | 29.5 | 21 | 28 | 142 | 81.7 | 5.644 | 209.564 - Fail | 0.228 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 2.1843 | 0.0133 | 7 | 2.159 | 2.197 | 0.037 | 0.063 |  | 0.525 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s | 2.6259 | *1.02 | 2.6107 | 0.015 | 15 | 2.582 | 2.627 | 0.042 | 0.071 | 0.341 - Pass | 0.731 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 836.1 | 0.2 | 7 | 835.7 | 836.3 | 0.541 | -5845.39 |  | 0.543 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Density at 15°C - ASTM D7042 | kg/m³ | 854 | *0.83 | 852.9 | 1.3 | 5 | 851 | 854 | 3.685 | 1.77 |  | 0.583 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  | 7.274 | 0.0374 | 4 | 7.218 | 7.296 | 0.104 |  |  | 0.923 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Biodiesel Content - ASTM D7371 | Volume % |  |  | 4.927 | 0.441 | 10 | 4.27 | 5.54 | 1.221 | 0.946 |  | 0.357 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 205.8 | *9.86 | 184.38 | 2.17 | 10 | 181 | 188.2 | 6.016 | 4.266 |  | 0.19 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 195 | *19.79 | 163.26 | 1.6 | 10 | 161 | 165.9 | 4.442 | 3.949 |  | 0.225 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 294.2 | *-5.29 | 309.42 | 2.88 | 10 | 304 | 314.7 | 7.976 | 4.3 |  | 0.266 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 314.9 | *-5.59 | 337.53 | 4.05 | 10 | 332 | 347.2 | 11.215 | 4.3 |  | 0.623 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 331.8 | *-19.3 | 354.6 | 1.18 | 9 | 353 | 356.5 | 3.272 | 5 |  | 0.315 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
D | D338 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 171.8 | *35.35 | 120.2 | 1.46 | 10 | 117.9 | 123.1 | 4.044 | 7.933 |  | 0.266 | Eric Salvail | 2024-03-13 |  | échantillon non valide, reprise D338r
F | F305 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 102 | -0.65 | 103.59 | 2.45 | 8 | 100 | 107.2 | 6.778 | 6.837 |  | 0.43 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 160.6 | -0.42 | 161.28 | 1.6 | 8 | 159 | 162.9 | 4.431 | 3.919 |  | 0.485 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 181.2 | 0.02 | 181.16 | 2.36 | 8 | 177 | 183.4 | 6.547 | 4.217 |  | 0.737 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 212 | -0.03 | 212.08 | 2.53 | 8 | 208 | 214.8 | 7.003 | 4.681 |  | 0.604 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 232.8 | -0.18 | 233.24 | 2.39 | 8 | 229 | 235.8 | 6.607 | 4.332 |  | 0.426 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 251.2 | -0.15 | 251.55 | 2.29 | 8 | 247 | 254.4 | 6.355 | 4.3 |  | 0.307 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 268.2 | -0.14 | 268.51 | 2.17 | 8 | 264 | 271.3 | 6.01 | 4.3 |  | 0.449 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 285.6 | -0.15 | 285.93 | 2.22 | 8 | 281 | 288.4 | 6.147 | 4.3 |  | 0.693 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 303.2 | -0.29 | 303.71 | 1.75 | 8 | 300 | 305.9 | 4.86 | 4.3 |  | 0.519 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 323.8 | -0.32 | 324.64 | 2.6 | 8 | 319 | 327.8 | 7.197 | 4.3 |  | 0.602 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 352.2 | -0.41 | 352.96 | 1.87 | 8 | 349 | 355 | 5.186 | 4.3 |  | 0.483 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 367 | -0.54 | 368.11 | 2.04 | 8 | 364 | 370.5 | 5.656 | 5 |  | 0.402 | Sarah Eid | 02-Aug-2022
F | F305 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 422.4 | -1.12 | 424.31 | 1.71 | 7 | 422.4 | 427.5 | 4.735 | 11.8 |  | 0.332 | Sarah Eid | 02-Aug-2022
F | F305 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 156.3 | -0.06 | 156.45 | 2.66 | 14 | 151.7 | 162.3 | 7.365 | 8.605 | 0.733 - Pass | 0.185 | Sarah Eid | 02-Aug-2022
F | F305 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 204.7 | 1.73 | 202.18 | 1.46 | 14 | 199 | 204.7 | 4.039 | 4.448 | 0.825 - Pass | 0.215 | Sarah Eid | 02-Aug-2022
F | F305 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 266.7 | 1.41 | 265.71 | 0.7 | 14 | 264.6 | 266.8 | 1.943 | 3 | 0.420 - Pass | 0.337 | Sarah Eid | 02-Aug-2022
F | F305 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 331 | 0.03 | 330.95 | 1.53 | 14 | 328.9 | 334.6 | 4.24 | 4.964 | 0.729 - Pass | 0.28 | Sarah Eid | 02-Aug-2022
F | F305 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 360.7 | 0.34 | 360.05 | 1.88 | 14 | 357 | 363.6 | 5.22 | 7.1 | 0.541 - Pass | 0.347 | Sarah Eid | 02-Aug-2022
F | F305 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.5 | 0.79 | 98.33 | 0.21 | 12 | 98.1 | 98.7 | 0.583 |  |  | 0.657 | Sarah Eid | 02-Aug-2022
F | F305 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0 | 1.4 | 0.08 | 13 | 1.3 | 1.6 | 0.226 |  |  | 1.442 | Sarah Eid | 02-Aug-2022
F | F305 | Loss - ASTM D86 (Automatic) | Vol. % | 0.1 | -0.83 | 0.44 | 0.41 | 14 | 0 | 1.3 | 1.148 |  |  | 1.003 | Sarah Eid | 02-Aug-2022
F | F305 | Acid Number - ASTM D664 | mg KOH/g | 0.017 | 0.06 | 0.016 | 0.011 | 8 | 0 | 0.037 | 0.031 | 0.014 |  | 0.4 | Sarah Eid | 02-Aug-2022
F | F305 | Carbon Residue, Micro Conradson - ASTM D4530 | Mass % | 0.032 | 2.04 | 0.01 | 0.01 | 6 | 0 | 0.032 | 0.036 | 0.141 |  | 1.901 | Sarah Eid | 02-Aug-2022
F | F305 | Cloud Point - ASTM D2500 | °C |  |  | -19 | 0 | 6 | -19 | -19 | 0 | 4 |  |  | Sarah Eid | 02-Aug-2022
F | F305 | Cloud Point - ASTM D5773 | °C | -18.2 | 1.01 | -19.06 | 0.86 | 13 | -20.7 | -17.6 | 2.371 | 2.5 | 0.899 - Pass | 0.253 | Sarah Eid | 02-Aug-2022
F | F305 | Density, at 15°C - ASTM D4052 | kg/m³ | 856.8 | 0 | 856.8 | 0.04 | 15 | 856.7 | 856.9 | 0.105 | 0.5 | 0.044 - Pass | 4 | Sarah Eid | 02-Aug-2022
F | F305 | Flash Point - ASTM D93 | °C | 56.3 | -1.1 | 57.4 | 1 | 15 | 55.5 | 59 | 2.758 | 4.075 | 0.458 - Pass | 0.391 | Sarah Eid | 02-Aug-2022
F | F305 | Nitrogen - ASTM D4629 | ppm wt/wt | 45.5 | -0.01 | 45.59 | 7.14 | 11 | 36 | 63 | 19.79 | 5.785 |  | 0.701 | Sarah Eid | 02-Aug-2022
F | F305 | Pour Point - ASTM D97 | °C |  |  | -29.1 | 4.8 | 7 | -39 | -24 | 13.326 | 9 |  | 0.749 | Sarah Eid | 02-Aug-2022
F | F305 | Pour Point - ASTM D5949 | °C | -29 | -0.61 | -27.2 | 3 | 12 | -32 | -21 | 8.338 | 6.3 |  | 0.215 | Sarah Eid | 02-Aug-2022
F | F305 | Sulfur - ASTM D5453 | mg/kg | 7.07 | 0.9 | 6.66 | 0.45 | 14 | 5.8 | 7.6 | 1.25 | 2.404 | 0.270 - Pass | 0.243 | Sarah Eid | 02-Aug-2022
F | F305 | Sulfur - ASTM D7039 | mg/kg |  |  | 6.14 | 0.34 | 5 | 5.8 | 6.7 | 0.931 | 1.967 |  | 0.582 | Sarah Eid | 02-Aug-2022
F | F305 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00012 | 8e-05 | 5 | 0 | 0.0002 | 0 | 0 |  | 0.452 | Sarah Eid | 02-Aug-2022
F | F305 | Viscosity at 40°C - ASTM D445 | cSt | 2.651 | 1.67 | 2.6435 | 0.0045 | 10 | 2.638 | 2.651 | 0.012 | 0.03 |  | 0.497 | Sarah Eid | 02-Aug-2022
F | F305 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.639 | 0.0173 | 6 | 2.618 | 2.658 | 0.048 | 0.015 |  | 0.374 | Sarah Eid | 02-Aug-2022
F | F306 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 115 | -0.8 | 116.66 | 2.09 | 8 | 114.1 | 120.8 | 5.789 | 7.7 |  | 0.402 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 143.2 | -0.57 | 146.11 | 5.13 | 9 | 140.5 | 156.5 | 14.2 | 3.692 |  | 0.891 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 158 | -0.53 | 161.01 | 5.72 | 9 | 154.9 | 174 | 15.847 | 3.915 |  | 1.024 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 180.6 | -0.5 | 183.99 | 6.77 | 9 | 176.4 | 200 | 18.758 | 4.26 |  | 1.058 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 200.2 | -0.46 | 203.54 | 7.23 | 9 | 194.5 | 220.5 | 20.022 | 3.946 |  | 0.976 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 216.8 | -0.31 | 217.94 | 3.66 | 8 | 211.4 | 224.6 | 10.125 | 4.3 |  | 0.672 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 232.4 | -0.35 | 233.79 | 4.01 | 8 | 226.4 | 240.8 | 11.108 | 4.3 |  | 0.567 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 247.4 | -0.26 | 248.48 | 4.11 | 8 | 241.1 | 255.8 | 11.372 | 4.3 |  | 0.569 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 261.2 | -0.32 | 262.51 | 4.12 | 8 | 255.3 | 270 | 11.422 | 4.3 |  | 0.539 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 275.6 | -0.4 | 277.33 | 4.31 | 8 | 270 | 285.2 | 11.949 | 4.3 |  | 0.477 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 293.4 | -0.45 | 295.21 | 4.01 | 8 | 289.1 | 303.2 | 11.11 | 4.3 |  | 0.503 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 304.6 | -0.43 | 306.4 | 4.14 | 8 | 300.7 | 315.1 | 11.468 | 5 |  | 0.588 | Martine Carrier | 2022-10-17
F | F306 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 330.2 | -0.39 | 334.13 | 9.97 | 8 | 328.2 | 358.5 | 27.627 | 11.8 |  | 1.887 | Martine Carrier | 2022-10-17
F | F306 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 157.9 | 1.88 | 152.39 | 2.93 | 15 | 147.3 | 157.9 | 8.115 | 8.382 | 0.937 - Pass | 0.243 | Martine Carrier | 2022-10-17
F | F306 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 181.5 | 1.78 | 179.85 | 0.93 | 15 | 178.6 | 181.5 | 2.564 | 3.957 | 0.420 - Pass | 0.403 | Martine Carrier | 2022-10-17
F | F306 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 232.8 | 1.24 | 231.99 | 0.65 | 15 | 230.7 | 233 | 1.81 | 3 | 0.364 - Pass | 0.163 | Martine Carrier | 2022-10-17
F | F306 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 278.7 | 0.18 | 278.57 | 0.7 | 15 | 277.4 | 279.9 | 1.932 | 4.179 | 0.214 - Pass | 0.212 | Martine Carrier | 2022-10-17
F | F306 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 301.2 | 2.31 | 297.75 | 1.5 | 15 | 295.4 | 301.2 | 4.15 | 7.1 | 0.342 - Pass | 0.224 | Martine Carrier | 2022-10-17
F | F306 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.8 | 1.3 | 98.35 | 0.34 | 15 | 97.4 | 98.8 | 0.953 |  |  | 1.022 | Martine Carrier | 2022-10-17
F | F306 | Residue - ASTM D86 (Automatic) | Vol. % | 0.9 | -2.25 | 1.27 | 0.16 | 15 | 0.9 | 1.6 | 0.452 |  |  | 1.953 | Martine Carrier | 2022-10-17
F | F306 | Loss - ASTM D86 (Automatic) | Vol. % | 0.3 | -0.12 | 0.32 | 0.18 | 14 | 0.1 | 0.8 | 0.512 |  |  | 1.307 | Martine Carrier | 2022-10-17
F | F306 | Acid Number - ASTM D664 | mg KOH/g | 0.038 | -0.52 | 0.042 | 0.008 | 10 | 0.03 | 0.057 | 0.022 | 0.03 |  | 0.741 | Martine Carrier | 2022-10-17
F | F306 | Cloud Point - ASTM D2500 | °C |  |  | -49.2 | 2.6 | 6 | -54 | -47 | 7.311 | 4 |  | 0.524 | Martine Carrier | 2022-10-17
F | F306 | Cloud Point - ASTM D5773 | °C | -51.6 | -0.8 | -49.49 | 2.64 | 14 | -54.1 | -43.2 | 7.323 | 2.5 | 8.579 - Fail | 0.449 | Martine Carrier | 2022-10-17
F | F306 | Density, at 15°C - ASTM D4052 | kg/m³ | 836.1 | 0.29 | 836.07 | 0.1 | 18 | 835.9 | 836.3 | 0.265 | 0.5 | 0.282 - Pass | 1.064 | Martine Carrier | 2022-10-17
F | F306 | Flash Point - ASTM D93 | °C | 52.2 | -0.64 | 53.01 | 1.27 | 17 | 51 | 55 | 3.523 | 3.764 | 0.876 - Pass | 0.499 | Martine Carrier | 2022-10-17
F | F306 | Nitrogen - ASTM D4629 | ppm wt/wt | 22 | 0.51 | 21 | 1.95 | 12 | 18 | 25 | 5.413 | 3.881 |  | 0.252 | Martine Carrier | 2022-10-17
F | F306 | Pour Point - ASTM D97 | °C |  |  | -65 | 1.7 | 3 | -66 | -63 | 4.798 | 9 |  | 0.732 | Martine Carrier | 2022-10-17
F | F306 | Pour Point - ASTM D5949 | °C | -67 | 0 | -67 | 5.5 | 10 | -76 | -56 | 15.228 | 6.3 |  | 0.305 | Martine Carrier | 2022-10-17
F | F306 | Sulfur - ASTM D5453 | mg/kg | 0.9 | 0.07 | 0.88 | 0.29 | 10 | 0.5 | 1.5 | 0.803 | 0.527 |  | 0.533 | Martine Carrier | 2022-10-17
F | F306 | Sulfur - ASTM D7039 | mg/kg |  |  | 0.93 | 0.54 | 4 | 0.4 | 1.6 | 1.49 | 0.708 |  | 0.289 | Martine Carrier | 2022-10-17
F | F306 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00017 | 0.00014 | 7 | 0 | 0.0004 | 0 | 0 |  | 0.432 | Martine Carrier | 2022-10-17
F | F306 | Viscosity at 40°C - ASTM D445 | cSt | 1.783 | 0.47 | 1.7803 | 0.0058 | 11 | 1.771 | 1.789 | 0.016 | 0.023 |  | 0.202 | Martine Carrier | 2022-10-17
F | F306 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 1.7749 | 0.0267 | 7 | 1.738 | 1.82 | 0.074 | 0.01 |  | 0.161 | Martine Carrier | 2022-10-17
F | F307 | Pour Point - ASTM D5949 | °C | -55 | -6.6 | -31.7 | 3.5 | 9 | -37 | -27 | 9.793 | 6.3 |  | 0.281 | Martine Carrier | 2022-10-21
F | F307 | Residue - ASTM D86 (Automatic) | Vol. % | 0.9 | -5.77 | 1.4 | 0.09 | 9 | 1.3 | 1.6 | 0.24 |  |  | 1.382 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 101.2 | -1.07 | 103.43 | 2.08 | 6 | 101.2 | 107.2 | 5.764 | 6.827 |  | 0.424 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 159.6 | -1.56 | 160.6 | 0.64 | 6 | 159.6 | 161.6 | 1.778 | 3.909 |  | 0.458 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 174.4 | -0.64 | 175.2 | 1.26 | 7 | 174.1 | 177.6 | 3.486 | 4.128 |  | 0.714 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 200.2 | -1.66 | 201.47 | 0.77 | 7 | 200.2 | 202.4 | 2.12 | 4.522 |  | 0.235 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 221.4 | -1.26 | 222.51 | 0.88 | 7 | 221.4 | 223.6 | 2.443 | 4.193 |  | 0.285 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 239.2 | -1.06 | 240.4 | 1.14 | 7 | 239.2 | 242.4 | 3.15 | 4.3 |  | 0.414 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 256.2 | -1.29 | 257.96 | 1.37 | 7 | 256.2 | 260.3 | 3.787 | 4.3 |  | 0.23 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 273.8 | -1.4 | 275.69 | 1.35 | 7 | 273.8 | 277.8 | 3.732 | 4.3 |  | 0.179 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 292.4 | -1.44 | 293.9 | 1.04 | 7 | 292.4 | 295.6 | 2.888 | 4.3 |  | 0.359 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 309.8 | -1.43 | 311.54 | 1.22 | 7 | 309.8 | 313.7 | 3.373 | 4.3 |  | 0.359 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 333.8 | -1.31 | 335.69 | 1.44 | 7 | 333.8 | 338.4 | 3.99 | 4.3 |  | 0.336 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 352.6 | -1.93 | 354.06 | 0.75 | 7 | 352.6 | 354.8 | 2.09 | 5 |  | 0.425 | Martine Carrier | 2022-10-21
F | F307 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 385.8 | -0.35 | 386.34 | 1.57 | 7 | 384 | 388.8 | 4.35 | 11.8 |  | 0.254 | Martine Carrier | 2022-10-21
F | F307 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 159.1 | 0.82 | 156.27 | 3.43 | 10 | 151.5 | 161.2 | 9.514 | 8.595 |  | 0.482 | Martine Carrier | 2022-10-21
F | F307 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 197 | 1.39 | 195.23 | 1.27 | 10 | 193 | 197 | 3.531 | 4.295 |  | 0.414 | Martine Carrier | 2022-10-21
F | F307 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 256.4 | 1.39 | 254.99 | 1.01 | 10 | 253.7 | 256.4 | 2.808 | 3 |  | 0.371 | Martine Carrier | 2022-10-21
F | F307 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 316.4 | 0.68 | 315.47 | 1.38 | 10 | 313.4 | 317.6 | 3.812 | 4.732 |  | 0.189 | Martine Carrier | 2022-10-21
F | F307 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 344.8 | 1 | 341.52 | 3.27 | 10 | 333.6 | 344.8 | 9.063 | 7.1 |  | 0.749 | Martine Carrier | 2022-10-21
F | F307 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | -0.41 | 98.3 | 0.24 | 10 | 98 | 98.7 | 0.679 |  |  | 0.617 | Martine Carrier | 2022-10-21
F | F307 | Loss - ASTM D86 (Automatic) | Vol. % | 0.9 | 2.05 | 0.35 | 0.27 | 10 | 0 | 0.9 | 0.742 |  |  | 0.432 | Martine Carrier | 2022-10-21
F | F307 | Acid Number - ASTM D664 | mg KOH/g | 0.026 | 0.53 | 0.016 | 0.019 | 6 | 0 | 0.05 | 0.053 | 0.013 |  | 0.606 | Martine Carrier | 2022-10-21
F | F307 | Cloud Point - ASTM D2500 | °C |  |  | -22.3 | 1.4 | 6 | -24 | -20 | 3.785 | 4 |  | 0.387 | Martine Carrier | 2022-10-21
F | F307 | Cloud Point - ASTM D5773 | °C | -22.4 | -0.04 | -22.37 | 0.76 | 11 | -23.8 | -20.9 | 2.114 | 2.5 |  | 0.337 | Martine Carrier | 2022-10-21
F | F307 | Density, at 15°C - ASTM D4052 | kg/m³ | 853.1 | 0.59 | 853.05 | 0.08 | 13 | 852.9 | 853.2 | 0.215 | 0.5 | 0.185 - Pass | 1.002 | Martine Carrier | 2022-10-21
F | F307 | Flash Point - ASTM D93 | °C | 55.4 | -0.24 | 55.95 | 2.3 | 12 | 53 | 62 | 6.384 | 3.972 |  | 0.883 | Martine Carrier | 2022-10-21
F | F307 | Nitrogen - ASTM D4629 | ppm wt/wt | 33 | -0.11 | 33.22 | 2.11 | 9 | 29 | 36 | 5.84 | 4.915 |  | 0.313 | Martine Carrier | 2022-10-21
F | F307 | Pour Point - ASTM D97 | °C |  |  | -31.5 | 3.1 | 6 | -36 | -27 | 8.716 | 9 |  | 0.293 | Martine Carrier | 2022-10-21
F | F307 | Sulfur - ASTM D5453 | mg/kg | 9.2 | 0.16 | 9.01 | 1.17 | 11 | 6.9 | 10.9 | 3.229 | 3.014 |  | 0.171 | Martine Carrier | 2022-10-21
F | F307 | Sulfur - ASTM D7039 | mg/kg |  |  | 9.47 | 0.71 | 3 | 8.7 | 10.1 | 1.965 | 2.486 |  | 0.33 | Martine Carrier | 2022-10-21
F | F307 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 8e-05 | 0.0001 | 4 | 0 | 0.0002 | 0 | 0 |  | 0.464 | Martine Carrier | 2022-10-21
F | F307 | Viscosity at 40°C - ASTM D445 | cSt | 2.279 | 0.32 | 2.2774 | 0.0052 | 8 | 2.27 | 2.286 | 0.014 | 0.027 |  | 0.295 | Martine Carrier | 2022-10-21
F | F307 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.267 | 0.0179 | 6 | 2.243 | 2.286 | 0.05 | 0.013 |  | 0.363 | Martine Carrier | 2022-10-21
F | F308 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 110.8 | -0.67 | 112.01 | 1.81 | 8 | 110 | 115.2 | 5.011 | 7.393 |  | 0.31 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 172 | 0.22 | 171.58 | 1.89 | 8 | 167.5 | 173.4 | 5.247 | 4.074 |  | 0.578 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 203 | -0.28 | 203.53 | 1.86 | 8 | 199.7 | 206.4 | 5.158 | 4.553 |  | 0.659 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 238.6 | -0.42 | 239.41 | 1.95 | 8 | 235.7 | 242.8 | 5.392 | 5.091 |  | 0.678 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 260 | -0.31 | 260.59 | 1.89 | 8 | 256.7 | 263.5 | 5.24 | 4.688 |  | 0.659 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 285.8 | 1.99 | 278.81 | 3.51 | 8 | 273.4 | 285.8 | 9.733 | 4.3 |  | 0.534 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 293.8 | -0.31 | 294.35 | 1.75 | 8 | 290.7 | 296.6 | 4.843 | 4.3 |  | 0.509 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 308 | -0.34 | 308.68 | 1.98 | 8 | 304.7 | 311.2 | 5.472 | 4.3 |  | 0.367 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 323.2 | -0.48 | 324.31 | 2.29 | 8 | 319.9 | 327.3 | 6.357 | 4.3 |  | 0.336 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 344.4 | -0.26 | 344.96 | 2.15 | 8 | 340.9 | 347.9 | 5.945 | 4.3 |  | 0.262 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 371.4 | -0.47 | 372.5 | 2.33 | 8 | 368.8 | 375.3 | 6.447 | 4.3 |  | 0.287 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 392.8 | -0.36 | 393.7 | 2.49 | 8 | 389.7 | 396.2 | 6.886 | 5 |  | 0.463 | Eric Salvail | 2023-02-13
F | F308 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 444.8 | -0.21 | 446.33 | 7.16 | 8 | 430.2 | 454 | 19.823 | 11.8 |  | 0.805 | Eric Salvail | 2023-02-13
F | F308 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 157.6 | -1.1 | 160.87 | 2.97 | 15 | 156.1 | 165.8 | 8.24 | 8.848 | 0.867 - Pass | 0.606 | Eric Salvail | 2023-02-13
F | F308 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 226.1 | 1.71 | 222.72 | 1.98 | 15 | 219.5 | 226.1 | 5.483 | 4.9 | 1.252 - Pass | 0.302 | Eric Salvail | 2023-02-13
F | F308 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 288.7 | 1.2 | 287.59 | 0.93 | 15 | 285.7 | 288.8 | 2.569 | 3 | 0.733 - Pass | 0.398 | Eric Salvail | 2023-02-13
F | F308 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 352 | 0.53 | 351.05 | 1.8 | 15 | 347.7 | 354.8 | 4.986 | 5.266 | 0.897 - Pass | 0.435 | Eric Salvail | 2023-02-13
F | F308 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 383 | 0.26 | 382.45 | 2.11 | 15 | 377.9 | 386.2 | 5.833 | 7.1 | 0.675 - Pass | 0.161 | Eric Salvail | 2023-02-13
F | F308 | Recovery - ASTM D86 (Automatic) | Vol. % | 97.9 | 0.03 | 97.89 | 0.24 | 13 | 97.4 | 98.3 | 0.664 |  |  | 0.685 | Eric Salvail | 2023-02-13
F | F308 | Residue - ASTM D86 (Automatic) | Vol. % | 1.5 | -0.29 | 1.52 | 0.06 | 12 | 1.5 | 1.7 | 0.16 |  |  | 4.332 | Eric Salvail | 2023-02-13
F | F308 | Loss - ASTM D86 (Automatic) | Vol. % | 0.6 | 0.35 | 0.5 | 0.28 | 15 | 0.1 | 1.1 | 0.783 |  |  | 0.697 | Eric Salvail | 2023-02-13
F | F308 | Acid Number - ASTM D664 | mg KOH/g | 0.017 | 0.4 | 0.012 | 0.013 | 9 | 0 | 0.04 | 0.035 | 0.011 |  | 0.56 | Eric Salvail | 2023-02-13
F | F308 | Carbon Residue, Micro Conradson - ASTM D4530 | Mass % | 0 | -0.59 | 0.03 | 0.05 | 7 | 0 | 0.1 | 0.135 | 0.142 |  | 1.645 | Eric Salvail | 2023-02-13
F | F308 | Cloud Point - ASTM D2500 | °C |  |  | -14 | 0.5 | 8 | -15 | -13 | 1.481 | 4 |  | 1.474 | Eric Salvail | 2023-02-13
F | F308 | Cloud Point - ASTM D5773 | °C | -13 | 1.57 | -14.11 | 0.7 | 14 | -15.5 | -13 | 1.948 | 2.5 | 0.607 - Pass | 0.213 | Eric Salvail | 2023-02-13
F | F308 | Density, at 15°C - ASTM D4052 | kg/m³ | 865.2 | -0.68 | 865.25 | 0.07 | 16 | 865.1 | 865.4 | 0.202 | 0.5 | 0.164 - Pass | 1.353 | Eric Salvail | 2023-02-13
F | F308 | Flash Point - ASTM D93 | °C | 64 | -1.17 | 65.13 | 0.97 | 18 | 63.3 | 66.7 | 2.682 | 4.624 | 0.336 - Pass | 0.421 | Eric Salvail | 2023-02-13
F | F308 | Nitrogen - ASTM D4629 | ppm wt/wt | 62.1 | 0.43 | 59.81 | 5.35 | 10 | 49 | 67 | 14.832 | 6.653 |  | 0.355 | Eric Salvail | 2023-02-13
F | F308 | Pour Point - ASTM D97 | °C |  |  | -19.9 | 1.6 | 8 | -21 | -18 | 4.301 | 9 |  | 1.581 | Eric Salvail | 2023-02-13
F | F308 | Pour Point - ASTM D5949 | °C | -20 | -0.44 | -19.4 | 1.3 | 12 | -21 | -18 | 3.633 | 6.3 |  | 1.206 | Eric Salvail | 2023-02-13
F | F308 | Sulfur - ASTM D5453 | mg/kg | 9.2 | 0.37 | 9.06 | 0.37 | 13 | 8.3 | 9.6 | 1.027 | 3.028 | 0.115 - Pass | 0.234 | Eric Salvail | 2023-02-13
F | F308 | Sulfur - ASTM D7039 | mg/kg |  |  | 8.86 | 0.84 | 5 | 8.1 | 10.3 | 2.331 | 2.398 |  | 0.717 | Eric Salvail | 2023-02-13
F | F308 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 1.5 | 0.00013 | 0.00012 | 8 | 0 | 0.0003 | 0 | 0 |  | 0.863 | Eric Salvail | 2023-02-13
F | F308 | Viscosity at 40°C - ASTM D445 | cSt | 3.649 | 0.24 | 3.6439 | 0.0214 | 13 | 3.61 | 3.693 | 0.059 | 0.038 | 2.426 - Fail | 0.524 | Eric Salvail | 2023-02-13
F | F308 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.6415 | 0.0198 | 6 | 3.606 | 3.657 | 0.055 | 0.021 |  | 0.633 | Eric Salvail | 2023-02-13
F | F309 | Residue - ASTM D86 (Automatic) | Vol. % | 1.2 | -6.64 | 1.39 | 0.03 | 12 | 1.3 | 1.4 | 0.08 |  |  | 4.332 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 111.1 | -0.85 | 112.86 | 2.07 | 8 | 110.9 | 116.8 | 5.744 | 7.449 |  | 0.515 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 144.6 | -0.76 | 146.43 | 2.4 | 8 | 143.8 | 151.2 | 6.646 | 3.696 |  | 0.493 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 161.7 | -0.68 | 163.35 | 2.41 | 8 | 160.3 | 168.3 | 6.684 | 3.95 |  | 0.591 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 183.8 | -0.58 | 185.64 | 3.14 | 8 | 181.4 | 192 | 8.708 | 4.285 |  | 0.516 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 202.4 | -0.52 | 204.11 | 3.27 | 8 | 199.5 | 210.6 | 9.058 | 3.953 |  | 0.525 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 220.6 | -0.47 | 222.33 | 3.7 | 8 | 217 | 229.8 | 10.246 | 4.3 |  | 0.623 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 239.1 | -0.43 | 240.95 | 4.33 | 8 | 235.1 | 250.1 | 11.989 | 4.3 |  | 0.681 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 259.5 | -0.37 | 261.14 | 4.43 | 8 | 254.8 | 270.3 | 12.269 | 4.3 |  | 0.742 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 280.4 | -0.39 | 282.01 | 4.15 | 8 | 275.9 | 290.5 | 11.498 | 4.3 |  | 0.694 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 302.8 | -0.54 | 304.74 | 3.59 | 8 | 299.3 | 311.8 | 9.947 | 4.3 |  | 0.428 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 330.3 | -0.56 | 332.15 | 3.29 | 8 | 327.1 | 338.6 | 9.115 | 4.3 |  | 0.506 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 349.9 | -0.75 | 351.99 | 2.79 | 8 | 347.6 | 357.2 | 7.716 | 5 |  | 0.431 | Eric Salvail | 2023-04-18
F | F309 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 385.3 | -1.83 | 388.2 | 1.59 | 8 | 385.3 | 390.6 | 4.4 | 11.8 |  | 0.383 | Eric Salvail | 2023-04-18
F | F309 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 155.2 | 1.06 | 153.33 | 1.76 | 14 | 150.2 | 156.2 | 4.884 | 8.433 | 0.335 - Pass | 0.28 | Eric Salvail | 2023-04-18
F | F309 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 183.7 | 1.74 | 182.13 | 0.9 | 14 | 180.4 | 183.7 | 2.505 | 4.007 | 0.391 - Pass | 0.254 | Eric Salvail | 2023-04-18
F | F309 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 239.1 | 1.33 | 237.77 | 1 | 14 | 236.6 | 239.4 | 2.765 | 3 | 0.849 - Pass | 0.516 | Eric Salvail | 2023-04-18
F | F309 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 313.5 | 0.14 | 313.25 | 1.76 | 14 | 310.7 | 317.6 | 4.865 | 4.699 | 1.072 - Pass | 0.536 | Eric Salvail | 2023-04-18
F | F309 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 342.4 | 0.68 | 341.66 | 1.09 | 14 | 340.5 | 343.8 | 3.012 | 7.1 | 0.180 - Pass | 0.88 | Eric Salvail | 2023-04-18
F | F309 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.8 | 2.1 | 98.19 | 0.29 | 14 | 97.7 | 98.8 | 0.802 |  |  | 0.874 | Eric Salvail | 2023-04-18
F | F309 | Loss - ASTM D86 (Automatic) | Vol. % | 0 | -1.95 | 0.41 | 0.21 | 14 | 0 | 0.8 | 0.58 |  |  | 1.229 | Eric Salvail | 2023-04-18
F | F309 | Acid Number - ASTM D664 | mg KOH/g | 0.015 | 0.27 | 0.012 | 0.012 | 8 | 0 | 0.03 | 0.034 | 0.01 |  | 0.664 | Eric Salvail | 2023-04-18
F | F309 | Cloud Point - ASTM D2500 | °C |  |  | -33 | 2.2 | 7 | -36 | -30 | 5.984 | 4 |  | 0.154 | Eric Salvail | 2023-04-18
F | F309 | Cloud Point - ASTM D5773 | °C | -35.2 | -0.06 | -35.08 | 1.99 | 13 | -37.4 | -31 | 5.517 | 2.5 | 4.869 - Fail | 0.515 | Eric Salvail | 2023-04-18
F | F309 | Density, at 15°C - ASTM D4052 | kg/m³ | 839.3 | -1.34 | 839.44 | 0.1 | 16 | 839.2 | 839.6 | 0.284 | 0.5 | 0.322 - Pass | 0.869 | Eric Salvail | 2023-04-18
F | F309 | Flash Point - ASTM D93 | °C | 50.8 | -1.42 | 52.87 | 1.46 | 18 | 50.8 | 56.5 | 4.032 | 3.754 | 1.154 - Pass | 0.272 | Eric Salvail | 2023-04-18
F | F309 | Nitrogen - ASTM D4629 | ppm wt/wt | 1.6 | 0.15 | 1.53 | 0.49 | 11 | 0.4 | 2.3 | 1.352 | 1.007 |  | 0.389 | Eric Salvail | 2023-04-18
F | F309 | Pour Point - ASTM D97 | °C |  |  | -59.4 | 3.9 | 5 | -63 | -54 | 10.835 | 9 |  | 0.356 | Eric Salvail | 2023-04-18
F | F309 | Pour Point - ASTM D5949 | °C | -53 | 0.54 | -55.6 | 4.9 | 11 | -64 | -48 | 13.472 | 6.3 |  | 0.174 | Eric Salvail | 2023-04-18
F | F309 | Sulfur - ASTM D5453 | mg/kg | 3.5 | 0.41 | 3.35 | 0.38 | 13 | 2.7 | 4 | 1.04 | 1.434 | 0.526 - Pass | 0.201 | Eric Salvail | 2023-04-18
F | F309 | Sulfur - ASTM D7039 | mg/kg |  |  | 3.73 | 0.45 | 7 | 3.2 | 4.5 | 1.246 | 1.503 |  | 0.218 | Eric Salvail | 2023-04-18
F | F309 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00014 | 0.00013 | 7 | 0 | 0.0003 | 0 | 0 |  | 0.434 | Eric Salvail | 2023-04-18
F | F309 | Viscosity at 40°C - ASTM D445 | cSt | 2.042 | 0.47 | 2.0393 | 0.0056 | 12 | 2.032 | 2.054 | 0.016 | 0.025 |  | 0.64 | Eric Salvail | 2023-04-18
F | F309 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.0446 | 0.0111 | 7 | 2.027 | 2.062 | 0.031 | 0.012 |  | 0.181 | Eric Salvail | 2023-04-18
F | F310 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 120.5 | -0.53 | 123.56 | 5.81 | 5 | 120.5 | 133.9 | 16.089 | 8.155 |  | 1.192 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 151 | -0.48 | 153.26 | 4.72 | 5 | 151 | 161.7 | 13.082 | 3.799 |  | 1.359 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 164.2 | -0.58 | 167.24 | 5.2 | 5 | 164.2 | 176.5 | 14.397 | 4.009 |  | 1.204 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 182.6 | -0.63 | 185.88 | 5.24 | 5 | 182.6 | 195.2 | 14.525 | 4.288 |  | 1.135 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 198.6 | -0.65 | 201.72 | 4.84 | 5 | 198.6 | 210.3 | 13.396 | 3.922 |  | 1.078 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 212.1 | -0.65 | 215.24 | 4.83 | 5 | 212.1 | 223.8 | 13.375 | 4.3 |  | 0.946 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 224.8 | -0.66 | 227.96 | 4.78 | 5 | 224.8 | 236.4 | 13.23 | 4.3 |  | 1.004 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 238.1 | -0.69 | 241.26 | 4.61 | 5 | 238.1 | 249.4 | 12.774 | 4.3 |  | 0.978 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 253.8 | -0.57 | 256.2 | 4.21 | 5 | 253.5 | 263.6 | 11.659 | 4.3 |  | 0.927 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 269.4 | -0.7 | 272.14 | 3.92 | 5 | 269.4 | 279 | 10.848 | 4.3 |  | 0.917 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 288 | -0.72 | 290.56 | 3.58 | 5 | 288 | 296.8 | 9.908 | 4.3 |  | 0.778 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 300.3 | -0.8 | 302.98 | 3.33 | 5 | 300.3 | 308.7 | 9.227 | 5 |  | 0.701 | Eric Salvail | 2023-05-17
F | F310 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 322.7 | -0.78 | 325.06 | 3.03 | 5 | 322.7 | 330.3 | 8.385 | 11.8 |  | 0.736 | Eric Salvail | 2023-05-17
F | F310 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 161.5 | 0.55 | 159.83 | 3.03 | 8 | 154.7 | 164.3 | 8.398 | 8.79 |  | 0.296 | Eric Salvail | 2023-05-17
F | F310 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 185.5 | 1.22 | 184.41 | 0.89 | 8 | 183.1 | 185.7 | 2.471 | 4.057 |  | 0.258 | Eric Salvail | 2023-05-17
F | F310 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 225.6 | 1.76 | 224.43 | 0.67 | 8 | 223.8 | 225.6 | 1.854 | 3 |  | 0.57 | Eric Salvail | 2023-05-17
F | F310 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 273.5 | 0.35 | 273.31 | 0.53 | 8 | 272.7 | 274.4 | 1.476 | 4.1 |  | 0.436 | Eric Salvail | 2023-05-17
F | F310 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 294.3 | 1.3 | 292.4 | 1.46 | 8 | 289.6 | 294.3 | 4.055 | 7.1 |  | 0.439 | Eric Salvail | 2023-05-17
F | F310 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.5 | 0.5 | 98.35 | 0.3 | 8 | 97.9 | 98.8 | 0.824 |  |  | 0.271 | Eric Salvail | 2023-05-17
F | F310 | Residue - ASTM D86 (Automatic) | Vol. % | 1.3 | 0.15 | 1.29 | 0.08 | 8 | 1.1 | 1.4 | 0.231 |  |  | 1.642 | Eric Salvail | 2023-05-17
F | F310 | Loss - ASTM D86 (Automatic) | Vol. % | 0.2 | -0.74 | 0.39 | 0.25 | 8 | 0.1 | 0.8 | 0.701 |  |  | 0.316 | Eric Salvail | 2023-05-17
F | F310 | Acid Number - ASTM D664 | mg KOH/g |  |  | 0.035 | 0.006 | 4 | 0.029 | 0.04 | 0.017 | 0.026 |  | 0.672 | Eric Salvail | 2023-05-17
F | F310 | Cloud Point - ASTM D2500 | °C |  |  | -51 | 1.6 | 4 | -53 | -49 | 4.523 | 4 |  | 0.376 | Eric Salvail | 2023-05-17
F | F310 | Cloud Point - ASTM D5773 | °C |  |  | -52.29 | 2.46 | 7 | -57.1 | -50 | 6.82 | 2.5 |  | 0.497 | Eric Salvail | 2023-05-17
F | F310 | Density, at 15°C - ASTM D4052 | kg/m³ |  |  | 836.69 | 0.06 | 8 | 836.6 | 836.8 | 0.178 | 0.5 |  | 1 | Eric Salvail | 2023-05-17
F | F310 | Flash Point - ASTM D93 | °C |  |  | 56.7 | 1.87 | 8 | 53.3 | 59 | 5.167 | 4.026 |  | 0.185 | Eric Salvail | 2023-05-17
F | F310 | Nitrogen - ASTM D4629 | ppm wt/wt |  |  | 43.5 | 4.72 | 6 | 36 | 48 | 13.081 | 5.647 |  | 0.371 | Eric Salvail | 2023-05-17
F | F310 | Pour Point - ASTM D5949 | °C |  |  | -66.8 | 4.1 | 5 | -70 | -60 | 11.32 | 6.3 |  | 0.699 | Eric Salvail | 2023-05-17
F | F310 | Sulfur - ASTM D5453 | mg/kg |  |  | 0.66 | 0.32 | 5 | 0.1 | 0.9 | 0.889 | 0.424 |  | 0.845 | Eric Salvail | 2023-05-17
F | F310 | Sulfur - ASTM D7039 | mg/kg |  |  | 0.73 | 0.54 | 4 | 0.3 | 1.5 | 1.507 | 0.621 |  | 0.454 | Eric Salvail | 2023-05-17
F | F310 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 8e-05 | 0.00011 | 5 | 0 | 0.0002 | 0 | 0 |  | 0.945 | Eric Salvail | 2023-05-17
F | F310 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 1.7122 | 0.0092 | 6 | 1.704 | 1.729 | 0.025 | 0.022 |  | 0.518 | Eric Salvail | 2023-05-17
F | F310 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 1.712 | 0.0087 | 3 | 1.702 | 1.717 | 0.024 | 0.01 |  | 0.732 | Eric Salvail | 2023-05-17
F | F311 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 119.3 | -0.17 | 119.64 | 2.07 | 7 | 117.5 | 123.4 | 5.721 | 7.896 |  | 0.404 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 169.5 | -0.77 | 170.6 | 1.43 | 7 | 168.7 | 173.1 | 3.956 | 4.059 |  | 0.19 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 194.7 | -0.89 | 196.04 | 1.52 | 7 | 193.5 | 197.7 | 4.197 | 4.441 |  | 0.316 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 222.3 | -0.98 | 224.13 | 1.87 | 7 | 221.3 | 226.7 | 5.172 | 4.862 |  | 0.213 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 242.5 | -0.88 | 244.3 | 2.04 | 7 | 241 | 247.1 | 5.652 | 4.476 |  | 0.223 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 259.8 | -0.82 | 261.46 | 2.03 | 7 | 258.1 | 264.3 | 5.629 | 4.3 |  | 0.267 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 275.3 | -0.82 | 277.17 | 2.29 | 7 | 273.4 | 280.3 | 6.344 | 4.3 |  | 0.261 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 290.7 | -0.93 | 292.67 | 2.12 | 7 | 289.3 | 295.4 | 5.862 | 4.3 |  | 0.29 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 305.7 | -0.93 | 307.73 | 2.19 | 7 | 304.2 | 310.3 | 6.07 | 4.3 |  | 0.274 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 322.9 | -0.99 | 325 | 2.12 | 7 | 321.6 | 327.6 | 5.88 | 4.3 |  | 0.317 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 346.4 | -0.87 | 349.78 | 3.9 | 8 | 345.4 | 358.3 | 10.791 | 4.3 |  | 0.672 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 364.1 | -1.06 | 367.15 | 2.88 | 8 | 363.8 | 373.1 | 7.986 | 5 |  | 0.521 | Eric Salvail | 2023-08-23
F | F311 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 416.4 | -0.55 | 417.63 | 2.23 | 7 | 415.4 | 422.4 | 6.187 | 11.8 |  | 0.981 | Eric Salvail | 2023-08-23
F | F311 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 170.7 | 1.02 | 167.04 | 3.61 | 14 | 158.8 | 171.7 | 9.999 | 9.187 | 1.185 - Pass | 0.267 | Eric Salvail | 2023-08-23
F | F311 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 217.3 | 1.14 | 215.47 | 1.61 | 14 | 212.4 | 218.2 | 4.455 | 4.74 | 0.883 - Pass | 0.178 | Eric Salvail | 2023-08-23
F | F311 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 274.9 | 1.68 | 273.56 | 0.8 | 14 | 272.1 | 274.9 | 2.213 | 3 | 0.544 - Pass | 0.269 | Eric Salvail | 2023-08-23
F | F311 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 330.1 | 0.44 | 329.44 | 1.49 | 14 | 327.6 | 332.4 | 4.127 | 4.942 | 0.697 - Pass | 0.398 | Eric Salvail | 2023-08-23
F | F311 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 358.8 | 1.28 | 355.99 | 2.2 | 14 | 350.3 | 358.8 | 6.084 | 7.1 | 0.734 - Pass | 0.437 | Eric Salvail | 2023-08-23
F | F311 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | 0.69 | 98 | 0.29 | 14 | 97.4 | 98.4 | 0.798 |  |  | 0.284 | Eric Salvail | 2023-08-23
F | F311 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 |  | 1.4 | 0 | 12 | 1.4 | 1.4 | 0 |  |  |  | Eric Salvail | 2023-08-23
F | F311 | Loss - ASTM D86 (Automatic) | Vol. % | 0.4 | -0.8 | 0.62 | 0.28 | 14 | 0.2 | 1.2 | 0.77 |  |  | 0.429 | Eric Salvail | 2023-08-23
F | F311 | Acid Number - ASTM D664 | mg KOH/g | 0.041 | 0.49 | 0.034 | 0.015 | 9 | 0 | 0.05 | 0.042 | 0.025 |  | 0.672 | Eric Salvail | 2023-08-23
F | F311 | Cloud Point - ASTM D2500 | °C |  |  | -30.1 | 1.3 | 7 | -32 | -28 | 3.726 | 4 |  | 0.246 | Eric Salvail | 2023-08-23
F | F311 | Cloud Point - ASTM D5773 | °C | -32.6 | -0.61 | -32.04 | 0.92 | 12 | -33.3 | -30.3 | 2.545 | 2.5 |  | 0.389 | Eric Salvail | 2023-08-23
F | F311 | Density, at 15°C - ASTM D4052 | kg/m³ | 860.6 | -1.04 | 860.67 | 0.07 | 15 | 860.6 | 860.8 | 0.195 | 0.5 | 0.152 - Pass | 1.403 | Eric Salvail | 2023-08-23
F | F311 | Flash Point - ASTM D93 | °C | 66.6 | -0.13 | 66.75 | 1.17 | 16 | 65 | 68.3 | 3.237 | 4.739 | 0.466 - Pass | 0.581 | Eric Salvail | 2023-08-23
F | F311 | Nitrogen - ASTM D4629 | ppm wt/wt | 4 | 0.73 | 3.53 | 0.64 | 11 | 2.5 | 4.5 | 1.783 | 1.549 |  | 0.224 | Eric Salvail | 2023-08-23
F | F311 | Pour Point - ASTM D97 | °C |  |  | -47.5 | 5.2 | 6 | -54 | -39 | 14.313 | 9 |  | 0.254 | Eric Salvail | 2023-08-23
F | F311 | Pour Point - ASTM D5949 | °C | -45 | 0.07 | -45.3 | 4.1 | 11 | -52 | -39 | 11.494 | 6.3 |  | 0.294 | Eric Salvail | 2023-08-23
F | F311 | Sulfur - ASTM D5453 | mg/kg | 5.3 | -0.14 | 5.38 | 0.56 | 14 | 4.5 | 6.5 | 1.56 | 2.047 | 0.581 - Pass | 0.173 | Eric Salvail | 2023-08-23
F | F311 | Sulfur - ASTM D7039 | mg/kg |  |  | 5.17 | 0.88 | 6 | 4.1 | 6.2 | 2.444 | 1.792 |  | 0.302 | Eric Salvail | 2023-08-23
F | F311 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00028 | 0.00022 | 8 | 0 | 0.0007 | 0.001 | 0 |  | 0.465 | Eric Salvail | 2023-08-23
F | F311 | Viscosity at 40°C - ASTM D445 | cSt | 3.24 | 0.28 | 3.2366 | 0.0121 | 12 | 3.216 | 3.256 | 0.033 | 0.035 |  | 0.336 | Eric Salvail | 2023-08-23
F | F311 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.2485 | 0.0054 | 6 | 3.242 | 3.256 | 0.015 | 0.019 |  | 0.241 | Eric Salvail | 2023-08-23
F | F312 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 107 | 0.57 | 105.64 | 2.39 | 8 | 100.8 | 108.6 | 6.621 | 6.972 |  | 0.389 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 159.5 | -0.22 | 159.78 | 1.22 | 8 | 157.5 | 161.9 | 3.392 | 3.897 |  | 0.583 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 178.6 | 0.09 | 178.45 | 1.72 | 8 | 174.7 | 180.6 | 4.773 | 4.177 |  | 0.703 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 210.7 | -0.23 | 211.23 | 2.3 | 8 | 207.4 | 215.7 | 6.37 | 4.668 |  | 0.596 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 234.3 | -0.4 | 235.23 | 2.29 | 8 | 231.1 | 239.4 | 6.339 | 4.358 |  | 0.666 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 253.9 | -0.2 | 254.36 | 2.36 | 8 | 250.5 | 259.1 | 6.531 | 4.3 |  | 0.77 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 271 | -0.26 | 271.63 | 2.4 | 8 | 267.7 | 276.4 | 6.649 | 4.3 |  | 0.709 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 288.1 | -0.36 | 288.94 | 2.35 | 8 | 285.7 | 294 | 6.523 | 4.3 |  | 0.858 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 304.9 | -0.41 | 305.91 | 2.46 | 8 | 302.2 | 310.9 | 6.824 | 4.3 |  | 0.546 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 324.4 | -0.4 | 325.58 | 2.92 | 8 | 321 | 331.3 | 8.092 | 4.3 |  | 0.423 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 354.2 | -0.25 | 354.94 | 2.99 | 8 | 350.6 | 360.8 | 8.294 | 4.3 |  | 0.366 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 378.7 | -0.13 | 379.16 | 3.58 | 8 | 375.1 | 386.4 | 9.91 | 5 |  | 0.491 | Martine Carrier | 2023-07-20
F | F312 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 442 | 0.54 | 436.48 | 10.26 | 8 | 419.1 | 453.1 | 28.417 | 11.8 |  | 0.247 | Martine Carrier | 2023-07-20
F | F312 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 157.2 | 0.5 | 155.93 | 2.56 | 14 | 151.2 | 159 | 7.1 | 8.576 | 0.685 - Pass | 0.488 | Martine Carrier | 2023-07-20
F | F312 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 200.6 | 1.25 | 199.02 | 1.27 | 14 | 196.7 | 200.6 | 3.511 | 4.378 | 0.643 - Pass | 0.314 | Martine Carrier | 2023-07-20
F | F312 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 269 | 1.21 | 267.69 | 1.08 | 14 | 265.7 | 269.1 | 2.998 | 3 | 0.999 - Pass | 0.334 | Martine Carrier | 2023-07-20
F | F312 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 336.7 | 0.72 | 335.31 | 1.93 | 14 | 332.2 | 338.9 | 5.332 | 5.03 | 1.124 - Pass | 0.127 | Martine Carrier | 2023-07-20
F | F312 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 372.5 | 0.28 | 371.51 | 3.56 | 14 | 364.9 | 377.2 | 9.85 | 7.1 | 1.925 - Pass | 0.284 | Martine Carrier | 2023-07-20
F | F312 | Recovery - ASTM D86 (Automatic) | Vol. % | 97.7 | -1.05 | 98.04 | 0.33 | 14 | 97.6 | 98.6 | 0.907 |  |  | 0.374 | Martine Carrier | 2023-07-20
F | F312 | Residue - ASTM D86 (Automatic) | Vol. % | 1.5 | 1.4 | 1.38 | 0.08 | 12 | 1.2 | 1.5 | 0.231 |  |  | 1.06 | Martine Carrier | 2023-07-20
F | F312 | Loss - ASTM D86 (Automatic) | Vol. % | 0.8 | 0.99 | 0.57 | 0.23 | 14 | 0.1 | 0.9 | 0.638 |  |  | 0.392 | Martine Carrier | 2023-07-20
F | F312 | Acid Number - ASTM D664 | mg KOH/g | 0.02 | 2.1 | 0.005 | 0.007 | 8 | 0 | 0.02 | 0.02 | 0.005 |  | 1.037 | Martine Carrier | 2023-07-20 |  | recheck demandé
F | F312 | Cloud Point - ASTM D2500 | °C |  |  | -20.3 | 0.7 | 8 | -21 | -19 | 1.959 | 4 |  | 0.805 | Martine Carrier | 2023-07-20
F | F312 | Cloud Point - ASTM D5773 | °C | -19.7 | 0.98 | -20.42 | 0.74 | 13 | -21.8 | -19.4 | 2.052 | 2.5 | 0.673 - Pass | 0.279 | Martine Carrier | 2023-07-20
F | F312 | Density, at 15°C - ASTM D4052 | kg/m³ | 856.6 | -0.07 | 856.61 | 0.09 | 16 | 856.4 | 856.8 | 0.257 | 0.5 | 0.265 - Pass | 1.194 | Martine Carrier | 2023-07-20
F | F312 | Flash Point - ASTM D93 | °C | 56.2 | -0.81 | 57.45 | 1.55 | 17 | 53.3 | 59.5 | 4.281 | 4.079 | 1.102 - Pass | 0.51 | Martine Carrier | 2023-07-20
F | F312 | Nitrogen - ASTM D4629 | ppm wt/wt | 52.61 | 0.6 | 49.86 | 4.61 | 11 | 42 | 57 | 12.779 | 6.058 |  | 0.209 | Martine Carrier | 2023-07-20
F | F312 | Pour Point - ASTM D97 | °C |  |  | -26.3 | 2.1 | 8 | -30 | -24 | 5.876 | 9 |  | 0.805 | Martine Carrier | 2023-07-20
F | F312 | Pour Point - ASTM D5949 | °C | -26 | 0.43 | -26.7 | 1.7 | 11 | -30 | -24 | 4.65 | 6.3 |  | 0.976 | Martine Carrier | 2023-07-20
F | F312 | Sulfur - ASTM D5453 | mg/kg | 8.7 | 0.33 | 8.46 | 0.71 | 14 | 6.7 | 9.9 | 1.976 | 2.877 | 0.472 - Pass | 0.603 | Martine Carrier | 2023-07-20
F | F312 | Sulfur - ASTM D7039 | mg/kg |  |  | 7.7 | 0.94 | 6 | 6.4 | 8.8 | 2.598 | 2.223 |  | 0.509 | Martine Carrier | 2023-07-20
F | F312 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00013 | 0.00014 | 6 | 0 | 0.0004 | 0 | 0 |  | 1.163 | Martine Carrier | 2023-07-20
F | F312 | Viscosity at 40°C - ASTM D445 | cSt | 2.704 | 0.49 | 2.6986 | 0.011 | 11 | 2.68 | 2.717 | 0.03 | 0.03 |  | 0.172 | Martine Carrier | 2023-07-20
F | F312 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.7173 | 0.0151 | 6 | 2.697 | 2.738 | 0.042 | 0.016 |  | 0.234 | Martine Carrier | 2023-07-20
F | F313 | Density, at 15°C - ASTM D4052 | kg/m³ | 835.5 | -16.74 | 838.49 | 0.18 | 16 | 838.2 | 838.9 | 0.494 | 0.5 | 0.977 - Pass | 1.619 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 102.4 | -0.32 | 103.09 | 2.13 | 9 | 100.1 | 108.1 | 5.887 | 6.804 |  | 1.006 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 143.7 | -1.05 | 144.63 | 0.89 | 9 | 143.4 | 146.4 | 2.458 | 3.669 |  | 0.381 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 161.6 | -0.91 | 162.62 | 1.12 | 9 | 160.5 | 164.4 | 3.102 | 3.939 |  | 0.323 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 189.2 | -0.68 | 190.07 | 1.28 | 9 | 187.5 | 192 | 3.542 | 4.351 |  | 0.299 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 212.9 | -1.94 | 214.44 | 0.79 | 8 | 212.9 | 215.4 | 2.196 | 4.088 |  | 0.314 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 233 | -1.92 | 234.74 | 0.9 | 8 | 233 | 235.6 | 2.504 | 4.3 |  | 0.416 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 252.7 | -1.36 | 253.86 | 0.86 | 8 | 252.7 | 255 | 2.373 | 4.3 |  | 0.316 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 271 | -0.94 | 271.96 | 1.02 | 8 | 271 | 273.7 | 2.828 | 4.3 |  | 0.437 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 291 | -1.6 | 293.04 | 1.27 | 8 | 291 | 294.8 | 3.525 | 4.3 |  | 0.197 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 311.8 | -1.81 | 314.09 | 1.26 | 8 | 311.8 | 315.6 | 3.502 | 4.3 |  | 0.354 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 340.3 | -1.82 | 342.59 | 1.26 | 8 | 340.3 | 344.1 | 3.48 | 4.3 |  | 0.3 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 362 | -1.95 | 364.09 | 1.07 | 8 | 362 | 365.6 | 2.963 | 5 |  | 0.314 | Eric Salvail | 2023-10-16
F | F313 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 415.9 | 0.36 | 414.57 | 3.74 | 9 | 406 | 417.4 | 10.371 | 11.8 |  | 1.215 | Eric Salvail | 2023-10-16
F | F313 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 151.7 | 1.01 | 147.87 | 3.8 | 14 | 140.6 | 154.6 | 10.537 | 8.133 | 1.679 - Pass | 0.162 | Eric Salvail | 2023-10-16
F | F313 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 181.4 | 0.17 | 181.18 | 1.34 | 14 | 178.3 | 183.2 | 3.712 | 3.986 | 0.867 - Pass | 0.277 | Eric Salvail | 2023-10-16
F | F313 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 250.6 | 0.56 | 250.06 | 0.98 | 14 | 248.3 | 252 | 2.703 | 3 | 0.812 - Pass | 0.339 | Eric Salvail | 2023-10-16
F | F313 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 322.4 | -0.58 | 323.04 | 1.1 | 14 | 321.5 | 324.6 | 3.045 | 4.846 | 0.395 - Pass | 0.481 | Eric Salvail | 2023-10-16
F | F313 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 355.6 | -0.28 | 356.19 | 2.1 | 14 | 353.9 | 361.6 | 5.816 | 7.1 | 0.671 - Pass | 0.497 | Eric Salvail | 2023-10-16
F | F313 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.4 | 1.02 | 98.16 | 0.24 | 14 | 97.6 | 98.6 | 0.658 |  |  | 0.462 | Eric Salvail | 2023-10-16
F | F313 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -1.61 | 1.34 | 0.15 | 14 | 1 | 1.6 | 0.417 |  |  | 1.76 | Eric Salvail | 2023-10-16
F | F313 | Loss - ASTM D86 (Automatic) | Vol. % | 0.5 | -0.03 | 0.51 | 0.23 | 14 | 0.2 | 1 | 0.629 |  |  | 0.569 | Eric Salvail | 2023-10-16
F | F313 | Acid Number - ASTM D664 | mg KOH/g |  |  | 0.005 | 0.008 | 6 | 0 | 0.02 | 0.022 | 0.005 |  | 1.043 | Eric Salvail | 2023-10-16
F | F313 | Cloud Point - ASTM D2500 | °C |  |  | -17.4 | 1.8 | 8 | -20 | -14 | 4.897 | 4 |  | 0.59 | Eric Salvail | 2023-10-16
F | F313 | Cloud Point - ASTM D5773 | °C | -17.7 | 0.54 | -18.23 | 0.98 | 13 | -19.4 | -15.8 | 2.715 | 2.5 | 1.179 - Pass | 0.649 | Eric Salvail | 2023-10-16
F | F313 | Flash Point - ASTM D93 | °C | 46.9 | -1.33 | 48.36 | 1.1 | 17 | 46.9 | 50.2 | 3.058 | 3.434 | 0.793 - Pass | 0.591 | Eric Salvail | 2023-10-16
F | F313 | Nitrogen - ASTM D4629 | ppm wt/wt | 8.9 | 0.97 | 8.21 | 0.71 | 11 | 6.9 | 8.9 | 1.964 | 2.393 |  | 0.553 | Eric Salvail | 2023-10-16
F | F313 | Pour Point - ASTM D97 | °C |  |  | -31.5 | 4.5 | 8 | -39 | -27 | 12.564 | 9 |  | 0.437 | Eric Salvail | 2023-10-16
F | F313 | Pour Point - ASTM D5949 | °C | -34 | -0.61 | -31.3 | 4.4 | 10 | -37 | -24 | 12.183 | 6.3 |  | 0.424 | Eric Salvail | 2023-10-16
F | F313 | Sulfur - ASTM D5453 | mg/kg | 9.5 | 0.36 | 9.19 | 0.87 | 14 | 8.1 | 11 | 2.405 | 3.059 | 0.618 - Pass | 0.391 | Eric Salvail | 2023-10-16
F | F313 | Sulfur - ASTM D7039 | mg/kg |  |  | 8.5 | 1.2 | 6 | 6.3 | 9.5 | 3.319 | 2.345 |  | 0.538 | Eric Salvail | 2023-10-16
F | F313 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0001 | -0.5 | 0.00016 | 0.00011 | 7 | 0 | 0.0003 | 0 | 0 |  | 0.499 | Eric Salvail | 2023-10-16
F | F313 | Viscosity at 40°C - ASTM D445 | cSt | 2.209 | 0.74 | 2.2038 | 0.0071 | 12 | 2.192 | 2.217 | 0.02 | 0.026 |  | 0.181 | Eric Salvail | 2023-10-16
F | F313 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 2.2018 | 0.0188 | 8 | 2.168 | 2.219 | 0.052 | 0.013 |  | 0.528 | Eric Salvail | 2023-10-16
F | F314 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -7.87 | 1.42 | 0.04 | 11 | 1.4 | 1.5 | 0.112 |  |  | 3.173 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 114.8 | -0.11 | 114.94 | 1.34 | 7 | 113 | 116.8 | 3.702 | 7.586 |  | 0.352 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 171.6 | 0.34 | 171.26 | 1.01 | 7 | 169.7 | 172.2 | 2.788 | 4.069 |  | 0.699 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 198.6 | 0.1 | 198.46 | 1.38 | 7 | 195.8 | 199.7 | 3.814 | 4.477 |  | 0.572 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 235 | -0.12 | 235.23 | 1.97 | 7 | 231.2 | 237.6 | 5.444 | 5.028 |  | 0.748 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 260.5 | -0.1 | 260.73 | 2.27 | 7 | 256.1 | 263.6 | 6.3 | 4.689 |  | 0.8 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 279.1 | -0.16 | 279.5 | 2.44 | 7 | 274.6 | 282.6 | 6.751 | 4.3 |  | 0.676 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 294.2 | -0.19 | 294.61 | 2.16 | 7 | 290.2 | 297.1 | 5.995 | 4.3 |  | 0.707 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 306 | -0.36 | 306.83 | 2.28 | 7 | 302.4 | 309.6 | 6.318 | 4.3 |  | 0.416 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 319.2 | -0.5 | 320.37 | 2.36 | 7 | 315.9 | 323.3 | 6.527 | 4.3 |  | 0.37 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 335.9 | -0.42 | 337.09 | 2.82 | 7 | 331.5 | 340.1 | 7.811 | 4.3 |  | 0.481 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 357 | -0.46 | 358.09 | 2.35 | 7 | 353.6 | 360.8 | 6.505 | 4.3 |  | 0.372 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 372.1 | -0.49 | 373.29 | 2.43 | 7 | 368.7 | 376.1 | 6.722 | 5 |  | 0.429 | Eric Salvail | 2023-11-10
F | F314 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 402.3 | -0.48 | 404.39 | 4.31 | 7 | 399.7 | 412.3 | 11.929 | 11.8 |  | 0.313 | Eric Salvail | 2023-11-10
F | F314 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 165.1 | 0.59 | 163.66 | 2.46 | 13 | 158.1 | 167.1 | 6.811 | 9.001 | 0.573 - Pass | 0.34 | Eric Salvail | 2023-11-10
F | F314 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 220.6 | 0.56 | 219.81 | 1.4 | 13 | 217 | 221.8 | 3.89 | 4.836 | 0.647 - Pass | 0.362 | Eric Salvail | 2023-11-10
F | F314 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 289.4 | 0.87 | 288.69 | 0.82 | 13 | 287.3 | 290 | 2.26 | 3 | 0.568 - Pass | 0.246 | Eric Salvail | 2023-11-10
F | F314 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 336.2 | -0.06 | 336.28 | 1.46 | 13 | 334.2 | 338.3 | 4.055 | 5.044 | 0.646 - Pass | 0.413 | Eric Salvail | 2023-11-10
F | F314 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 359.8 | 1.04 | 357.5 | 2.22 | 13 | 352 | 359.8 | 6.148 | 7.1 | 0.750 - Pass | 0.72 | Eric Salvail | 2023-11-10
F | F314 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.3 | 0.67 | 98.05 | 0.37 | 13 | 97.5 | 98.7 | 1.015 |  |  | 0.309 | Eric Salvail | 2023-11-10
F | F314 | Loss - ASTM D86 (Automatic) | Vol. % | 0.6 | 0.07 | 0.58 | 0.31 | 13 | 0 | 1.1 | 0.87 |  |  | 0.384 | Eric Salvail | 2023-11-10
F | F314 | Acid Number - ASTM D664 | mg KOH/g |  |  | 0.003 | 0.005 | 8 | 0 | 0.01 | 0.013 | 0.003 |  | 2.008 | Eric Salvail | 2023-11-10
F | F314 | Carbon Residue, Micro Conradson - ASTM D4530 | Mass % | 0 | -0.41 | 0.02 | 0.04 | 6 | 0 | 0.1 | 0.113 | 0.141 |  | 1.901 | Eric Salvail | 2023-11-10
F | F314 | Cloud Point - ASTM D2500 | °C |  |  | -13.9 | 1.2 | 8 | -16 | -12 | 3.453 | 4 |  | 0.311 | Eric Salvail | 2023-11-10
F | F314 | Cloud Point - ASTM D5773 | °C | -13.5 | 1.08 | -14 | 0.46 | 11 | -14.8 | -13.5 | 1.287 | 2.5 |  | 0.482 | Eric Salvail | 2023-11-10
F | F314 | Density, at 15°C - ASTM D4052 | kg/m³ | 860.5 | 1.24 | 860.43 | 0.06 | 15 | 860.3 | 860.5 | 0.164 | 0.5 | 0.108 - Pass | 1.976 | Eric Salvail | 2023-11-10
F | F314 | Flash Point - ASTM D93 | °C | 65 | -0.5 | 65.52 | 1.03 | 17 | 64 | 67.8 | 2.849 | 4.652 | 0.375 - Pass | 0.689 | Eric Salvail | 2023-11-10
F | F314 | Nitrogen - ASTM D4629 | ppm wt/wt | 36 | 0.5 | 35 | 2 | 10 | 32 | 38 | 5.54 | 5.049 |  | 0.349 | Eric Salvail | 2023-11-10
F | F314 | Pour Point - ASTM D97 | °C |  |  | -18.4 | 2.5 | 8 | -21 | -15 | 6.935 | 9 |  | 0.656 | Eric Salvail | 2023-11-10
F | F314 | Pour Point - ASTM D5949 | °C | -17 | 0.74 | -18.4 | 1.9 | 10 | -21 | -15 | 5.256 | 6.3 |  | 0.327 | Eric Salvail | 2023-11-10
F | F314 | Sulfur - ASTM D5453 | mg/kg | 9.6 | -0.32 | 9.79 | 0.61 | 13 | 8.8 | 11.1 | 1.69 | 3.209 | 0.278 - Pass | 0.335 | Eric Salvail | 2023-11-10
F | F314 | Sulfur - ASTM D7039 | mg/kg |  |  | 9.08 | 0.96 | 5 | 7.4 | 9.7 | 2.667 | 2.43 |  | 0.849 | Eric Salvail | 2023-11-10
F | F314 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0 | -0.82 | 0.0002 | 0.00024 | 7 | 0 | 0.0007 | 0.001 | 0 |  | 0.677 | Eric Salvail | 2023-11-10
F | F314 | Viscosity at 40°C - ASTM D445 | cSt | 3.519 | 0.82 | 3.5112 | 0.0095 | 13 | 3.496 | 3.533 | 0.026 | 0.037 | 0.511 - Pass | 0.488 | Eric Salvail | 2023-11-10
F | F314 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.5278 | 0.0087 | 6 | 3.513 | 3.536 | 0.024 | 0.021 |  | 0.445 | Eric Salvail | 2023-11-10
F | F315 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 106.7 | 0.1 | 106.52 | 1.8 | 6 | 104 | 109.6 | 4.984 | 7.03 |  | 0.537 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 186.6 | 0 | 186.6 | 0.78 | 6 | 185.9 | 187.6 | 2.153 | 4.299 |  | 0.644 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 216.1 | -0.4 | 216.33 | 0.58 | 6 | 215.6 | 217.3 | 1.612 | 4.745 |  | 0.303 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 248.1 | -0.4 | 248.76 | 1.62 | 7 | 245.9 | 251.3 | 4.495 | 5.231 |  | 0.386 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 266.8 | -0.44 | 267.46 | 1.48 | 7 | 264.7 | 269.3 | 4.111 | 4.777 |  | 0.291 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 282.8 | -0.48 | 283.46 | 1.38 | 7 | 280.8 | 284.6 | 3.831 | 4.3 |  | 0.603 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 297 | -0.63 | 297.87 | 1.38 | 7 | 295.3 | 299 | 3.817 | 4.3 |  | 0.553 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 309.8 | -0.68 | 310.66 | 1.26 | 7 | 308.4 | 312 | 3.478 | 4.3 |  | 0.301 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 324.3 | -0.54 | 325.23 | 1.72 | 7 | 322.2 | 327.5 | 4.754 | 4.3 |  | 0.255 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 343.3 | -0.47 | 344.03 | 1.56 | 7 | 341.2 | 346 | 4.335 | 4.3 |  | 0.287 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 368.5 | -0.5 | 369.37 | 1.73 | 7 | 367 | 372 | 4.786 | 4.3 |  | 0.177 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 388.7 | -0.32 | 389.41 | 2.21 | 7 | 386.4 | 392.5 | 6.118 | 5 |  | 0.26 | Eric Salvail | 2024-02-07
F | F315 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 443.1 | 0.12 | 442.21 | 7.25 | 7 | 428.4 | 453 | 20.07 | 11.8 |  | 0.735 | Eric Salvail | 2024-02-07
F | F315 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 166.2 | 1.09 | 161.82 | 4.02 | 13 | 155.5 | 168.3 | 11.132 | 8.9 | 1.564 - Pass | 0.37 | Eric Salvail | 2024-02-07
F | F315 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 240.2 | 1.76 | 236.54 | 2.08 | 12 | 233.5 | 240.2 | 5.751 | 5.204 |  | 0.267 | Eric Salvail | 2024-02-07
F | F315 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 292.3 | 1.59 | 290.96 | 0.84 | 12 | 289.9 | 292.3 | 2.337 | 3 |  | 0.58 | Eric Salvail | 2024-02-07
F | F315 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 348.8 | 0.45 | 348 | 1.8 | 13 | 344.8 | 352 | 4.976 | 5.22 | 0.909 - Pass | 0.352 | Eric Salvail | 2024-02-07
F | F315 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 379.1 | 0.87 | 377.23 | 2.14 | 13 | 373.5 | 380.9 | 5.932 | 7.1 | 0.698 - Pass | 0.239 | Eric Salvail | 2024-02-07
F | F315 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | 0.43 | 98.04 | 0.38 | 13 | 97.1 | 98.6 | 1.046 |  |  | 0.508 | Eric Salvail | 2024-02-07
F | F315 | Residue - ASTM D86 (Automatic) | Vol. % | 1.5 | 0.32 | 1.45 | 0.15 | 13 | 1.2 | 1.8 | 0.402 |  |  | 1.042 | Eric Salvail | 2024-02-07
F | F315 | Loss - ASTM D86 (Automatic) | Vol. % | 0.3 | -0.54 | 0.52 | 0.4 | 13 | 0 | 1.5 | 1.107 |  |  | 0.484 | Eric Salvail | 2024-02-07
F | F315 | Acid Number - ASTM D664 | mg KOH/g | 0.02 | 1.03 | 0.009 | 0.011 | 8 | 0 | 0.03 | 0.031 | 0.008 |  | 0.716 | Eric Salvail | 2024-02-07
F | F315 | Cloud Point - ASTM D2500 | °C |  |  | -11.9 | 0.7 | 7 | -13 | -11 | 1.911 | 4 |  | 0.755 | Eric Salvail | 2024-02-07
F | F315 | Cloud Point - ASTM D5773 | °C | -11.2 | 0.52 | -11.48 | 0.54 | 11 | -12.3 | -10.8 | 1.491 | 2.5 |  | 0.797 | Eric Salvail | 2024-02-07
F | F315 | Density, at 15°C - ASTM D4052 | kg/m³ | 864.4 | 1.14 | 864.29 | 0.09 | 16 | 864.1 | 864.5 | 0.257 | 0.5 | 0.265 - Pass | 1.194 | Eric Salvail | 2024-02-07
F | F315 | Flash Point - ASTM D93 | °C | 66.8 | -0.04 | 66.87 | 1.77 | 15 | 64.5 | 71.7 | 4.889 | 4.748 | 1.061 - Pass | 0.704 | Eric Salvail | 2024-02-07
F | F315 | Nitrogen - ASTM D4629 | ppm wt/wt | 27 | 1.23 | 25.11 | 1.54 | 9 | 23 | 27 | 4.256 | 4.256 |  | 0.36 | Eric Salvail | 2024-02-07
F | F315 | Pour Point - ASTM D97 | °C |  |  | -18.2 | 0.4 | 6 | -19 | -18 | 1.131 | 9 |  | 1.901 | Eric Salvail | 2024-02-07
F | F315 | Pour Point - ASTM D5949 | °C | -13 | 1.39 | -16.6 | 2.6 | 11 | -20 | -12 | 7.252 | 6.3 |  | 0.457 | Eric Salvail | 2024-02-07
F | F315 | Sulfur - ASTM D5453 | mg/kg | 9.5 | 0 | 9.5 | 0.79 | 13 | 8.1 | 11.6 | 2.193 | 3.137 | 0.489 - Pass | 0.795 | Eric Salvail | 2024-02-07
F | F315 | Sulfur - ASTM D7039 | mg/kg |  |  | 8.48 | 0.96 | 5 | 7.1 | 9.4 | 2.667 | 2.342 |  | 0.343 | Eric Salvail | 2024-02-07
F | F315 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00024 | 0.00024 | 7 | 0 | 0.0007 | 0.001 | 0 |  | 0.567 | Eric Salvail | 2024-02-07
F | F315 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 4.0316 | 0.0317 | 10 | 4.006 | 4.098 | 0.088 | 0.041 |  | 1.317 | Eric Salvail | 2024-02-07
F | F315 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s |  |  | 3.925 | 0.2614 | 6 | 3.392 | 4.046 | 0.724 | 0.023 |  | 1.693 | Eric Salvail | 2024-02-07
M | M967 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.05 | 8.56 | 0.0413 | 0.001 | 16 | 0.04 | 0.043 | 0.003 | 0.006 | 0.243 - Pass | 0.768 | Eric Salvail | 2023-06-12
M | M967 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 29.7 | 1.04 | 28.33 | 1.32 | 22 | 25.8 | 30.6 | 3.668 | 4.7 | 0.609 - Pass | 0.361 | Eric Salvail | 2023-06-12
M | M967 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 47.9 | 0.26 | 47.7 | 0.78 | 21 | 46.6 | 49.6 | 2.167 |  |  | 0.537 | Eric Salvail | 2023-06-12
M | M967 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 96.7 | 1.23 | 95.91 | 0.64 | 22 | 94.5 | 97 | 1.775 |  |  | 0.184 | Eric Salvail | 2023-06-12
M | M967 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 135.8 | 1.56 | 133.4 | 1.54 | 22 | 131 | 136.3 | 4.263 |  |  | 0.436 | Eric Salvail | 2023-06-12
M | M967 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 194.5 | 0.8 | 192.7 | 2.26 | 22 | 186.5 | 196.7 | 6.258 | 7.1 | 0.777 - Pass | 0.325 | Eric Salvail | 2023-06-12
M | M967 | Recovery - ASTM D86 (Automatic) | Vol.% | 96.2 | -0.57 | 96.62 | 0.74 | 21 | 94.9 | 97.5 | 2.047 |  |  | 0.836 | Eric Salvail | 2023-06-12
M | M967 | Residue - ASTM D86 (Automatic) | Vol.% | 1.2 | 1.47 | 1.08 | 0.09 | 20 | 0.9 | 1.3 | 0.236 |  |  | 1.69 | Eric Salvail | 2023-06-12
M | M967 | Loss - ASTM D86 (Automatic) | Vol.% | 2.6 | 0.61 | 2.19 | 0.68 | 22 | 0.9 | 3.5 | 1.871 |  |  | 0.715 | Eric Salvail | 2023-06-12
M | M967 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 46.9 | -1.61 | 47.92 | 0.64 | 22 | 46.9 | 49.2 | 1.76 |  |  | 0.205 | Eric Salvail | 2023-06-12
M | M967 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 92.1 | -1.98 | 92.67 | 0.29 | 22 | 92.1 | 93.1 | 0.795 |  |  | 0.456 | Eric Salvail | 2023-06-12
M | M967 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 0.8 | 0.17 | 3 | 0.6 | 0.9 | 0.48 |  |  | 0.732 | Eric Salvail | 2023-06-12
M | M967 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 3.8 | -0.3 | 4.24 | 1.46 | 18 | 2.7 | 7.1 | 4.038 | 2.706 | 2.226 - Fail | 0.951 | Eric Salvail | 2023-06-12
M | M967 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 1.4 | -0.94 | 2.01 | 0.65 | 19 | 1 | 3.7 | 1.787 | 0.786 | 5.169 - Fail | 0.427 | Eric Salvail | 2023-06-12
M | M967 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  | 20 | 0.88 | 19.1 | 1 | 12 | 17 | 20 | 2.77 |  |  | 0.825 | Eric Salvail | 2023-06-12
M | M967 | Oxygen Content - All Methods | Mass % | 3.53 | 0.2 | 3.5 | 0.16 | 14 | 3.2 | 3.7 | 0.436 |  |  | 0.965 | Eric Salvail | 2023-06-12
M | M967 | Ethanol - All Methods | Vol. % | 8.94 | -0.54 | 9.125 | 0.344 | 17 | 8.56 | 9.9 | 0.953 |  |  | 0.235 | Eric Salvail | 2023-06-12
M | M967 | Density, at 15°C - ASTM D4052 | kg/m³ | 699.7 | 1.96 | 698.72 | 0.5 | 22 | 698.2 | 699.8 | 1.384 | 3.565 | 0.151 - Pass | 1.557 | Eric Salvail | 2023-06-12
M | M967 | Driveablity Index (ethanol corrected) - CAN/CGSB-3.511 | °C | 510 | 0.79 | 502.5 | 9.5 | 16 | 476 | 512 | 26.201 |  |  | 0.761 | Eric Salvail | 2023-06-12
M | M967 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 0.96 | 0.46 | 9 | 0 | 1.5 | 1.27 | 1.893 |  | 0.536 | Eric Salvail | 2023-06-12
M | M967 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.34 | 0.43 | 8 | 0 | 1.2 | 1.184 | 1.8 |  | 0.897 | Eric Salvail | 2023-06-12
M | M967 | Sulfur - ASTM D5453 | mg/kg | 1.9 | -0.02 | 1.91 | 0.45 | 18 | 1 | 2.8 | 1.256 | 0.942 | 1.778 - Pass | 0.4 | Eric Salvail | 2023-06-12
M | M967 | Sulfur - ASTM D7039 | mg/kg |  |  | 1.97 | 0.46 | 7 | 1.2 | 2.6 | 1.266 | 1.065 |  | 0.317 | Eric Salvail | 2023-06-12
M | M967 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 0.0003 | 0.00051 | 5 | 0 | 0.0012 | 0.001 | 0 |  | 0.902 | Eric Salvail | 2023-06-12
M | M967 | Vapor Pressure - ASTM D5191 | kPa | 87.8 | 1 | 87.1 | 0.7 | 23 | 85.7 | 88.4 | 1.943 |  |  | 0.242 | Eric Salvail | 2023-06-12
M | M967 | Octane Number, Research - ASTM D2699 |  | 98.4 | 2.73 | 97.82 | 0.21 | 22 | 97.5 | 98.4 | 0.586 | 0.7 | 0.700 - Pass | 0.677 | Eric Salvail | 2023-06-12 |  | recheck demandé
M | M967 | Octane Number, Motor - ASTM D2700 |  | 91.7 | 1.13 | 91.2 | 0.45 | 20 | 90.2 | 92 | 1.24 | 0.9 | 1.899 - Pass | 0.255 | Eric Salvail | 2023-06-12
M | M968 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 30.1 | -0.19 | 30.38 | 1.43 | 26 | 27.7 | 34.5 | 3.951 | 4.7 | 0.707 - Pass | 0.383 | Eric Salvail | 2023-07-14
M | M968 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 56.7 | -0.62 | 57.63 | 1.5 | 26 | 54 | 60.4 | 4.166 |  |  | 0.21 | Eric Salvail | 2023-07-14
M | M968 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 109.1 | 0.61 | 108.82 | 0.46 | 27 | 107.9 | 109.8 | 1.272 |  |  | 0.278 | Eric Salvail | 2023-07-14
M | M968 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 158.4 | -2.35 | 160.37 | 0.84 | 27 | 158.3 | 161.7 | 2.315 |  |  | 0.727 | Eric Salvail | 2023-07-14
M | M968 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 208 | 0.13 | 207.71 | 2.21 | 27 | 203.7 | 213.6 | 6.135 | 7.1 | 0.747 - Pass | 0.236 | Eric Salvail | 2023-07-14
M | M968 | Recovery - ASTM D86 (Automatic) | Vol.% | 96.9 | -1.01 | 97.5 | 0.6 | 26 | 96.1 | 98.5 | 1.65 |  |  | 0.254 | Eric Salvail | 2023-07-14
M | M968 | Residue - ASTM D86 (Automatic) | Vol.% | 0.8 | -2.86 | 1.06 | 0.09 | 26 | 0.8 | 1.3 | 0.25 |  |  | 2.173 | Eric Salvail | 2023-07-14
M | M968 | Loss - ASTM D86 (Automatic) | Vol.% | 2.3 | 1.54 | 1.4 | 0.59 | 27 | 0.5 | 2.8 | 1.629 |  |  | 0.419 | Eric Salvail | 2023-07-14
M | M968 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 31.9 | -0.68 | 32.33 | 0.63 | 27 | 30.8 | 33.6 | 1.757 |  |  | 0.329 | Eric Salvail | 2023-07-14
M | M968 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 85.8 | 1.49 | 85.38 | 0.28 | 26 | 84.8 | 85.9 | 0.77 |  |  | 0.456 | Eric Salvail | 2023-07-14
M | M968 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 18.14 | 1.44 | 5 | 16.7 | 20.2 | 3.979 |  |  | 0.278 | Eric Salvail | 2023-07-14
M | M968 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 8.36 | 1.56 | 5 | 6.5 | 10.2 | 4.312 |  |  | 0.349 | Eric Salvail | 2023-07-14
M | M968 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 20.7 | 0.36 | 20.3 | 1.1 | 22 | 19.1 | 23.2 | 3.039 | 2.706 | 1.261 - Pass | 1.184 | Eric Salvail | 2023-07-14
M | M968 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 7.4 | -0.69 | 8.02 | 0.9 | 20 | 6.8 | 10.7 | 2.481 | 1.571 | 2.494 - Fail | 1.193 | Eric Salvail | 2023-07-14
M | M968 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.243 | 1.08 | 0.2347 | 0.0077 | 21 | 0.219 | 0.255 | 0.021 | 0.032 | 0.433 - Pass | 0.359 | Eric Salvail | 2023-07-14
M | M968 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 29.6 | 0.7 | 15 | 29 | 31 | 2.041 |  |  | 1.691 | Eric Salvail | 2023-07-14
M | M968 | Density, at 15°C - ASTM D4052 | kg/m³ | 729.1 | 2.16 | 728.13 | 0.45 | 27 | 727.5 | 729.2 | 1.245 | 2.639 | 0.223 - Pass | 1.173 | Eric Salvail | 2023-07-14
M | M968 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 572.6 | 3.8 | 18 | 567 | 578 | 10.5 |  |  | 0.51 | Eric Salvail | 2023-07-14
M | M968 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 17.93 | 2.94 | 20 | 11 | 23 | 8.135 | 6.116 | 1.769 - Pass | 0.281 | Eric Salvail | 2023-07-14
M | M968 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.64 | 0.74 | 10 | 0 | 1.8 | 2.045 | 2.181 |  | 0.933 | Eric Salvail | 2023-07-14
M | M968 | Induction Period at 100°C, ASTM D525 | min | 960 | -0.58 | 973 | 22.5 | 3 | 960 | 999 | 62.371 | 97.3 |  | 0.732 | Eric Salvail | 2023-07-14
M | M968 | Sulfur - ASTM D5453 | mg/kg | 10.7 | -0.69 | 11.25 | 0.8 | 24 | 9.8 | 13.4 | 2.205 | 3.56 | 0.384 - Pass | 0.557 | Eric Salvail | 2023-07-14
M | M968 | Sulfur - ASTM D7039 | mg/kg |  |  | 11.39 | 0.74 | 7 | 10.4 | 12.5 | 2.051 | 2.746 |  | 0.338 | Eric Salvail | 2023-07-14
M | M968 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0004 | 0.99 | 0.00027 | 0.00013 | 15 | 0.0001 | 0.0005 | 0 | 0 | 1.216 - Pass | 0.755 | Eric Salvail | 2023-07-14
M | M968 | Vapor Pressure - ASTM D5191 | kPa | 59.6 | 0.21 | 59.5 | 0.5 | 23 | 58.4 | 60.4 | 1.384 |  |  | 0.241 | Eric Salvail | 2023-07-14
M | M968 | Octane Number, Research - ASTM D2699 |  | 96 | 1.74 | 95.67 | 0.19 | 23 | 95.3 | 96 | 0.532 | 0.7 | 0.578 - Pass | 0.829 | Eric Salvail | 2023-07-14
M | M968 | Octane Number, Motor - ASTM D2700 |  | 91.7 | 9.7 | 87.69 | 0.41 | 21 | 86.6 | 88.6 | 1.146 | 0.9 | 1.623 - Pass | 0.42 | Eric Salvail | 2023-07-14 |  | recheck demandé
M | M969 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 27.4 | -0.7 | 28.11 | 1.01 | 21 | 26 | 29.7 | 2.81 | 4.7 | 0.357 - Pass | 0.527 | Eric Salvail | 2023-07-14
M | M969 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 42.5 | -2.46 | 44.34 | 0.75 | 23 | 42.5 | 46.5 | 2.069 |  |  | 1.138 | Eric Salvail | 2023-07-14
M | M969 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 69 | 0.05 | 68.95 | 0.95 | 22 | 67.5 | 71.2 | 2.64 |  |  | 1.283 | Eric Salvail | 2023-07-14
M | M969 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 175.8 | -2.23 | 177.17 | 0.61 | 22 | 175.8 | 178.2 | 1.7 |  |  | 0.39 | Eric Salvail | 2023-07-14
M | M969 | Recovery - ASTM D86 (Automatic) | Vol.% | 97.4 | -0.22 | 97.51 | 0.5 | 22 | 96.4 | 98.2 | 1.383 |  |  | 0.49 | Eric Salvail | 2023-07-14
M | M969 | Loss - ASTM D86 (Automatic) | Vol.% | 1.9 | 1.01 | 1.41 | 0.48 | 23 | 0.7 | 2.5 | 1.337 |  |  | 0.662 | Eric Salvail | 2023-07-14
M | M969 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 56.4 | -0.29 | 56.49 | 0.32 | 22 | 55.9 | 57.1 | 0.876 |  |  | 0.243 | Eric Salvail | 2023-07-14
M | M969 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 80.4 | 1.34 | 80.03 | 0.27 | 22 | 79.5 | 80.4 | 0.762 |  |  | 0.442 | Eric Salvail | 2023-07-14
M | M969 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 20.38 | 1.12 | 5 | 18.6 | 21.4 | 3.099 |  |  | 0.355 | Eric Salvail | 2023-07-14
M | M969 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 3.26 | 1.32 | 5 | 2.1 | 4.8 | 3.662 |  |  | 0.747 | Eric Salvail | 2023-07-14
M | M969 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 23.3 | 1.22 | 22.24 | 0.87 | 19 | 20.7 | 23.8 | 2.399 | 2.706 | 0.786 - Pass | 0.376 | Eric Salvail | 2023-07-14
M | M969 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 6.5 | 2.4 | 5.15 | 0.57 | 19 | 4 | 6.5 | 1.566 | 1.259 | 1.547 - Pass | 0.705 | Eric Salvail | 2023-07-14
M | M969 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.8 | 0.98 | 0.7822 | 0.0182 | 19 | 0.747 | 0.823 | 0.051 | 0.108 | 0.219 - Pass | 0.397 | Eric Salvail | 2023-07-14
M | M969 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 30.6 | 0.5 | 11 | 30 | 31 | 1.357 |  |  | 1.669 | Eric Salvail | 2023-07-14
M | M969 | Oxygen Content - All Methods | Mass % | 3.63 | 1.4 | 3.5 | 0.09 | 14 | 3.3 | 3.63 | 0.248 |  |  | 0.377 | Eric Salvail | 2023-07-14
M | M969 | Ethanol - All Methods | Vol. % | 9.65 | 0.45 | 9.53 | 0.269 | 17 | 9 | 9.97 | 0.744 |  |  | 0.253 | Eric Salvail | 2023-07-14
M | M969 | Density, at 15°C - ASTM D4052 | kg/m³ | 734.7 | 1.83 | 734.14 | 0.31 | 20 | 733.6 | 734.7 | 0.847 | 2.45 | 0.120 - Pass | 0.304 | Eric Salvail | 2023-07-14
M | M969 | Driveablity Index (ethanol corrected) - CAN/CGSB-3.511 | °C |  |  | 463.5 | 3.9 | 14 | 457 | 471 | 10.687 |  |  | 0.354 | Eric Salvail | 2023-07-14
M | M969 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 2.18 | 0.81 | 15 | 1.4 | 4 | 2.241 | 2.633 | 0.724 - Pass | 0.963 | Eric Salvail | 2023-07-14
M | M969 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.52 | 0.65 | 10 | 0 | 2 | 1.808 | 2.05 |  | 0.851 | Eric Salvail | 2023-07-14
M | M969 | Induction Period at 100°C, ASTM D525 | min | 960 | 0.83 | 810 | 180 | 4 | 600 | 960 | 498.6 | 81 |  | 0.494 | Eric Salvail | 2023-07-14
M | M969 | Sulfur - ASTM D5453 | mg/kg | 5.2 | 0.18 | 5.1 | 0.59 | 20 | 4 | 6.2 | 1.641 | 1.966 | 0.697 - Pass | 0.184 | Eric Salvail | 2023-07-14
M | M969 | Sulfur - ASTM D7039 | mg/kg |  |  | 4.86 | 0.46 | 7 | 4.2 | 5.4 | 1.278 | 1.734 |  | 0.388 | Eric Salvail | 2023-07-14
M | M969 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 0.91 | 0.00016 | 0.00015 | 8 | 0 | 0.0004 | 0 | 0 |  | 0.596 | Eric Salvail | 2023-07-14
M | M969 | Vapor Pressure - ASTM D5191 | kPa | 82.8 | 0.55 | 82.47 | 0.59 | 23 | 80.8 | 83.6 | 1.643 |  |  | 0.522 | Eric Salvail | 2023-07-14
M | M969 | Octane Number, Research - ASTM D2699 |  | 91.6 | 1.46 | 91.26 | 0.23 | 22 | 90.9 | 91.6 | 0.637 | 0.7 | 0.828 - Pass | 0.645 | Eric Salvail | 2023-07-14
M | M969 | Octane Number, Motor - ASTM D2700 |  | 83.8 | -0.35 | 83.98 | 0.52 | 22 | 83 | 85 | 1.435 | 0.9 | 2.541 - Fail | 0.571 | Eric Salvail | 2023-07-14
M | M969 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 218 | 4.89 | 212.53 | 1.12 | 22 | 210.7 | 214.9 | 3.1 | 7.1 | 0.191 - Pass | 0.242 | Eric Salvail | 2023-07-14 |  | Recheck demandé
M | M969 | Residue - ASTM D86 (Automatic) | Vol.% | 0.7 | -6.4 | 1.11 | 0.06 | 20 | 1 | 1.3 | 0.177 |  |  | 3.423 | Eric Salvail | 2023-07-14
M | M970 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 29.5 | 0.03 | 29.45 | 1.54 | 26 | 26.6 | 33.2 | 4.277 | 4.7 | 0.828 - Pass | 0.559 | Eric Salvail | 2023-09-25
M | M970 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 49 | -1.21 | 50.52 | 1.25 | 25 | 47.6 | 53.8 | 3.469 |  |  | 0.332 | Eric Salvail | 2023-09-25
M | M970 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 114.9 | -0.77 | 115.4 | 0.65 | 26 | 114.2 | 116.7 | 1.806 |  |  | 0.292 | Eric Salvail | 2023-09-25
M | M970 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 210.3 | 1.09 | 207.08 | 2.94 | 26 | 198.7 | 211.5 | 8.15 | 7.1 | 1.318 - Pass | 0.513 | Eric Salvail | 2023-09-25
M | M970 | Recovery - ASTM D86 (Automatic) | Vol.% | 97.4 | -0.69 | 97.77 | 0.53 | 25 | 96.8 | 99 | 1.467 |  |  | 0.164 | Eric Salvail | 2023-09-25
M | M970 | Loss - ASTM D86 (Automatic) | Vol.% | 1.8 | 1.17 | 1.23 | 0.49 | 26 | 0.4 | 2.2 | 1.347 |  |  | 0.287 | Eric Salvail | 2023-09-25
M | M970 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 37.7 | 0.83 | 37.33 | 0.45 | 26 | 36.5 | 38 | 1.243 |  |  | 0.357 | Eric Salvail | 2023-09-25
M | M970 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 81.1 | 1 | 80.64 | 0.46 | 26 | 79.5 | 81.5 | 1.264 |  |  | 0.325 | Eric Salvail | 2023-09-25
M | M970 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 46.17 | 1.35 | 7 | 44.1 | 47.9 | 3.736 |  |  | 0.188 | Eric Salvail | 2023-09-25
M | M970 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 0.91 | 0.2 | 7 | 0.6 | 1.1 | 0.541 |  |  | 0.528 | Eric Salvail | 2023-09-25
M | M970 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 48.2 | -0.63 | 48.73 | 0.85 | 21 | 47.1 | 50.4 | 2.359 | 2.706 | 0.760 - Pass | 0.259 | Eric Salvail | 2023-09-25
M | M970 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 0.8 | -0.48 | 0.96 | 0.34 | 21 | 0.4 | 1.7 | 0.937 | 0.545 | 2.954 - Fail | 0.7 | Eric Salvail | 2023-09-25
M | M970 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.206 | 1.06 | 0.1993 | 0.0063 | 21 | 0.188 | 0.215 | 0.017 | 0.028 | 0.404 - Pass | 0.476 | Eric Salvail | 2023-09-25
M | M970 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  | 47 | -0.23 | 47.2 | 0.8 | 16 | 46 | 49 | 2.351 |  |  | 0.665 | Eric Salvail | 2023-09-25
M | M970 | Density, at 15°C - ASTM D4052 | kg/m³ | 770.6 | 1.23 | 769.97 | 0.51 | 26 | 769.4 | 771.2 | 1.409 | 1.321 | 1.138 - Pass | 1.353 | Eric Salvail | 2023-09-25
M | M970 | Driveability Index - CAN/CGSB-3.5 | °C | 579 | -1.61 | 584.7 | 3.6 | 19 | 579 | 590 | 9.895 |  |  | 0.47 | Eric Salvail | 2023-09-25
M | M970 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 15.18 | 5.18 | 18 | 5 | 26 | 14.359 | 5.723 | 6.296 - Fail | 0.931 | Eric Salvail | 2023-09-25
M | M970 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.9 | 0.86 | 12 | 0 | 2.5 | 2.392 | 2.416 |  | 0.484 | Eric Salvail | 2023-09-25
M | M970 | Induction Period at 100°C, ASTM D525 | min | 999 | 0.7 | 893 | 151.1 | 3 | 720 | 999 | 418.508 | 89.3 |  | 0.539 | Eric Salvail | 2023-09-25
M | M970 | Sulfur - ASTM D5453 | mg/kg | 14.3 | 0.65 | 13.75 | 0.85 | 23 | 11.7 | 15 | 2.35 | 4.139 | 0.322 - Pass | 0.667 | Eric Salvail | 2023-09-25
M | M970 | Sulfur - ASTM D7039 | mg/kg |  |  | 13.63 | 1.07 | 7 | 11.5 | 14.8 | 2.974 | 3.026 |  | 0.537 | Eric Salvail | 2023-09-25
M | M970 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0013 | -0.22 | 0.00135 | 0.00024 | 17 | 0.0009 | 0.002 | 0.001 | 0 | 3.358 - Fail | 0.836 | Eric Salvail | 2023-09-25
M | M970 | Vapor Pressure - ASTM D5191 | kPa | 63.3 | 0.68 | 62.99 | 0.46 | 26 | 61.8 | 63.7 | 1.269 |  |  | 0.593 | Eric Salvail | 2023-09-25
M | M970 | Octane Number, Research - ASTM D2699 |  | 96.2 | 0.67 | 96.05 | 0.23 | 23 | 95.6 | 96.4 | 0.63 | 0.7 | 0.810 - Pass | 0.494 | Eric Salvail | 2023-09-25
M | M970 | Octane Number, Motor - ASTM D2700 |  | 85.8 | -0.84 | 86.13 | 0.39 | 22 | 85.2 | 86.7 | 1.082 | 0.9 | 1.445 - Pass | 0.327 | Eric Salvail | 2023-09-25
M | M970 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 160.8 | -4.35 | 163.2 | 0.55 | 25 | 161.8 | 164.3 | 1.531 |  |  | 0.241 | Eric Salvail | 2023-09-25
M | M970 | Residue - ASTM D86 (Automatic) | Vol.% | 0.8 | -4.9 | 1.01 | 0.04 | 22 | 0.9 | 1.1 | 0.118 |  |  | 5.119 | Eric Salvail | 2023-09-25
M | M971 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 32.2 | -0.25 | 32.5 | 1.17 | 23 | 30 | 34.7 | 3.245 | 4.7 | 0.477 - Pass | 0.424 | Eric Salvail | 2023-10-16
M | M971 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 49.7 | 0.42 | 49.54 | 0.39 | 22 | 48.6 | 50.1 | 1.087 |  |  | 0.306 | Eric Salvail | 2023-10-16
M | M971 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 69.8 | -1.15 | 70.84 | 0.9 | 23 | 68.7 | 72.6 | 2.502 |  |  | 0.321 | Eric Salvail | 2023-10-16
M | M971 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 161.7 | -2.24 | 163.91 | 0.99 | 23 | 161.7 | 165.6 | 2.734 |  |  | 0.586 | Eric Salvail | 2023-10-16
M | M971 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 215.1 | 2.2 | 211.83 | 1.49 | 23 | 208.6 | 215.1 | 4.115 | 7.1 | 0.336 - Pass | 0.309 | Eric Salvail | 2023-10-16 |  | recheck demandé
M | M971 | Recovery - ASTM D86 (Automatic) | Vol.% | 97.3 | -0.52 | 97.65 | 0.67 | 22 | 96.1 | 98.6 | 1.857 |  |  | 0.487 | Eric Salvail | 2023-10-16
M | M971 | Loss - ASTM D86 (Automatic) | Vol.% | 2 | 1.07 | 1.25 | 0.7 | 23 | 0.4 | 2.9 | 1.941 |  |  | 0.596 | Eric Salvail | 2023-10-16
M | M971 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 56.1 | -1.26 | 56.48 | 0.3 | 22 | 56 | 57.2 | 0.828 |  |  | 0.359 | Eric Salvail | 2023-10-16
M | M971 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 85.2 | 1.14 | 84.77 | 0.37 | 22 | 84.1 | 85.7 | 1.037 |  |  | 0.39 | Eric Salvail | 2023-10-16
M | M971 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 21.88 | 1.78 | 6 | 20 | 24.8 | 4.935 |  |  | 0.32 | Eric Salvail | 2023-10-16
M | M971 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 13.77 | 1.93 | 6 | 12 | 16.3 | 5.359 |  |  | 0.585 | Eric Salvail | 2023-10-16
M | M971 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 21.5 | 0.16 | 21.36 | 0.9 | 20 | 19.6 | 23.7 | 2.496 | 2.706 | 0.851 - Pass | 1.004 | Eric Salvail | 2023-10-16
M | M971 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 9.6 | -0.47 | 9.79 | 0.41 | 20 | 8.8 | 10.4 | 1.134 | 1.737 | 0.426 - Pass | 0.34 | Eric Salvail | 2023-10-16
M | M971 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.45 | 2.16 | 0.3918 | 0.027 | 20 | 0.34 | 0.45 | 0.075 | 0.054 | 1.909 - Pass | 0.329 | Eric Salvail | 2023-10-16 |  | recheck demandé
M | M971 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  | 26 | 1.01 | 25.3 | 0.7 | 13 | 24 | 27 | 2.008 |  |  | 1.426 | Eric Salvail | 2023-10-16
M | M971 | Oxygen Content - All Methods | Mass % | 3.77 | -0.16 | 3.82 | 0.34 | 15 | 3.405 | 4.6 | 0.939 |  |  | 0.885 | Eric Salvail | 2023-10-16
M | M971 | Ethanol - All Methods | Vol. % | 10.06 | -0.26 | 10.266 | 0.788 | 19 | 9.15 | 12.3 | 2.184 |  |  | 0.992 | Eric Salvail | 2023-10-16
M | M971 | Density, at 15°C - ASTM D4052 | kg/m³ | 741.7 | 2.47 | 741 | 0.28 | 23 | 740.6 | 741.7 | 0.781 | 2.233 | 0.122 - Pass | 0.66 | Eric Salvail | 2023-10-16
M | M971 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 463.3 | 6.5 | 15 | 452 | 473 | 18.033 |  |  | 0.494 | Eric Salvail | 2023-10-16
M | M971 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 6.78 | 0.69 | 15 | 5.8 | 8 | 1.905 | 4.146 | 0.211 - Pass | 0.442 | Eric Salvail | 2023-10-16
M | M971 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.56 | 0.32 | 10 | 0 | 1 | 0.878 | 2.096 |  | 0.456 | Eric Salvail | 2023-10-16
M | M971 | Induction Period at 100°C, ASTM D525 | min | 853 | 0.07 | 844.3 | 120.2 | 3 | 720 | 960 | 333.05 | 84.433 |  | 0.29 | Eric Salvail | 2023-10-16
M | M971 | Sulfur - ASTM D5453 | mg/kg | 5 | 0.47 | 4.81 | 0.4 | 20 | 4.3 | 5.8 | 1.111 | 1.883 | 0.348 - Pass | 0.573 | Eric Salvail | 2023-10-16
M | M971 | Sulfur - ASTM D7039 | mg/kg |  |  | 4.69 | 0.62 | 7 | 3.7 | 5.8 | 1.726 | 1.7 |  | 0.514 | Eric Salvail | 2023-10-16
M | M971 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 0.76 | 0.00019 | 0.00014 | 10 | 0 | 0.0005 | 0 | 0 |  | 0.569 | Eric Salvail | 2023-10-16
M | M971 | Vapor Pressure - ASTM D5191 | kPa | 69.2 | -0.14 | 69.26 | 0.39 | 23 | 68.4 | 70.1 | 1.088 |  |  | 0.443 | Eric Salvail | 2023-10-16
M | M971 | Octane Number, Research - ASTM D2699 |  | 93.4 | 0.48 | 93.32 | 0.17 | 22 | 93 | 93.7 | 0.473 | 0.7 | 0.457 - Pass | 0.591 | Eric Salvail | 2023-10-16
M | M971 | Octane Number, Motor - ASTM D2700 |  | 82.8 | -1.42 | 83.38 | 0.41 | 22 | 82.8 | 84.4 | 1.124 | 0.9 | 1.560 - Pass | 1.223 | Eric Salvail | 2023-10-16
M | M971 | Residue - ASTM D86 (Automatic) | Vol.% | 0.7 | -5.02 | 1.07 | 0.07 | 21 | 0.9 | 1.2 | 0.202 |  |  | 2.956 | Eric Salvail | 2023-10-16
M | M972 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 77.6 | -5.57 | 79.73 | 0.38 | 26 | 78.9 | 80.5 | 1.059 |  |  | 0.491 | Eric Salvail | 2023-11-10
M | M972 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 50 | -6.06 | 51.9 | 0.31 | 26 | 51.3 | 52.6 | 0.871 |  |  | 0.335 | Eric Salvail | 2023-11-10
M | M972 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 32.2 | 0.25 | 31.84 | 1.45 | 27 | 28.8 | 34.5 | 4.005 | 4.7 | 0.726 - Pass | 0.12 | Eric Salvail | 2023-11-10
M | M972 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 49.6 | -0.74 | 49.95 | 0.47 | 27 | 48.6 | 51 | 1.305 |  |  | 0.355 | Eric Salvail | 2023-11-10
M | M972 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 84.8 | -0.77 | 85.71 | 1.19 | 27 | 82.5 | 87.6 | 3.284 |  |  | 0.366 | Eric Salvail | 2023-11-10
M | M972 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 171.4 | 0.1 | 171.3 | 0.98 | 27 | 169.1 | 173.1 | 2.713 |  |  | 0.837 | Eric Salvail | 2023-11-10
M | M972 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 207.1 | -0.88 | 208.5 | 1.59 | 27 | 206.4 | 213.4 | 4.391 | 7.1 | 0.382 - Pass | 0.796 | Eric Salvail | 2023-11-10
M | M972 | Recovery - ASTM D86 (Automatic) | Vol.% | 97 | -1.43 | 97.66 | 0.46 | 26 | 96.5 | 98.5 | 1.271 |  |  | 0.627 | Eric Salvail | 2023-11-10
M | M972 | Loss - ASTM D86 (Automatic) | Vol.% | 2 | 1.54 | 1.22 | 0.51 | 27 | 0.4 | 2.7 | 1.403 |  |  | 0.722 | Eric Salvail | 2023-11-10
M | M972 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 20.34 | 2.48 | 5 | 18.2 | 24.6 | 6.857 |  |  | 0.731 | Eric Salvail | 2023-11-10
M | M972 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 15.62 | 2.87 | 5 | 12.9 | 20.3 | 7.94 |  |  | 0.484 | Eric Salvail | 2023-11-10
M | M972 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 19.8 | 0.43 | 19.42 | 0.89 | 23 | 17 | 21.1 | 2.458 | 2.706 | 0.825 - Pass | 0.891 | Eric Salvail | 2023-11-10
M | M972 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 12.4 | -0.42 | 12.55 | 0.36 | 22 | 11.5 | 13.1 | 0.988 | 1.966 | 0.253 - Pass | 0.457 | Eric Salvail | 2023-11-10
M | M972 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.698 | 1.96 | 0.6232 | 0.0382 | 24 | 0.559 | 0.698 | 0.106 | 0.086 | 1.511 - Pass | 0.288 | Eric Salvail | 2023-11-10
M | M972 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 26.5 | 0.8 | 16 | 25 | 28 | 2.231 |  |  | 0.98 | Eric Salvail | 2023-11-10
M | M972 | Oxygen Content - All Methods | Mass % | 3.7787 | 0.33 | 3.73 | 0.15 | 17 | 3.38 | 4 | 0.413 |  |  | 0.462 | Eric Salvail | 2023-11-10
M | M972 | Ethanol - All Methods | Vol. % | 10.1 | 0.21 | 10.013 | 0.421 | 21 | 9.11 | 10.8 | 1.166 |  |  | 0.505 | Eric Salvail | 2023-11-10
M | M972 | Density, at 15°C - ASTM D4052 | kg/m³ | 743 | 2.05 | 742.53 | 0.23 | 24 | 742.2 | 743 | 0.635 | 2.185 | 0.084 - Pass | 0.491 | Eric Salvail | 2023-11-10
M | M972 | Driveablity Index (ethanol corrected) - CAN/CGSB-3.511 | °C |  |  | 515.2 | 8.5 | 21 | 492 | 532 | 23.498 |  |  | 0.576 | Eric Salvail | 2023-11-10
M | M972 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 11.11 | 1.12 | 19 | 9.4 | 13.5 | 3.091 | 5.051 | 0.375 - Pass | 0.302 | Eric Salvail | 2023-11-10
M | M972 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.86 | 1 | 11 | 0 | 3.4 | 2.757 | 2.387 |  | 1.344 | Eric Salvail | 2023-11-10
M | M972 | Induction Period at 100°C, ASTM D525 | min | 1015 | 0.83 | 838.8 | 212.1 | 4 | 600 | 1020 | 587.647 | 83.875 |  | 0.507 | Eric Salvail | 2023-11-10
M | M972 | Sulfur - ASTM D5453 | mg/kg | 9.7 | -0.53 | 10.3 | 1.12 | 23 | 8.4 | 12.6 | 3.111 | 3.332 | 0.872 - Pass | 0.34 | Eric Salvail | 2023-11-10
M | M972 | Sulfur - ASTM D7039 | mg/kg |  |  | 9.63 | 0.95 | 7 | 8.5 | 10.8 | 2.642 | 2.509 |  | 0.409 | Eric Salvail | 2023-11-10
M | M972 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 0.00014 | 0.0001 | 9 | 0 | 0.0003 | 0 | 0 |  | 0.931 | Eric Salvail | 2023-11-10
M | M972 | Vapor Pressure - ASTM D5191 | kPa | 70.6 | 1.05 | 70.12 | 0.46 | 27 | 69.2 | 71.2 | 1.263 |  |  | 0.307 | Eric Salvail | 2023-11-10
M | M972 | Octane Number, Research - ASTM D2699 |  | 92.2 | 0.53 | 92.09 | 0.22 | 21 | 91.8 | 92.6 | 0.602 | 0.7 | 0.740 - Pass | 0.553 | Eric Salvail | 2023-11-10
M | M972 | Octane Number, Motor - ASTM D2700 |  | 82.2 | -0.65 | 82.46 | 0.4 | 21 | 81.9 | 83.6 | 1.11 | 0.9 | 1.520 - Pass | 0.616 | Eric Salvail | 2023-11-10
M | M972 | Residue - ASTM D86 (Automatic) | Vol.% | 1 | -4.48 | 1.1 | 0.02 | 22 | 1 | 1.1 | 0.059 |  |  | 6.531 | Eric Salvail | 2023-11-10
M | M973 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 130.8 | 11.29 | 124.71 | 0.54 | 22 | 123.1 | 125.5 | 1.494 |  |  | 0.492 | Eric Salvail | 2023-11-24 |  | recheck demandé
M | M973 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 33 | 0.66 | 32.07 | 1.4 | 23 | 29.4 | 35.7 | 3.889 | 4.7 | 0.685 - Pass | 0.223 | Eric Salvail | 2023-11-24
M | M973 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 64.2 | 2.1 | 61.92 | 1.08 | 23 | 60.1 | 64.2 | 3.004 |  |  | 0.494 | Eric Salvail | 2023-11-24 |  | recheck demandé
M | M973 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 104 | 2.34 | 103.07 | 0.4 | 23 | 102.5 | 104 | 1.103 |  |  | 0.554 | Eric Salvail | 2023-11-24 |  | recheck demandé
M | M973 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 185.9 | -0.38 | 186.83 | 2.47 | 21 | 182.7 | 191.6 | 6.848 | 7.1 | 0.930 - Pass | 0.192 | Eric Salvail | 2023-11-24
M | M973 | Recovery - ASTM D86 (Automatic) | Vol.% | 96.7 | -1.49 | 97.7 | 0.67 | 22 | 96.1 | 98.7 | 1.857 |  |  | 0.445 | Eric Salvail | 2023-11-24
M | M973 | Residue - ASTM D86 (Automatic) | Vol.% | 1 | -0.39 | 1.03 | 0.06 | 20 | 1 | 1.2 | 0.177 |  |  | 5.77 | Eric Salvail | 2023-11-24
M | M973 | Loss - ASTM D86 (Automatic) | Vol.% | 2.3 | 2.02 | 1.18 | 0.55 | 23 | 0.3 | 2.3 | 1.535 |  |  | 0.158 | Eric Salvail | 2023-11-24
M | M973 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 35.2 | -0.16 | 35.28 | 0.5 | 23 | 34.1 | 36 | 1.38 |  |  | 0.523 | Eric Salvail | 2023-11-24
M | M973 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 95.8 | -0.19 | 95.84 | 0.19 | 22 | 95.4 | 96.1 | 0.531 |  |  | 0.52 | Eric Salvail | 2023-11-24
M | M973 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 35.74 | 2.26 | 5 | 33.4 | 39.4 | 6.27 |  |  | 0.397 | Eric Salvail | 2023-11-24
M | M973 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 1.14 | 0.18 | 5 | 0.9 | 1.4 | 0.503 |  |  | 0.304 | Eric Salvail | 2023-11-24
M | M973 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 40.4 | 0.08 | 40.33 | 0.89 | 19 | 38.2 | 41.3 | 2.457 | 2.706 | 0.824 - Pass | 1.614 | Eric Salvail | 2023-11-24
M | M973 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.503 | 0 | 0.503 | 0.0135 | 21 | 0.478 | 0.54 | 0.037 | 0.069 | 0.288 - Pass | 0.336 | Eric Salvail | 2023-11-24
M | M973 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 47.7 | 0.9 | 13 | 46 | 49 | 2.389 |  |  | 1.926 | Eric Salvail | 2023-11-24
M | M973 | Density, at 15°C - ASTM D4052 | kg/m³ | 753.4 | 1.71 | 752.82 | 0.34 | 22 | 752.4 | 753.7 | 0.936 | 1.861 | 0.253 - Pass | 0.794 | Eric Salvail | 2023-11-24
M | M973 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 526.3 | 2.2 | 18 | 523 | 530 | 5.99 |  |  | 0.429 | Eric Salvail | 2023-11-24
M | M973 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 0.92 | 0.42 | 15 | 0.4 | 1.6 | 1.176 | 1.865 | 0.398 - Pass | 0.92 | Eric Salvail | 2023-11-24
M | M973 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.23 | 0.32 | 8 | 0 | 0.8 | 0.898 | 1.594 |  | 1.23 | Eric Salvail | 2023-11-24
M | M973 | Induction Period at 100°C, ASTM D525 | min | >960 |  | 784.7 | 224.1 | 3 | 600 | 1034 | 620.785 | 78.467 |  | 0.391 | Eric Salvail | 2023-11-24
M | M973 | Sulfur - ASTM D5453 | mg/kg | 4.7 | -0.56 | 4.96 | 0.46 | 20 | 4.1 | 6.2 | 1.285 | 1.927 | 0.445 - Pass | 0.644 | Eric Salvail | 2023-11-24
M | M973 | Sulfur - ASTM D7039 | mg/kg |  |  | 4.79 | 0.61 | 7 | 3.5 | 5.3 | 1.703 | 1.72 |  | 0.741 | Eric Salvail | 2023-11-24
M | M973 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 0.00011 | 8e-05 | 8 | 0 | 0.0002 | 0 | 0 |  | 0.656 | Eric Salvail | 2023-11-24
M | M973 | Vapor Pressure - ASTM D5191 | kPa | 53.9 | 0.02 | 53.89 | 0.56 | 24 | 52.9 | 55.6 | 1.538 |  |  | 0.449 | Eric Salvail | 2023-11-24
M | M973 | Octane Number, Research - ASTM D2699 |  | 97.2 | 1.51 | 96.88 | 0.21 | 22 | 96.4 | 97.3 | 0.584 | 0.7 | 0.696 - Pass | 0.339 | Eric Salvail | 2023-11-24
M | M973 | Octane Number, Motor - ASTM D2700 |  | 87.2 | -2.54 | 88.17 | 0.38 | 20 | 87.2 | 89.1 | 1.053 | 0.9 | 1.369 - Pass | 0.823 | Eric Salvail | 2023-11-24 |  | reche3ck demandé
M | M973 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 4.7 | 11.59 | 1.17 | 0.3 | 18 | 0.5 | 1.6 | 0.844 | 0.599 | 1.984 - Fail | 0.292 | Eric Salvail | 2023-11-24
M | M974 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 84.6 | -20.31 | 89.63 | 0.25 | 25 | 88.9 | 90.1 | 0.686 |  |  | 0.712 | Martine Carrier | 2023-12-21
M | M974 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 46.5 | -10.15 | 51.85 | 0.53 | 26 | 50.8 | 52.8 | 1.46 |  |  | 0.417 | Martine Carrier | 2023-12-21
M | M974 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 22.6 | -1.01 | 24.16 | 1.54 | 28 | 21.6 | 28.6 | 4.269 | 4.7 | 0.825 - Pass | 0.616 | Martine Carrier | 2023-12-21
M | M974 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 35.1 | -0.67 | 35.96 | 1.28 | 28 | 32.5 | 38.8 | 3.556 |  |  | 0.373 | Martine Carrier | 2023-12-21
M | M974 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 90.9 | -0.26 | 91.07 | 0.64 | 28 | 89.6 | 92.1 | 1.779 |  |  | 0.392 | Martine Carrier | 2023-12-21
M | M974 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 149.6 | -0.51 | 149.91 | 0.61 | 27 | 148.5 | 151.4 | 1.678 |  |  | 0.734 | Martine Carrier | 2023-12-21
M | M974 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 194.6 | -0.87 | 196.28 | 1.94 | 28 | 193.6 | 200.9 | 5.363 | 7.1 | 0.571 - Pass | 0.514 | Martine Carrier | 2023-12-21
M | M974 | Recovery - ASTM D86 (Automatic) | Vol.% | 93.8 | -1.85 | 95.59 | 0.97 | 27 | 93.6 | 96.8 | 2.68 |  |  | 0.824 | Martine Carrier | 2023-12-21
M | M974 | Residue - ASTM D86 (Automatic) | Vol.% | 1 | -1.7 | 1.08 | 0.04 | 24 | 1 | 1.1 | 0.123 |  |  | 5.645 | Martine Carrier | 2023-12-21
M | M974 | Loss - ASTM D86 (Automatic) | Vol.% | 5.2 | 1.83 | 3.36 | 1.01 | 28 | 1.9 | 5.6 | 2.787 |  |  | 0.796 | Martine Carrier | 2023-12-21
M | M974 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 14.81 | 2.1 | 9 | 12.7 | 18.9 | 5.818 |  |  | 0.67 | Martine Carrier | 2023-12-21
M | M974 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 11.27 | 1.4 | 9 | 9.1 | 13.3 | 3.871 |  |  | 0.178 | Martine Carrier | 2023-12-21
M | M974 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 15.8 | 0.69 | 15.45 | 0.51 | 23 | 14.4 | 16.4 | 1.415 | 2.706 | 0.274 - Pass | 0.713 | Martine Carrier | 2023-12-21
M | M974 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 12.4 | 1.78 | 11.48 | 0.52 | 23 | 10.1 | 12.5 | 1.433 | 1.881 | 0.581 - Pass | 0.429 | Martine Carrier | 2023-12-21
M | M974 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.965 | 1.35 | 0.9217 | 0.0321 | 23 | 0.842 | 1.006 | 0.089 | 0.127 | 0.489 - Pass | 0.828 | Martine Carrier | 2023-12-21
M | M974 | Benzene Emissions Number, Winter - Regulation SOR/97-493 |  | 48 | -0.85 | 48.6 | 0.7 | 14 | 47 | 50 | 2.048 |  |  | 1.171 | Martine Carrier | 2023-12-21
M | M974 | Density, at 15°C - ASTM D4052 | kg/m³ | 710.8 | 1.61 | 709.76 | 0.64 | 29 | 707.7 | 711.1 | 1.785 | 3.217 | 0.308 - Pass | 1.111 | Martine Carrier | 2023-12-21
M | M974 | Driveability Index - CAN/CGSB-3.5 | °C | 475 | -1.01 | 477.6 | 2.5 | 21 | 472 | 482 | 7.039 |  |  | 0.227 | Martine Carrier | 2023-12-21
M | M974 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 1.36 | 0.52 | 20 | 0.5 | 2 | 1.442 | 2.18 | 0.437 - Pass | 0.927 | Martine Carrier | 2023-12-21
M | M974 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.3 | 0.38 | 10 | 0 | 1 | 1.045 | 1.738 |  | 0.912 | Martine Carrier | 2023-12-21
M | M974 | Induction Period at 100°C, ASTM D525 | min | 1149 | 1.19 | 830.7 | 268.4 | 6 | 480 | 1149 | 743.533 | 83.067 |  | 0.336 | Martine Carrier | 2023-12-21
M | M974 | Sulfur - ASTM D5453 | mg/kg | 6 | -0.37 | 6.4 | 1.09 | 26 | 3.9 | 8.6 | 3.019 | 2.333 | 1.675 - Pass | 0.26 | Martine Carrier | 2023-12-21
M | M974 | Sulfur - ASTM D7039 | mg/kg |  |  | 7.09 | 1.28 | 9 | 5.4 | 10.1 | 3.554 | 2.126 |  | 0.921 | Martine Carrier | 2023-12-21
M | M974 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 0.11 | 0.00028 | 0.00014 | 19 | 0 | 0.0006 | 0 | 0 | 1.420 - Pass | 1.796 | Martine Carrier | 2023-12-21
M | M974 | Vapor Pressure - ASTM D5191 | kPa | 102 | -0.08 | 102.06 | 0.7 | 28 | 100.4 | 103.2 | 1.929 |  |  | 0.278 | Martine Carrier | 2023-12-21
M | M974 | Octane Number, Research - ASTM D2699 |  | 88.3 | 0.06 | 88.28 | 0.35 | 23 | 87.3 | 88.7 | 0.967 | 0.7 | 1.907 - Pass | 0.705 | Martine Carrier | 2023-12-21
M | M974 | Octane Number, Motor - ASTM D2700 |  | 81.5 | -1.33 | 81.92 | 0.31 | 21 | 81.4 | 82.6 | 0.87 | 0.9 | 0.934 - Pass | 0.375 | Martine Carrier | 2023-12-21
M | M975 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 78.4 | -4.5 | 80.3 | 0.42 | 23 | 79.2 | 81.4 | 1.168 |  |  | 0.84 | Eric Salvail | 2024-02-12
M | M975 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 34.4 | 1.16 | 32.75 | 1.42 | 24 | 30.1 | 36.2 | 3.924 | 4.7 | 0.697 - Pass | 0.281 | Eric Salvail | 2024-02-12
M | M975 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 56.9 | -0.26 | 57.15 | 0.98 | 24 | 55.2 | 59.7 | 2.712 |  |  | 0.562 | Eric Salvail | 2024-02-12
M | M975 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 105.9 | -0.12 | 105.96 | 0.48 | 24 | 105.2 | 107 | 1.337 |  |  | 0.415 | Eric Salvail | 2024-02-12
M | M975 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 172.1 | 0.27 | 171.88 | 0.81 | 24 | 169.7 | 173.7 | 2.248 |  |  | 0.831 | Eric Salvail | 2024-02-12
M | M975 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 218.1 | 0.65 | 217 | 1.7 | 24 | 213 | 219.9 | 4.722 | 7.1 | 0.442 - Pass | 0.472 | Eric Salvail | 2024-02-12
M | M975 | Recovery - ASTM D86 (Automatic) | Vol.% | 97.1 | -1.82 | 98.01 | 0.5 | 23 | 96.8 | 98.7 | 1.385 |  |  | 0.418 | Eric Salvail | 2024-02-12
M | M975 | Residue - ASTM D86 (Automatic) | Vol.% | 1.1 | 0.09 | 1.1 | 0.05 | 23 | 1 | 1.2 | 0.131 |  |  | 4.484 | Eric Salvail | 2024-02-12
M | M975 | Loss - ASTM D86 (Automatic) | Vol.% | 1.8 | 1.86 | 0.83 | 0.52 | 24 | 0 | 2.1 | 1.449 |  |  | 0.341 | Eric Salvail | 2024-02-12
M | M975 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 38 | -2.83 | 39.88 | 0.66 | 24 | 38 | 40.7 | 1.841 |  |  | 1.292 | Eric Salvail | 2024-02-12
M | M975 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 29.88 | 0.56 | 5 | 29.2 | 30.6 | 1.56 |  |  | 0.291 | Eric Salvail | 2024-02-12
M | M975 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 10.74 | 1.81 | 5 | 8.7 | 13.5 | 5.027 |  |  | 0.318 | Eric Salvail | 2024-02-12
M | M975 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 30 | 0.97 | 29.13 | 0.9 | 21 | 26.9 | 30.1 | 2.479 | 2.706 | 0.839 - Pass | 1.196 | Eric Salvail | 2024-02-12
M | M975 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 10.6 | 0.19 | 10.55 | 0.29 | 20 | 10 | 11.1 | 0.791 | 1.802 | 0.193 - Pass | 0.369 | Eric Salvail | 2024-02-12
M | M975 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.912 | 2.02 | 0.855 | 0.0282 | 21 | 0.797 | 0.912 | 0.078 | 0.118 | 0.439 - Pass | 0.204 | Eric Salvail | 2024-02-12 |  | recheck demandé
M | M975 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  | 40 | 0.46 | 39.6 | 0.8 | 15 | 38 | 41 | 2.254 |  |  | 1.613 | Eric Salvail | 2024-02-12
M | M975 | Density, at 15°C - ASTM D4052 | kg/m³ | 748 | 0.52 | 747.81 | 0.36 | 24 | 746.7 | 748.5 | 0.998 | 2.019 | 0.244 - Pass | 0.864 | Eric Salvail | 2024-02-12
M | M975 | Driveability Index - CAN/CGSB-3.5 | °C | 575 | -0.2 | 575.7 | 3.4 | 19 | 570 | 583 | 9.328 |  |  | 0.291 | Eric Salvail | 2024-02-12
M | M975 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 1.95 | 1.31 | 15 | 0 | 4.8 | 3.633 | 2.517 | 2.083 - Fail | 0.427 | Eric Salvail | 2024-02-12
M | M975 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.74 | 0.82 | 7 | 0 | 2.2 | 2.261 | 2.281 |  | 0.528 | Eric Salvail | 2024-02-12
M | M975 | Sulfur - ASTM D5453 | mg/kg | 15.9 | 0.17 | 15.71 | 1.16 | 20 | 13.6 | 17.4 | 3.218 | 4.573 | 0.495 - Pass | 0.45 | Eric Salvail | 2024-02-12
M | M975 | Sulfur - ASTM D7039 | mg/kg |  |  | 15.54 | 0.81 | 7 | 14.5 | 16.5 | 2.232 | 3.249 |  | 0.326 | Eric Salvail | 2024-02-12
M | M975 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | -0.11 | 0.00032 | 0.0002 | 14 | 0 | 0.0007 | 0.001 | 0 | 2.840 - Fail | 0.271 | Eric Salvail | 2024-02-12
M | M975 | Vapor Pressure - ASTM D5191 | kPa | 53.9 | 0.54 | 53.56 | 0.64 | 24 | 52.6 | 55.5 | 1.763 |  |  | 0.502 | Eric Salvail | 2024-02-12
M | M975 | Octane Number, Research - ASTM D2699 |  | 88.7 | 1.84 | 87.98 | 0.39 | 23 | 87.3 | 89.1 | 1.078 | 0.7 | 2.372 - Fail | 0.775 | Eric Salvail | 2024-02-12
M | M975 | Octane Number, Motor - ASTM D2700 |  | 80 | -1.41 | 80.29 | 0.2 | 22 | 79.8 | 80.7 | 0.562 | 0.9 | 0.391 - Pass | 0.438 | Eric Salvail | 2024-02-12
M | M976 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 24.7 | 0.12 | 24.55 | 1.28 | 25 | 21.8 | 27 | 3.546 | 4.7 | 0.569 - Pass | 0.236 | Eric Salvail | 2024-03-11
M | M976 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 38.2 | -0.57 | 39.3 | 1.92 | 25 | 34.6 | 43.6 | 5.321 |  |  | 0.324 | Eric Salvail | 2024-03-11
M | M976 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 97.4 | 0.25 | 97.3 | 0.41 | 25 | 96.6 | 98.2 | 1.132 |  |  | 0.248 | Eric Salvail | 2024-03-11
M | M976 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 138.8 | 0.81 | 137.47 | 1.65 | 25 | 133.7 | 140.6 | 4.557 |  |  | 0.128 | Eric Salvail | 2024-03-11
M | M976 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 203.9 | 1.75 | 196.64 | 4.15 | 24 | 190.4 | 207.9 | 11.484 | 7.1 | 2.616 - Fail | 0.863 | Eric Salvail | 2024-03-11
M | M976 | Recovery - ASTM D86 (Automatic) | Vol.% | 93 | -1.32 | 94.81 | 1.37 | 24 | 92 | 98.5 | 3.789 |  |  | 0.252 | Eric Salvail | 2024-03-11
M | M976 | Loss - ASTM D86 (Automatic) | Vol.% | 6.2 | 1.36 | 4.27 | 1.42 | 25 | 0.7 | 6.9 | 3.923 |  |  | 0.244 | Eric Salvail | 2024-03-11
M | M976 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 41.4 | -1.18 | 42.49 | 0.92 | 24 | 40.6 | 44 | 2.546 |  |  | 0.499 | Eric Salvail | 2024-03-11
M | M976 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 91.9 | -0.3 | 92 | 0.33 | 24 | 91.5 | 92.9 | 0.913 |  |  | 0.417 | Eric Salvail | 2024-03-11
M | M976 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 1.75 | 0.3 | 4 | 1.5 | 2.1 | 0.831 |  |  | 0.494 | Eric Salvail | 2024-03-11
M | M976 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 1.35 | 0.21 | 6 | 1 | 1.6 | 0.574 |  |  | 0.307 | Eric Salvail | 2024-03-11
M | M976 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 4.1 | -0.09 | 4.24 | 1.56 | 21 | 2.5 | 7 | 4.334 | 2.706 | 2.566 - Fail | 0.863 | Eric Salvail | 2024-03-11
M | M976 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 1.2 | -1.43 | 2.38 | 0.82 | 21 | 0.4 | 4.3 | 2.271 | 0.856 | 7.044 - Fail | 0.484 | Eric Salvail | 2024-03-11
M | M976 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.043 | 1.1 | 0.0417 | 0.0011 | 19 | 0.04 | 0.043 | 0.003 | 0.006 | 0.304 - Pass | 1.113 | Eric Salvail | 2024-03-11
M | M976 | Benzene Emissions Number, Winter - Regulation SOR/97-493 |  |  |  | 37.5 | 1.2 | 12 | 35 | 38 | 3.207 |  |  | 3.183 | Eric Salvail | 2024-03-11
M | M976 | Density, at 15°C - ASTM D4052 | kg/m³ | 686.4 | 1.4 | 685.69 | 0.51 | 24 | 685.1 | 686.7 | 1.413 | 3.976 | 0.126 - Pass | 1.174 | Eric Salvail | 2024-03-11
M | M976 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 489 | 4.9 | 17 | 481 | 498 | 13.641 |  |  | 0.764 | Eric Salvail | 2024-03-11
M | M976 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 1.57 | 0.69 | 17 | 0.5 | 3 | 1.916 | 2.31 | 0.688 - Pass | 0.74 | Eric Salvail | 2024-03-11
M | M976 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.73 | 0.55 | 10 | 0 | 2 | 1.523 | 2.269 |  | 1.008 | Eric Salvail | 2024-03-11
M | M976 | Induction Period at 100°C, ASTM D525 | min |  |  | 853.3 | 231.7 | 4 | 600 | 1094 | 641.889 | 85.325 |  | 0.345 | Eric Salvail | 2024-03-11
M | M976 | Sulfur - ASTM D5453 | mg/kg | 6 | 0.91 | 5.34 | 0.73 | 23 | 4.2 | 7.2 | 2.009 | 2.037 | 0.973 - Pass | 0.631 | Eric Salvail | 2024-03-11
M | M976 | Sulfur - ASTM D7039 | mg/kg |  |  | 5.57 | 0.47 | 6 | 4.8 | 6.1 | 1.307 | 1.866 |  | 0.248 | Eric Salvail | 2024-03-11
M | M976 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0001 | 0.33 | 8e-05 | 8e-05 | 12 | 0 | 0.0002 | 0 | 0 |  | 1.08 | Eric Salvail | 2024-03-11
M | M976 | Vapor Pressure - ASTM D5191 | kPa | 99.9 | 1.27 | 98.65 | 0.99 | 24 | 96.3 | 101 | 2.746 |  |  | 0.23 | Eric Salvail | 2024-03-11
M | M976 | Octane Number, Research - ASTM D2699 |  | 93.3 | 0.9 | 93.11 | 0.21 | 22 | 92.6 | 93.6 | 0.585 | 0.7 | 0.700 - Pass | 0.579 | Eric Salvail | 2024-03-11
M | M976 | Octane Number, Motor - ASTM D2700 |  | 90.1 | -0.09 | 90.13 | 0.31 | 21 | 89.5 | 90.8 | 0.868 | 0.9 | 0.930 - Pass | 0.342 | Eric Salvail | 2024-03-11
M | M976 | Residue - ASTM D86 (Automatic) | Vol.% | 0.8 | -7.78 | 1.09 | 0.04 | 20 | 1 | 1.1 | 0.101 |  |  | 6.072 | Eric Salvail | 2024-03-11 |  | recheck demandé
P | P118 | Composition - Ethane - All Methods | Vol % | 0 | -27.69 | 0.05 | 0 | 14 | 0.047 | 0.053 | 0.01 |  |  | 0.7 | Eric Salvail | 2024-02-22
P | P118 | Composition - Ethane - ASTM D2163 | Vol % | 0 | -30.34 | 0.05 | 0 | 27 | 0.047 | 0.054 | 0.01 | 0.05 | 0.010 - Pass | 1.29 | Eric Salvail | 2024-02-22
P | P118 | Composition - Propane - All Methods | Vol % | 92.786 | 1.01 | 92.37 | 0.41 | 14 | 91.814 | 93.33 | 1.13 |  |  | 0.43 | Eric Salvail | 2024-02-22
P | P118 | Composition - C3 - All Methods | Vol % | 92.786 | 1.01 | 92.4 | 0.38 | 16 | 91.864 | 93.38 | 1.06 |  |  | 0.42 | Eric Salvail | 2024-02-22
P | P118 | Composition - Isobutane - All Methods | Vol % | 1.149 | -0.44 | 1.17 | 0.04 | 16 | 1.072 | 1.239 | 0.11 |  |  | 0.5 | Eric Salvail | 2024-02-22
P | P118 | Composition - n-Butane - All Methods | Vol % | 6.066 | -0.9 | 6.36 | 0.33 | 17 | 5.54 | 6.957 | 0.92 |  |  | 0.39 | Eric Salvail | 2024-02-22
P | P118 | Density, LPG, at 15°C - All Methods | kg/m³ | 512.5 | -0.51 | 512.69 | 0.37 | 23 | 511.984 | 513.3 | 1.02 |  |  | 0.64 | Eric Salvail | 2024-02-22
P | P118 |  |  | 2163 | -1.5 | 2489.25 | 217.5 | 4 | 2163 | 2598 | 602.48 |  |  | 1.1 | Eric Salvail | 2024-02-22
P | P118 | Sulfur, LPG - All Methods | ppmw | 11.01 | -1.36 | 15 | 2.9 | 24 | 7.921 | 21.56 | 8.11 |  |  | 0.26 | Eric Salvail | 2024-02-22
P | P118 | Vapor Pressure LPG at 37.8°C - All Methods | kPa | 1133 | 0.15 | 1132.1 | 6 | 25 | 1116.95 | 1147.54 | 16.71 |  |  | 1.45 | Eric Salvail | 2024-02-22
P | P118 | Butane and Heavier - All Methods | Vol. % | 7.214 | -1.03 | 7.58 | 0.35 | 21 | 6.58 | 8.139 | 0.97 |  |  | 0.42 | Eric Salvail | 2024-02-22
P | P118 | Composition - Propene - All Methods | Vol % | 0 | -10.63 | 0.04 | 0 | 11 | 0.04 | 0.05 | 0.01 |  |  | 2.36 | Eric Salvail | 2024-02-22
P | P118 | Composition - Propene - ASTM D2163 | Vol % | 0 | -21.55 | 0.04 | 0 | 27 | 0.038 | 0.047 | 0.01 | 0.04 | 0.023 - Pass | 1.44 | Eric Salvail | 2024-02-22
P | P118 | Composition - Propane - ASTM D2163 | Vol % | 92.786 | 1.58 | 92.26 | 0.33 | 30 | 91.67 | 93.33 | 0.92 | 4.14 | 0.050 - Pass | 0.81 | Eric Salvail | 2024-02-22
P | P118 | Composition - C3 - ASTM D2163 | Vol % | 92.786 | 1.28 | 92.33 | 0.36 | 25 | 91.71 | 93.38 | 1 |  |  | 0.55 | Eric Salvail | 2024-02-22
P | P118 | Composition - Isobutane - ASTM D2163 | Vol % | 1.149 | -0.53 | 1.16 | 0.02 | 29 | 1.113 | 1.22 | 0.07 | 0.18 | 0.153 - Pass | 0.25 | Eric Salvail | 2024-02-22
P | P118 | Composition - n-Butane - ASTM D2163 | Vol % | 6.066 | -1.63 | 6.43 | 0.22 | 30 | 5.938 | 6.98 | 0.62 | 0.32 | 3.769 - Fail | 0.44 | Eric Salvail | 2024-02-22
T | T330 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 151.7 | 0.16 | 151.35 | 2.15 | 17 | 147.8 | 154.6 | 5.968 | 8.324 | 0.514 - Pass | 0.44 | Martine Carrier | 2022-10-13
T | T330 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 174.3 | 0.9 | 173.83 | 0.53 | 16 | 172.9 | 174.4 | 1.461 | 3.824 | 0.146 - Pass | 0.619 | Martine Carrier | 2022-10-13
T | T330 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 205.8 | 0.89 | 205.24 | 0.63 | 16 | 204.3 | 206.3 | 1.734 | 3 | 0.334 - Pass | 0.377 | Martine Carrier | 2022-10-13
T | T330 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 255 | 0.15 | 254.83 | 1.19 | 15 | 253.4 | 257.1 | 3.307 | 3.822 | 0.748 - Pass | 0.784 | Martine Carrier | 2022-10-13
T | T330 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 286.2 | 0.39 | 285.69 | 1.3 | 16 | 282.5 | 287.3 | 3.602 | 7.1 | 0.257 - Pass | 0.416 | Martine Carrier | 2022-10-13
T | T330 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.4 | 0.53 | 98.29 | 0.21 | 16 | 97.9 | 98.6 | 0.589 |  |  | 0.382 | Martine Carrier | 2022-10-13
T | T330 | Residue - ASTM D86 (Automatic) | Vol. % | 1.2 | 0 | 1.2 | 0 | 15 | 1.2 | 1.2 | 0 |  |  |  | Martine Carrier | 2022-10-13
T | T330 | Loss - ASTM D86 (Automatic) | Vol. % | 0.4 | -0.43 | 0.49 | 0.22 | 15 | 0.2 | 0.9 | 0.597 |  |  | 0.619 | Martine Carrier | 2022-10-13
T | T330 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0015 | 0.0007 | 11 | 0.001 | 0.003 | 0.002 | 0.002 |  | 1.675 | Martine Carrier | 2022-10-13
T | T330 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 16.8 | 0.57 | 5 | 16 | 17.5 | 1.591 | 1.825 |  | 0.191 | Martine Carrier | 2022-10-13
T | T330 | Aromatics - ASTM D1319 | Vol. % |  |  | 16.39 | 1.01 | 9 | 15 | 18.4 | 2.806 |  |  | 0.344 | Martine Carrier | 2022-10-13
T | T330 | Olefins - ASTM D1319 | Vol. % |  |  | 1.46 | 0.53 | 10 | 0.9 | 2.2 | 1.461 |  |  | 0.732 | Martine Carrier | 2022-10-13
T | T330 | Density, at 15°C - ASTM D4052 | kg/m³ | 823 | 0.33 | 822.98 | 0.05 | 17 | 822.9 | 823.1 | 0.146 | 0.5 | 0.086 - Pass | 2.857 | Martine Carrier | 2022-10-13
T | T330 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 140 | 0.13 | 130.3 | 73.2 | 12 | 75 | 289 | 202.895 |  |  | 1.761 | Martine Carrier | 2022-10-13
T | T330 | Existent Gum Content - IP540 | mg/100mL |  |  | 3 | 3.1 | 10 | 0 | 10 | 8.606 | 4.871 |  | 0.988 | Martine Carrier | 2022-10-13
T | T330 | Flash Point - ASTM D56 | °C |  |  | 44.84 | 1.02 | 14 | 43.5 | 46.5 | 2.822 | 4.3 | 0.431 - Pass | 0.406 | Martine Carrier | 2022-10-13
T | T330 | Freezing Point - ASTM D5972 | °C | -60.7 | -0.08 | -60.07 | 7.97 | 13 | -66.8 | -43.4 | 22.09 | 0.8 | 762.442 - Fail | 0.987 | Martine Carrier | 2022-10-13
T | T330 | Micro Separometer - ASTM D3948 |  |  |  | 86.4 | 6.6 | 9 | 78 | 97 | 18.38 | 13.719 |  | 0.28 | Martine Carrier | 2022-10-13
T | T330 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 90 | 0.03 | 89.9 | 2.2 | 14 | 86 | 93 | 5.997 | 13.787 | 0.189 - Pass | 0.532 | Martine Carrier | 2022-10-13
T | T330 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.578 | 0.094 | 10 | 0.35 | 0.74 | 0.26 | 0.068 |  | 1.73 | Martine Carrier | 2022-10-13
T | T330 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.0846 | 0.012 | 9 | 43.06 | 43.103 | 0.033 | 0.046 |  | 0.376 | Martine Carrier | 2022-10-13
T | T330 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 20.79 | 1.33 | 13 | 18.3 | 24.1 | 3.677 | 3.445 | 1.140 - Pass | 1.093 | Martine Carrier | 2022-10-13
T | T330 | Sulfur - ASTM D4294 | mg/kg |  |  | 31.03 | 6.06 | 3 | 27 | 38 | 16.782 | 17.56 |  | 0.586 | Martine Carrier | 2022-10-13
T | T330 | Sulfur - ASTM D5453 | mg/kg | 22.1 | -0.88 | 22.75 | 0.75 | 11 | 21.8 | 24 | 2.067 | 6.04 |  | 0.522 | Martine Carrier | 2022-10-13
T | T330 | Sulfur - ASTM D7039 | mg/kg |  |  | 23.02 | 0.94 | 5 | 22 | 23.9 | 2.601 | 4.016 |  | 0.756 | Martine Carrier | 2022-10-13
T | T330 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 1 | 0.0002 | 0.0001 | 7 | 0 | 0.0003 | 0 | 0 |  | 0.891 | Martine Carrier | 2022-10-13
T | T330 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 10 | 0 | 0 | 0 |  |  |  | Martine Carrier | 2022-10-13
T | T330 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.8836 | 0.0153 | 9 | 4.865 | 4.916 | 0.042 | 0.027 |  | 0.376 | Martine Carrier | 2022-10-13
T | T330 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.9102 | 0.0203 | 5 | 4.893 | 4.943 | 0.056 | 0.149 |  | 0.433 | Martine Carrier | 2022-10-13
T | T331 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 167.4 | 1.13 | 164.74 | 2.36 | 18 | 160.8 | 169.8 | 6.534 | 9.061 | 0.520 - Pass | 0.196 | Martine Carrier | 2022-10-17
T | T331 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 187 | 0.4 | 186.7 | 0.76 | 20 | 184.8 | 187.8 | 2.105 | 4.107 | 0.263 - Pass | 0.658 | Martine Carrier | 2022-10-17
T | T331 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 221.1 | 1.14 | 220.16 | 0.83 | 20 | 217.9 | 221.5 | 2.287 | 3 | 0.581 - Pass | 0.391 | Martine Carrier | 2022-10-17
T | T331 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 272 | 1.69 | 270.26 | 1.03 | 20 | 267.9 | 272 | 2.852 | 4.054 | 0.495 - Pass | 0.322 | Martine Carrier | 2022-10-17
T | T331 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 291.4 | 0.4 | 291 | 1 | 20 | 288.4 | 293.3 | 2.779 | 7.1 | 0.153 - Pass | 1.011 | Martine Carrier | 2022-10-17
T | T331 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | -0.43 | 98.31 | 0.25 | 20 | 97.9 | 98.8 | 0.681 |  |  | 0.44 | Martine Carrier | 2022-10-17
T | T331 | Residue - ASTM D86 (Automatic) | Vol. % | 1 | -2.95 | 1.26 | 0.09 | 20 | 1 | 1.4 | 0.244 |  |  | 2.652 | Martine Carrier | 2022-10-17
T | T331 | Loss - ASTM D86 (Automatic) | Vol. % | 0.8 | 1.43 | 0.47 | 0.23 | 20 | 0 | 0.9 | 0.65 |  |  | 0.319 | Martine Carrier | 2022-10-17
T | T331 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0011 | 0.0006 | 15 | 0 | 0.002 | 0.002 | 0.001 | 1.682 - Pass | 1.665 | Martine Carrier | 2022-10-17
T | T331 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 19.91 | 0.95 | 8 | 18.7 | 21.6 | 2.638 | 2.044 |  | 0.238 | Martine Carrier | 2022-10-17
T | T331 | Aromatics - ASTM D1319 | Vol. % |  |  | 18.79 | 1.25 | 11 | 15.9 | 20.2 | 3.463 |  |  | 0.672 | Martine Carrier | 2022-10-17
T | T331 | Olefins - ASTM D1319 | Vol. % |  |  | 1.61 | 1.06 | 11 | 0 | 3.4 | 2.927 |  |  | 0.465 | Martine Carrier | 2022-10-17
T | T331 | Density, at 15°C - ASTM D4052 | kg/m³ | 836.3 | 0.06 | 836.3 | 0.08 | 21 | 836.2 | 836.5 | 0.223 | 0.5 | 0.199 - Pass | 1.656 | Martine Carrier | 2022-10-17
T | T331 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 322 | 0.69 | 295.1 | 38.9 | 15 | 200 | 340 | 107.758 |  |  | 0.852 | Martine Carrier | 2022-10-17
T | T331 | Existent Gum Content - IP540 | mg/100mL |  |  | 6.1 | 4.1 | 16 | 1 | 15 | 11.393 | 7.43 | 2.351 - Fail | 0.659 | Martine Carrier | 2022-10-17
T | T331 | Flash Point - ASTM D56 | °C |  |  | 54.97 | 1.06 | 19 | 53 | 57 | 2.925 | 4.3 | 0.463 - Pass | 0.393 | Martine Carrier | 2022-10-17
T | T331 | Freezing Point - ASTM D5972 | °C | -54.1 | -0.33 | -53.58 | 1.59 | 17 | -55.6 | -50.1 | 4.392 | 0.8 | 30.145 - Fail | 0.993 | Martine Carrier | 2022-10-17
T | T331 | Micro Separometer - ASTM D3948 |  |  |  | 90 | 4.6 | 13 | 78 | 95 | 12.844 | 12.02 | 1.142 - Pass | 0.563 | Martine Carrier | 2022-10-17
T | T331 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 90 | 0.63 | 88.3 | 2.8 | 16 | 85 | 94 | 7.736 | 14.736 | 0.276 - Pass | 0.504 | Martine Carrier | 2022-10-17
T | T331 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.246 | 0.015 | 16 | 0.22 | 0.27 | 0.043 | 0.041 | 1.112 - Pass | 1.201 | Martine Carrier | 2022-10-17
T | T331 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 42.9726 | 0.0143 | 13 | 42.953 | 43.001 | 0.04 | 0.046 | 0.738 - Pass | 0.364 | Martine Carrier | 2022-10-17
T | T331 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 42.9042 | 0.0326 | 5 | 42.88 | 42.96 | 0.09 | 0.035 |  | 0.671 | Martine Carrier | 2022-10-17
T | T331 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 19.44 | 1.08 | 17 | 17 | 21 | 2.98 | 3.318 | 0.807 - Pass | 0.324 | Martine Carrier | 2022-10-17
T | T331 | Sulfur - ASTM D4294 | mg/kg |  |  | 4.97 | 5 | 3 | 0 | 10 | 13.851 | 5.39 |  | 0.284 | Martine Carrier | 2022-10-17
T | T331 | Sulfur - ASTM D5453 | mg/kg | 0.6 | -0.57 | 0.76 | 0.29 | 8 | 0.5 | 1.3 | 0.796 | 0.473 |  | 0.958 | Martine Carrier | 2022-10-17
T | T331 | Sulfur - ASTM D7039 | mg/kg |  |  | 0.83 | 0.38 | 3 | 0.4 | 1.1 | 1.049 | 0.669 |  | 0.537 | Martine Carrier | 2022-10-17
T | T331 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00014 | 0.00022 | 8 | 0 | 0.0006 | 0.001 | 0 |  | 1.184 | Martine Carrier | 2022-10-17
T | T331 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 14 | 0 | 0 | 0 |  |  |  | Martine Carrier | 2022-10-17
T | T331 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 6.8296 | 0.0315 | 13 | 6.78 | 6.887 | 0.087 | 0.043 | 4.169 - Fail | 0.323 | Martine Carrier | 2022-10-17
T | T331 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 6.4558 | 0.9035 | 6 | 4.612 | 6.855 | 2.503 | 0.149 |  | 1.719 | Martine Carrier | 2022-10-17
T | T332 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 171.3 | -0.32 | 171.97 | 2.13 | 18 | 167.5 | 175.6 | 5.893 | 9.458 | 0.388 - Pass | 0.542 | Martine Carrier | 2022-12-09
T | T332 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 184.8 | 0.41 | 184.52 | 0.68 | 18 | 183 | 185.7 | 1.897 | 4.059 | 0.218 - Pass | 0.194 | Martine Carrier | 2022-12-09
T | T332 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 200.8 | -0.17 | 200.9 | 0.57 | 18 | 199.4 | 201.9 | 1.59 | 3 | 0.281 - Pass | 0.784 | Martine Carrier | 2022-12-09
T | T332 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 236.7 | -0.92 | 237.37 | 0.72 | 18 | 236.5 | 238.8 | 2 | 3.561 | 0.315 - Pass | 0.651 | Martine Carrier | 2022-12-09
T | T332 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 274.6 | 0.6 | 273.44 | 1.94 | 18 | 270.4 | 278.2 | 5.376 | 7.1 | 0.573 - Pass | 0.393 | Martine Carrier | 2022-12-09
T | T332 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.8 | 1.35 | 98.39 | 0.3 | 18 | 97.7 | 98.9 | 0.844 |  |  | 0.383 | Martine Carrier | 2022-12-09
T | T332 | Residue - ASTM D86 (Automatic) | Vol. % | 1.2 | 0 | 1.2 | 0.04 | 16 | 1.1 | 1.3 | 0.101 |  |  | 4.373 | Martine Carrier | 2022-12-09
T | T332 | Loss - ASTM D86 (Automatic) | Vol. % | 0 | -1.47 | 0.39 | 0.27 | 18 | 0 | 0.8 | 0.735 |  |  | 0.444 | Martine Carrier | 2022-12-09
T | T332 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0011 | 0.0005 | 12 | 0 | 0.002 | 0.001 | 0.001 |  | 2.195 | Martine Carrier | 2022-12-09
T | T332 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 17.35 | 0.91 | 8 | 15.6 | 18.6 | 2.534 | 1.865 |  | 0.482 | Martine Carrier | 2022-12-09
T | T332 | Aromatics - ASTM D1319 | Vol. % |  |  | 16.84 | 0.98 | 9 | 14.6 | 17.8 | 2.718 |  |  | 0.609 | Martine Carrier | 2022-12-09
T | T332 | Olefins - ASTM D1319 | Vol. % |  |  | 1.61 | 0.68 | 9 | 0.7 | 3 | 1.886 |  |  | 0.317 | Martine Carrier | 2022-12-09
T | T332 | Density, at 15°C - ASTM D4052 | kg/m³ | 813.7 | -0.92 | 813.78 | 0.08 | 17 | 813.7 | 814 | 0.23 | 0.5 | 0.212 - Pass | 1.606 | Martine Carrier | 2022-12-09
T | T332 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 105 | 0.91 | 87.6 | 19 | 13 | 55 | 130 | 52.709 |  |  | 0.326 | Martine Carrier | 2022-12-09
T | T332 | Existent Gum Content - IP540 | mg/100mL |  |  | 0.6 | 0.5 | 8 | 0 | 1 | 1.371 | 1.797 |  | 1.091 | Martine Carrier | 2022-12-09
T | T332 | Flash Point - ASTM D56 | °C |  |  | 55.39 | 0.88 | 16 | 54 | 56.5 | 2.437 | 4.3 | 0.321 - Pass | 0.634 | Martine Carrier | 2022-12-09
T | T332 | Freezing Point - ASTM D5972 | °C | -64.4 | 0.44 | -64.56 | 0.36 | 15 | -65.3 | -63.7 | 1.009 | 0.8 | 1.589 - Pass | 0.524 | Martine Carrier | 2022-12-09
T | T332 | Micro Separometer - ASTM D3948 |  |  |  | 88.1 | 7 | 8 | 75 | 100 | 19.386 | 12.928 |  | 0.454 | Martine Carrier | 2022-12-09
T | T332 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 91 | -0.94 | 93.3 | 2.4 | 14 | 87 | 96 | 6.735 | 11.74 | 0.329 - Pass | 0.598 | Martine Carrier | 2022-12-09
T | T332 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.692 | 0.017 | 14 | 0.67 | 0.72 | 0.046 | 0.075 | 0.378 - Pass | 0.479 | Martine Carrier | 2022-12-09
T | T332 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.1406 | 0.0119 | 11 | 43.125 | 43.166 | 0.033 | 0.046 |  | 0.253 | Martine Carrier | 2022-12-09
T | T332 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 43.124 | 0.0034 | 4 | 43.12 | 43.128 | 0.009 | 0.035 |  | 0.228 | Martine Carrier | 2022-12-09
T | T332 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 21.71 | 0.58 | 12 | 20.4 | 22.5 | 1.617 | 3.531 |  | 0.611 | Martine Carrier | 2022-12-09
T | T332 | Sulfur - ASTM D4294 | mg/kg | 627 | -1.5 | 668.34 | 27.6 | 8 | 627 | 704 | 76.465 | 127.024 |  | 0.412 | Martine Carrier | 2022-12-09
T | T332 | Sulfur - ASTM D5453 | mg/kg |  |  | 629.26 | 43.13 | 8 | 585.5 | 705.7 | 119.481 | 79.728 |  | 0.617 | Martine Carrier | 2022-12-09
T | T332 | Sulfur - ASTM D7039 | mg/kg |  |  | 641.74 | 13.51 | 5 | 621.4 | 652.9 | 37.436 | 24.225 |  | 0.527 | Martine Carrier | 2022-12-09
T | T332 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0004 | -0.64 | 0.00046 | 0.0001 | 13 | 0.0004 | 0.0007 | 0 | 0 | 0.653 - Pass | 1.79 | Martine Carrier | 2022-12-09
T | T332 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 12 | 0 | 0 | 0 |  |  |  | Martine Carrier | 2022-12-09
T | T332 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.5944 | 0.0227 | 8 | 4.55 | 4.625 | 0.063 | 0.025 |  | 0.301 | Martine Carrier | 2022-12-09
T | T332 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.5964 | 0.0224 | 7 | 4.572 | 4.631 | 0.062 | 0.149 |  | 0.534 | Martine Carrier | 2022-12-09
T | T333 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C |  |  | 152.87 | 3.07 | 19 | 145.2 | 159.4 | 8.504 | 8.408 | 1.023 - Pass | 0.312 | Eric Salvail | 2023-02-13
T | T333 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C |  |  | 177.67 | 0.6 | 19 | 176.4 | 178.8 | 1.67 | 3.909 | 0.183 - Pass | 0.35 | Eric Salvail | 2023-02-13
T | T333 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C |  |  | 205.02 | 0.35 | 18 | 204.4 | 205.6 | 0.962 | 3 | 0.103 - Pass | 0.232 | Eric Salvail | 2023-02-13
T | T333 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C |  |  | 241.93 | 0.62 | 18 | 240.8 | 243 | 1.721 | 3.629 | 0.225 - Pass | 0.184 | Eric Salvail | 2023-02-13
T | T333 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C |  |  | 265.63 | 1.41 | 19 | 262.8 | 268.8 | 3.907 | 7.1 | 0.303 - Pass | 0.332 | Eric Salvail | 2023-02-13
T | T333 | Recovery - ASTM D86 (Automatic) | Vol. % |  |  | 98.31 | 0.32 | 19 | 97.4 | 98.9 | 0.89 |  |  | 1.532 | Eric Salvail | 2023-02-13
T | T333 | Residue - ASTM D86 (Automatic) | Vol. % |  |  | 1.22 | 0.05 | 19 | 1.1 | 1.3 | 0.139 |  |  | 3.503 | Eric Salvail | 2023-02-13
T | T333 | Loss - ASTM D86 (Automatic) | Vol. % |  |  | 0.41 | 0.22 | 18 | 0 | 0.9 | 0.6 |  |  | 0.943 | Eric Salvail | 2023-02-13
T | T333 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0011 | 0.0005 | 15 | 0 | 0.002 | 0.001 | 0.001 | 1.095 - Pass | 2.672 | Eric Salvail | 2023-02-13
T | T333 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 19.74 | 0.69 | 10 | 18.4 | 20.8 | 1.916 | 2.033 |  | 0.256 | Eric Salvail | 2023-02-13
T | T333 | Aromatics - ASTM D1319 | Vol. % |  |  | 19.76 | 0.86 | 9 | 18.3 | 21.1 | 2.383 |  |  | 0.161 | Eric Salvail | 2023-02-13
T | T333 | Olefins - ASTM D1319 | Vol. % |  |  | 1.19 | 0.43 | 9 | 0.7 | 2 | 1.187 |  |  | 0.517 | Eric Salvail | 2023-02-13
T | T333 | Density, at 15°C - ASTM D4052 | kg/m³ | 810.3 | 0.66 | 810.26 | 0.07 | 20 | 810.2 | 810.4 | 0.19 | 0.5 | 0.145 - Pass | 2.412 | Eric Salvail | 2023-02-13
T | T333 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 230 | 2.16 | 161.3 | 31.8 | 14 | 117 | 230 | 88.162 |  |  | 0.413 | Eric Salvail | 2023-02-13
T | T333 | Existent Gum Content - IP540 | mg/100mL |  |  | 1.2 | 0.7 | 12 | 0 | 3 | 1.999 | 2.806 |  | 1.674 | Eric Salvail | 2023-02-13
T | T333 | Flash Point - ASTM D56 | °C |  |  | 45.73 | 0.97 | 18 | 43.5 | 47.5 | 2.681 | 4.3 | 0.389 - Pass | 0.559 | Eric Salvail | 2023-02-13
T | T333 | Freezing Point - ASTM D5972 | °C |  |  | -52.59 | 0.7 | 19 | -53.8 | -51.2 | 1.95 | 0.8 | 5.940 - Fail | 0.327 | Eric Salvail | 2023-02-13
T | T333 | Micro Separometer - ASTM D3948 |  |  |  | 71.4 | 9.7 | 11 | 56 | 87 | 26.749 | 19.8 |  | 0.183 | Eric Salvail | 2023-02-13
T | T333 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 91 | 0.69 | 89 | 2.9 | 17 | 83 | 94 | 8.076 | 14.318 | 0.318 - Pass | 0.303 | Eric Salvail | 2023-02-13
T | T333 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 1.096 | 0.058 | 16 | 0.99 | 1.26 | 0.162 | 0.099 | 2.646 - Fail | 1.662 | Eric Salvail | 2023-02-13
T | T333 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.1474 | 0.0131 | 13 | 43.129 | 43.176 | 0.036 | 0.046 | 0.625 - Pass | 0.219 | Eric Salvail | 2023-02-13
T | T333 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 43.1713 | 0.041 | 4 | 43.142 | 43.232 | 0.114 | 0.035 |  | 0.762 | Eric Salvail | 2023-02-13
T | T333 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 21.45 | 0.78 | 17 | 20.2 | 22.8 | 2.161 | 3.506 | 0.380 - Pass | 0.724 | Eric Salvail | 2023-02-13
T | T333 | Sulfur - ASTM D4294 | mg/kg |  |  | 660.91 | 25.58 | 9 | 629.9 | 695 | 70.852 | 126.112 |  | 0.483 | Eric Salvail | 2023-02-13
T | T333 | Sulfur - ASTM D5453 | mg/kg | 547.1 | -1.42 | 618.98 | 50.56 | 9 | 547.1 | 694.6 | 140.061 | 78.424 |  | 0.29 | Eric Salvail | 2023-02-13
T | T333 | Sulfur - ASTM D7039 | mg/kg |  |  | 647.08 | 13.44 | 5 | 636.6 | 667 | 37.239 | 24.334 |  | 0.576 | Eric Salvail | 2023-02-13
T | T333 | Sulfur - All Methods | Mass % |  |  | 0.0558 | 0.0228 | 4 | 0.0219 | 0.0704 | 0.063 |  |  | 0.816 | Eric Salvail | 2023-02-13
T | T333 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.002 | 0.66 | 0.00196 | 6e-05 | 14 | 0.0018 | 0.002 | 0 | 0 | 0.208 - Pass | 2.173 | Eric Salvail | 2023-02-13
T | T333 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 13 | 0 | 0 | 0 |  |  |  | Eric Salvail | 2023-02-13
T | T333 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.5099 | 0.0256 | 9 | 4.47 | 4.55 | 0.071 | 0.024 |  | 0.189 | Eric Salvail | 2023-02-13
T | T333 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.4986 | 0.0295 | 8 | 4.443 | 4.549 | 0.082 | 0.149 |  | 0.55 | Eric Salvail | 2023-02-13
T | T334 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C |  |  | 158.96 | 3.19 | 12 | 153.3 | 164 | 8.828 | 8.743 |  | 0.184
T | T334 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C |  |  | 184.33 | 0.76 | 12 | 183.3 | 185.5 | 2.094 | 4.055 |  | 0.302
T | T334 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C |  |  | 224.36 | 0.67 | 12 | 223.3 | 225.9 | 1.866 | 3 |  | 0.375
T | T334 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C |  |  | 273.75 | 0.87 | 12 | 272.8 | 276 | 2.413 | 4.106 |  | 0.822
T | T334 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C |  |  | 293.95 | 0.75 | 12 | 292.7 | 295.1 | 2.088 | 7.1 |  | 0.286
T | T334 | Recovery - ASTM D86 (Automatic) | Vol. % |  |  | 98.43 | 0.24 | 12 | 98 | 99 | 0.672 |  |  | 0.576
T | T334 | Residue - ASTM D86 (Automatic) | Vol. % |  |  | 1.3 | 0 | 11 | 1.3 | 1.3 | 0
T | T334 | Loss - ASTM D86 (Automatic) | Vol. % |  |  | 0.3 | 0.2 | 12 | 0 | 0.7 | 0.541 |  |  | 0.252
T | T334 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0009 | 0.0006 | 8 | 0 | 0.002 | 0.002 | 0.001 |  | 1
T | T334 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 22.54 | 0.44 | 5 | 22.1 | 23.1 | 1.217 | 2.221 |  | 0.381
T | T334 | Aromatics - ASTM D1319 | Vol. % |  |  | 21.73 | 0.61 | 4 | 21.1 | 22.5 | 1.698 |  |  | 0.255
T | T334 | Olefins - ASTM D1319 | Vol. % |  |  | 1.5 | 0.51 | 4 | 0.8 | 2 | 1.412 |  |  | 0.331
T | T334 | Density, at 15°C - ASTM D4052 | kg/m³ |  |  | 838.77 | 0.07 | 12 | 838.7 | 838.9 | 0.18 | 0.5 |  | 1.344
T | T334 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m |  |  | 325.5 | 44 | 10 | 265 | 386 | 121.924 |  |  | 0.365
T | T334 | Existent Gum Content - IP540 | mg/100mL |  |  | 3.8 | 1.5 | 6 | 2 | 6 | 4.077 | 5.609 |  | 0.262
T | T334 | Flash Point - ASTM D56 | °C |  |  | 52.34 | 1.47 | 11 | 49.5 | 54.7 | 4.066 | 4.3 |  | 0.293
T | T334 | Freezing Point - ASTM D5972 | °C |  |  | -51.73 | 0.64 | 11 | -53.3 | -51.1 | 1.77 | 0.8 |  | 0.647
T | T334 | Micro Separometer - ASTM D3948 |  |  |  | 84.6 | 7 | 7 | 75 | 92 | 19.314 | 14.573 |  | 0.557
T | T334 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating |  |  | 90 | 3.4 | 9 | 85 | 96 | 9.291 | 13.746 |  | 0.323
T | T334 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.39 | 0.04 | 8 | 0.35 | 0.45 | 0.11 | 0.053 |  | 0.681
T | T334 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 42.9177 | 0.0076 | 7 | 42.905 | 42.927 | 0.021 | 0.046 |  | 0.239
T | T334 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 19.03 | 1.4 | 8 | 17.9 | 21.9 | 3.886 | 3.279 |  | 0.745
T | T334 | Sulfur - ASTM D5453 | mg/kg |  |  | 0.44 | 0.21 | 5 | 0.1 | 0.6 | 0.574 | 0.313 |  | 0.527
T | T334 | Sulfur - ASTM D7039 | mg/kg |  |  | 0.2 | 0.08 | 4 | 0.1 | 0.3 | 0.226 | 0.31 |  | 0.376
T | T334 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 0.00054 | 0.00099 | 5 | 0 | 0.0023 | 0.003 | 0 |  | 1.252
T | T334 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0.02 | 0.04 | 5 | 0 | 0.1 | 0.124 |  |  | 1.498
T | T334 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 7.025 | 0.0174 | 6 | 7.005 | 7.05 | 0.048 | 0.044 |  | 0.289
T | T334 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 7.013 | 0.0026 | 3 | 7.01 | 7.015 | 0.007 | 0.149 |  | 0.468
T | T335 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 151.2 | 0.91 | 148.96 | 2.45 | 18 | 143.9 | 152 | 6.783 | 8.193 | 0.685 - Pass | 0.554 | Eric Salvail | 2023-04-21
T | T335 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 173.4 | 1.09 | 172.46 | 0.86 | 19 | 171.2 | 173.9 | 2.372 | 3.794 | 0.391 - Pass | 0.591 | Eric Salvail | 2023-04-21
T | T335 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 206.7 | 1.85 | 205.48 | 0.66 | 19 | 204.5 | 206.7 | 1.822 | 3 | 0.369 - Pass | 0.778 | Eric Salvail | 2023-04-21
T | T335 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 262.3 | 1.08 | 260.64 | 1.54 | 19 | 258.5 | 264.2 | 4.258 | 3.91 | 1.186 - Pass | 0.312 | Eric Salvail | 2023-04-21
T | T335 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 295.8 | 1.02 | 294.16 | 1.61 | 19 | 291.4 | 296.9 | 4.454 | 7.1 | 0.394 - Pass | 0.215 | Eric Salvail | 2023-04-21
T | T335 | Recovery - ASTM D86 (Automatic) | Vol. % | 98 | -0.98 | 98.19 | 0.2 | 17 | 97.8 | 98.7 | 0.549 |  |  | 0.836 | Eric Salvail | 2023-04-21
T | T335 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -0.98 | 1.21 | 0.11 | 19 | 0.9 | 1.3 | 0.299 |  |  | 1.835 | Eric Salvail | 2023-04-21
T | T335 | Loss - ASTM D86 (Automatic) | Vol. % | 0.9 | 0.91 | 0.6 | 0.33 | 19 | 0 | 1.5 | 0.909 |  |  | 0.631 | Eric Salvail | 2023-04-21
T | T335 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0014 | 0.0009 | 17 | 0 | 0.004 | 0.003 | 0.001 | 2.985 - Fail | 2.811 | Eric Salvail | 2023-04-21
T | T335 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 13.36 | 0.41 | 10 | 12.8 | 14.1 | 1.147 | 1.567 |  | 0.505 | Eric Salvail | 2023-04-21
T | T335 | Aromatics - ASTM D1319 | Vol. % |  |  | 12.26 | 0.82 | 7 | 11.1 | 13.4 | 2.267 |  |  | 0.397 | Eric Salvail | 2023-04-21
T | T335 | Olefins - ASTM D1319 | Vol. % |  |  | 1.46 | 0.72 | 8 | 0.7 | 2.9 | 1.997 |  |  | 0.483 | Eric Salvail | 2023-04-21
T | T335 | Density, at 15°C - ASTM D4052 | kg/m³ | 821.5 | -0.39 | 821.53 | 0.06 | 20 | 821.4 | 821.7 | 0.177 | 0.5 | 0.125 - Pass | 3.073 | Eric Salvail | 2023-04-21
T | T335 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 444 | 0.06 | 440.4 | 64.5 | 14 | 303 | 524 | 178.649 |  |  | 0.61 | Eric Salvail | 2023-04-21
T | T335 | Existent Gum Content - IP540 | mg/100mL |  |  | 1.8 | 0.6 | 11 | 1 | 3 | 1.67 | 3.585 |  | 1.463 | Eric Salvail | 2023-04-21
T | T335 | Flash Point - ASTM D56 | °C |  |  | 44.36 | 0.84 | 17 | 43 | 46 | 2.334 | 4.3 | 0.295 - Pass | 0.362 | Eric Salvail | 2023-04-21
T | T335 | Freezing Point - ASTM D5972 | °C | -63 | -0.17 | -62.48 | 3.05 | 17 | -67.4 | -57.4 | 8.454 | 0.8 | 111.670 - Fail | 0.331 | Eric Salvail | 2023-04-21
T | T335 | Micro Separometer - ASTM D3948 |  |  |  | 77.5 | 11.3 | 11 | 65 | 99 | 31.25 | 17.527 |  | 0.391 | Eric Salvail | 2023-04-21
T | T335 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 88 | 0.1 | 87.7 | 3.1 | 16 | 82 | 93 | 8.625 | 15.042 | 0.329 - Pass | 0.505 | Eric Salvail | 2023-04-21
T | T335 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.513 | 0.014 | 15 | 0.48 | 0.53 | 0.04 | 0.063 | 0.400 - Pass | 1.026 | Eric Salvail | 2023-04-21
T | T335 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.1629 | 0.0111 | 11 | 43.147 | 43.188 | 0.031 | 0.046 |  | 0.964 | Eric Salvail | 2023-04-21
T | T335 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 42.8744 | 0.5093 | 5 | 41.965 | 43.155 | 1.411 | 0.035 |  | 1.315 | Eric Salvail | 2023-04-21
T | T335 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 22.13 | 1.06 | 16 | 20 | 23.9 | 2.947 | 3.57 | 0.681 - Pass | 0.773 | Eric Salvail | 2023-04-21
T | T335 | Sulfur - ASTM D4294 | mg/kg |  |  | 19.9 | 12.22 | 5 | 2.5 | 33.7 | 33.851 | 13.187 |  | 0.293 | Eric Salvail | 2023-04-21
T | T335 | Sulfur - ASTM D5453 | mg/kg | 15.4 | 0.17 | 15.17 | 1.33 | 11 | 12.1 | 16.8 | 3.679 | 4.457 |  | 0.365 | Eric Salvail | 2023-04-21
T | T335 | Sulfur - ASTM D7039 | mg/kg |  |  | 14.85 | 0.45 | 6 | 14 | 15.3 | 1.236 | 3.17 |  | 0.809 | Eric Salvail | 2023-04-21
T | T335 | Sulfur - All Methods | Mass % |  |  | 0.0019 | 0.0016 | 4 | 0.0002 | 0.004 | 0.004 |  |  | 0.281 | Eric Salvail | 2023-04-21
T | T335 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | 0.76 | 0.00022 | 0.00011 | 11 | 0 | 0.0004 | 0 | 0 |  | 0.601 | Eric Salvail | 2023-04-21
T | T335 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 12 | 0 | 0 | 0 |  |  |  | Eric Salvail | 2023-04-21
T | T335 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 5.0826 | 0.0271 | 8 | 5.049 | 5.135 | 0.075 | 0.028 |  | 0.306 | Eric Salvail | 2023-04-21
T | T335 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 5.081 | 0.0879 | 8 | 4.915 | 5.24 | 0.243 | 0.149 |  | 0.939 | Eric Salvail | 2023-04-21
T | T336 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 147.7 | 0.4 | 146.61 | 2.76 | 23 | 141.7 | 153.2 | 7.634 | 8.063 | 0.896 - Pass | 0.212 | Eric Salvail | 2023-05-08
T | T336 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 169.9 | 1.63 | 168.38 | 0.93 | 23 | 166.2 | 169.9 | 2.581 | 3.704 | 0.485 - Pass | 0.232 | Eric Salvail | 2023-05-08
T | T336 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 213.5 | 1.74 | 212.37 | 0.65 | 22 | 210.8 | 213.5 | 1.8 | 3 | 0.360 - Pass | 0.226 | Eric Salvail | 2023-05-08
T | T336 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 276.4 | 0.06 | 276.34 | 0.92 | 22 | 274.9 | 278 | 2.538 | 4.145 | 0.375 - Pass | 0.332 | Eric Salvail | 2023-05-08
T | T336 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 296.5 | 0.6 | 296.01 | 0.82 | 22 | 294.3 | 297.8 | 2.274 | 7.1 | 0.103 - Pass | 0.376 | Eric Salvail | 2023-05-08
T | T336 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.9 | 1.76 | 98.21 | 0.39 | 22 | 97.4 | 98.9 | 1.084 |  |  | 0.491 | Eric Salvail | 2023-05-08
T | T336 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -1.66 | 1.27 | 0.1 | 23 | 1 | 1.5 | 0.282 |  |  | 2.644 | Eric Salvail | 2023-05-08
T | T336 | Loss - ASTM D86 (Automatic) | Vol. % | 0 | -1.27 | 0.5 | 0.39 | 23 | 0 | 1.3 | 1.089 |  |  | 0.711 | Eric Salvail | 2023-05-08
T | T336 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0011 | 0.0009 | 17 | 0 | 0.003 | 0.002 | 0.001 | 3.062 - Fail | 1.073 | Eric Salvail | 2023-05-08
T | T336 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 19.38 | 0.72 | 11 | 18.3 | 20.5 | 1.989 | 2.008 |  | 0.278 | Eric Salvail | 2023-05-08
T | T336 | Aromatics - ASTM D1319 | Vol. % |  |  | 19.23 | 0.66 | 8 | 18.4 | 20.2 | 1.83 |  |  | 0.358 | Eric Salvail | 2023-05-08
T | T336 | Olefins - ASTM D1319 | Vol. % |  |  | 1.69 | 1.02 | 10 | 0.8 | 4.1 | 2.835 |  |  | 0.702 | Eric Salvail | 2023-05-08
T | T336 | Density, at 15°C - ASTM D4052 | kg/m³ | 827.7 | -1.22 | 827.8 | 0.08 | 22 | 827.7 | 828 | 0.218 | 0.5 | 0.189 - Pass | 1.836 | Eric Salvail | 2023-05-08
T | T336 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 318 | 0.13 | 313.1 | 36.7 | 17 | 261 | 409 | 101.731 |  |  | 0.54 | Eric Salvail | 2023-05-08
T | T336 | Existent Gum Content - IP540 | mg/100mL |  |  | 6.5 | 5.2 | 21 | 1 | 14 | 14.356 | 7.683 | 3.492 - Fail | 1.768 | Eric Salvail | 2023-05-08
T | T336 | Flash Point - ASTM D56 | °C |  |  | 43.96 | 1.08 | 22 | 42.5 | 46.5 | 3.005 | 4.3 | 0.488 - Pass | 0.616 | Eric Salvail | 2023-05-08
T | T336 | Freezing Point - ASTM D5972 | °C | -51.4 | 0.25 | -51.71 | 1.23 | 21 | -54.1 | -48 | 3.416 | 0.8 | 18.234 - Fail | 1.076 | Eric Salvail | 2023-05-08
T | T336 | Micro Separometer - ASTM D3948 |  |  |  | 90.8 | 2.8 | 13 | 85 | 95 | 7.657 | 11.604 | 0.435 - Pass | 0.309 | Eric Salvail | 2023-05-08
T | T336 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 92 | 0.69 | 90.5 | 2.2 | 18 | 87 | 94 | 6.028 | 13.453 | 0.201 - Pass | 0.357 | Eric Salvail | 2023-05-08
T | T336 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.365 | 0.005 | 15 | 0.36 | 0.37 | 0.014 | 0.051 | 0.078 - Pass | 2.688 | Eric Salvail | 2023-05-08
T | T336 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.0304 | 0.0108 | 14 | 43.012 | 43.046 | 0.03 | 0.046 | 0.420 - Pass | 0.446 | Eric Salvail | 2023-05-08
T | T336 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 42.9638 | 0.0067 | 4 | 42.957 | 42.973 | 0.019 | 0.035 |  | 0.395 | Eric Salvail | 2023-05-08
T | T336 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 19.59 | 0.7 | 18 | 18.5 | 21 | 1.939 | 3.333 | 0.338 - Pass | 0.356 | Eric Salvail | 2023-05-08
T | T336 | Sulfur - ASTM D4294 | mg/kg |  |  | 4.52 | 9.76 | 6 | 0 | 24.3 | 27.025 | 5.07 |  | 1.575 | Eric Salvail | 2023-05-08
T | T336 | Sulfur - ASTM D5453 | mg/kg | 0.8 | 0.72 | 0.73 | 0.1 | 8 | 0.6 | 0.9 | 0.287 | 0.455 |  | 0.406 | Eric Salvail | 2023-05-08
T | T336 | Sulfur - ASTM D7039 | mg/kg |  |  | 0.68 | 0.34 | 4 | 0.2 | 1 | 0.943 | 0.597 |  | 0.364 | Eric Salvail | 2023-05-08
T | T336 | Sulfur - All Methods | Mass % |  |  | 0.0009 | 0.0011 | 6 | 0 | 0.0025 | 0.003 |  |  | 0.536 | Eric Salvail | 2023-05-08
T | T336 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00021 | 0.00022 | 7 | 0 | 0.0005 | 0.001 | 0 |  | 0.588 | Eric Salvail | 2023-05-08
T | T336 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0.02 | 0.04 | 16 | 0 | 0.1 | 0.112 |  |  | 4.509 | Eric Salvail | 2023-05-08
T | T336 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 5.6389 | 0.044 | 11 | 5.599 | 5.722 | 0.122 | 0.033 |  | 1.163 | Eric Salvail | 2023-05-08
T | T336 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 5.6321 | 0.0213 | 7 | 5.597 | 5.663 | 0.059 | 0.149 |  | 0.394 | Eric Salvail | 2023-05-08
T | T337 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 155.5 | 0.58 | 153.76 | 2.99 | 19 | 148.9 | 158.3 | 8.291 | 8.457 | 0.961 - Pass | 0.315 | Eric Salvail | 2023-08-23
T | T337 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 183 | 0.45 | 182.68 | 0.71 | 19 | 181.4 | 183.9 | 1.975 | 4.019 | 0.242 - Pass | 0.24 | Eric Salvail | 2023-08-23
T | T337 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 207.5 | 0.34 | 207.32 | 0.54 | 19 | 205.8 | 208 | 1.484 | 3 | 0.245 - Pass | 0.535 | Eric Salvail | 2023-08-23
T | T337 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 236.4 | 0.57 | 236.09 | 0.53 | 19 | 235 | 236.9 | 1.473 | 3.541 | 0.173 - Pass | 0.285 | Eric Salvail | 2023-08-23
T | T337 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 250 | 0.35 | 249.53 | 1.36 | 19 | 246.7 | 253 | 3.762 | 7.1 | 0.281 - Pass | 0.36 | Eric Salvail | 2023-08-23
T | T337 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.4 | 0.15 | 98.36 | 0.29 | 19 | 97.7 | 98.9 | 0.801 |  |  | 0.359 | Eric Salvail | 2023-08-23
T | T337 | Residue - ASTM D86 (Automatic) | Vol. % | 1 | -2.23 | 1.16 | 0.07 | 17 | 1 | 1.2 | 0.197 |  |  | 3.155 | Eric Salvail | 2023-08-23
T | T337 | Loss - ASTM D86 (Automatic) | Vol. % | 0.6 | 0.34 | 0.5 | 0.29 | 19 | 0 | 1.2 | 0.815 |  |  | 0.525 | Eric Salvail | 2023-08-23
T | T337 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0014 | 0.0005 | 14 | 0.001 | 0.002 | 0.001 | 0.001 | 0.848 - Pass | 2.806 | Eric Salvail | 2023-08-23
T | T337 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 18.36 | 1 | 9 | 16.3 | 19.7 | 2.777 | 1.936 |  | 0.386 | Eric Salvail | 2023-08-23
T | T337 | Aromatics - ASTM D1319 | Vol. % |  |  | 17.97 | 0.93 | 10 | 16.2 | 19 | 2.582 |  |  | 0.664 | Eric Salvail | 2023-08-23
T | T337 | Olefins - ASTM D1319 | Vol. % |  |  | 1.59 | 1.01 | 10 | 0.5 | 4 | 2.794 |  |  | 0.842 | Eric Salvail | 2023-08-23
T | T337 | Density, at 15°C - ASTM D4052 | kg/m³ | 812.8 | 0.11 | 812.79 | 0.06 | 19 | 812.66 | 812.92 | 0.164 | 0.5 | 0.108 - Pass | 3.03 | Eric Salvail | 2023-08-23
T | T337 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 141 | 0.24 | 135.2 | 24.1 | 13 | 96 | 185 | 66.856 |  |  | 0.845 | Eric Salvail | 2023-08-23
T | T337 | Existent Gum Content - IP540 | mg/100mL |  |  | 0.9 | 0.9 | 10 | 0 | 3 | 2.464 | 2.413 |  | 0.961 | Eric Salvail | 2023-08-23
T | T337 | Flash Point - ASTM D56 | °C |  |  | 46.32 | 0.9 | 17 | 44.5 | 48 | 2.486 | 4.3 | 0.334 - Pass | 0.27 | Eric Salvail | 2023-08-23
T | T337 | Freezing Point - ASTM D5972 | °C | -54.6 | 0.5 | -54.89 | 0.58 | 18 | -55.8 | -54 | 1.598 | 0.8 | 3.990 - Fail | 0.419 | Eric Salvail | 2023-08-23
T | T337 | Micro Separometer - ASTM D3948 |  |  |  | 55.1 | 25.5 | 10 | 0 | 100 | 70.579 | 24.319 |  | 0.649 | Eric Salvail | 2023-08-23
T | T337 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 88 | 0.04 | 87.9 | 3.5 | 16 | 78 | 93 | 9.695 | 14.941 | 0.421 - Pass | 0.94 | Eric Salvail | 2023-08-23
T | T337 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 1.225 | 0.022 | 15 | 1.19 | 1.27 | 0.061 | 0.106 | 0.329 - Pass | 0.206 | Eric Salvail | 2023-08-23
T | T337 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.1298 | 0.0164 | 12 | 43.106 | 43.159 | 0.045 | 0.046 |  | 0.386 | Eric Salvail | 2023-08-23
T | T337 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 43.1395 | 0.0093 | 4 | 43.13 | 43.152 | 0.026 | 0.035 |  | 0.295 | Eric Salvail | 2023-08-23
T | T337 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 21.08 | 0.72 | 17 | 19.6 | 22 | 1.989 | 3.471 | 0.328 - Pass | 0.342 | Eric Salvail | 2023-08-23
T | T337 | Sulfur - ASTM D4294 | mg/kg |  |  | 651.36 | 541.78 | 5 | 11.4 | 1113.1 | 1500.743 | 124.935 |  | 0.67 | Eric Salvail | 2023-08-23
T | T337 | Sulfur - ASTM D5453 | mg/kg | 1088 | 0.12 | 1080.25 | 64.54 | 10 | 1000.2 | 1202.4 | 178.779 | 136.868 |  | 0.223 | Eric Salvail | 2023-08-23
T | T337 | Sulfur - ASTM D7039 | mg/kg |  |  | 1092.58 | 47.86 | 5 | 1036.6 | 1157.2 | 132.566 | 32.289 |  | 0.252 | Eric Salvail | 2023-08-23
T | T337 | Sulfur - All Methods | Mass % |  |  | 169.5904 | 415.1442 | 6 | 0.101 | 1017 | 1149.949 |  |  | 1.901 | Eric Salvail | 2023-08-23
T | T337 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0017 | 0.85 | 0.00161 | 0.0001 | 16 | 0.0014 | 0.0018 | 0 | 0 | 0.565 - Pass | 0.696 | Eric Salvail | 2023-08-23
T | T337 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 13 | 0 | 0 | 0 |  |  |  | Eric Salvail | 2023-08-23
T | T337 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.6697 | 0.0164 | 9 | 4.651 | 4.704 | 0.046 | 0.025 |  | 0.567 | Eric Salvail | 2023-08-23
T | T337 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.667 | 0.0231 | 7 | 4.639 | 4.702 | 0.064 | 0.149 |  | 0.382 | Eric Salvail | 2023-08-23
T | T338 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 149.1 | 0.59 | 147.29 | 3.08 | 10 | 142.2 | 150.6 | 8.53 | 8.101 |  | 0.493 | Martine Carrier | 2023-08-04
T | T338 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 174.3 | 1.24 | 173.8 | 0.4 | 10 | 173.1 | 174.3 | 1.116 | 3.824 |  | 0.317 | Martine Carrier | 2023-08-04
T | T338 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 211.2 | 0.06 | 211.16 | 0.62 | 10 | 210 | 212 | 1.709 | 3 |  | 0.32 | Martine Carrier | 2023-08-04
T | T338 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 265.7 | -0.2 | 266.02 | 1.61 | 10 | 263.8 | 269.7 | 4.464 | 3.99 |  | 0.437 | Martine Carrier | 2023-08-04
T | T338 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 296.2 | 0.36 | 295.88 | 0.88 | 10 | 294.5 | 297 | 2.439 | 7.1 |  | 0.518 | Martine Carrier | 2023-08-04
T | T338 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.3 | 0.27 | 98.19 | 0.41 | 10 | 97.2 | 98.6 | 1.142 |  |  | 0.696 | Martine Carrier | 2023-08-04
T | T338 | Residue - ASTM D86 (Automatic) | Vol. % | 1.1 | -1.74 | 1.28 | 0.1 | 10 | 1.1 | 1.5 | 0.286 |  |  | 1.03 | Martine Carrier | 2023-08-04
T | T338 | Loss - ASTM D86 (Automatic) | Vol. % | 0.4 | 0 | 0.4 | 0.2 | 9 | 0.1 | 0.8 | 0.554 |  |  | 0.372 | Martine Carrier | 2023-08-04
T | T338 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0016 | 0.0005 | 7 | 0.001 | 0.002 | 0.001 | 0.002 |  | 1.312 | Martine Carrier | 2023-08-04
T | T338 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 17.93 | 0.7 | 3 | 17.2 | 18.6 | 1.946 | 1.906 |  | 0.294 | Martine Carrier | 2023-08-04
T | T338 | Aromatics - ASTM D1319 | Vol. % |  |  | 16.38 | 1.13 | 5 | 14.4 | 17.1 | 3.13 |  |  | 0.901 | Martine Carrier | 2023-08-04
T | T338 | Olefins - ASTM D1319 | Vol. % |  |  | 1.54 | 0.53 | 5 | 1 | 2.2 | 1.461 |  |  | 0.432 | Martine Carrier | 2023-08-04
T | T338 | Density, at 15°C - ASTM D4052 | kg/m³ | 826.7 | -0.25 | 826.72 | 0.08 | 10 | 826.6 | 826.9 | 0.219 | 0.5 |  | 1.557 | Martine Carrier | 2023-08-04
T | T338 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 218 | 0.6 | 201.4 | 27.8 | 8 | 144 | 223 | 77.106 |  |  | 0.824 | Martine Carrier | 2023-08-04
T | T338 | Existent Gum Content - IP540 | mg/100mL |  |  | 3.6 | 3.2 | 7 | 1 | 9 | 8.739 | 5.376 |  | 0.56 | Martine Carrier | 2023-08-04
T | T338 | Flash Point - ASTM D56 | °C |  |  | 44.39 | 0.6 | 9 | 43.5 | 45.5 | 1.665 | 4.3 |  | 0.398 | Martine Carrier | 2023-08-04
T | T338 | Freezing Point - ASTM D5972 | °C | -63.5 | 0.29 | -63.87 | 1.29 | 7 | -65.3 | -61.8 | 3.575 | 0.8 |  | 0.358 | Martine Carrier | 2023-08-04
T | T338 | Micro Separometer - ASTM D3948 |  |  |  | 71 | 5 | 4 | 65 | 76 | 13.757 | 19.924 |  | 0.285 | Martine Carrier | 2023-08-04
T | T338 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 86 | 0.19 | 85.6 | 2.3 | 7 | 83 | 89 | 6.368 | 16.145 |  | 0.375 | Martine Carrier | 2023-08-04
T | T338 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.773 | 0.039 | 7 | 0.69 | 0.81 | 0.108 | 0.081 |  | 0.944 | Martine Carrier | 2023-08-04
T | T338 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.0686 | 0.0145 | 7 | 43.054 | 43.099 | 0.04 | 0.046 |  | 0.858 | Martine Carrier | 2023-08-04
T | T338 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 20.06 | 0.99 | 7 | 19 | 22 | 2.741 | 3.376 |  | 0.443 | Martine Carrier | 2023-08-04
T | T338 | Sulfur - ASTM D5453 | mg/kg | 12.7 | 1.25 | 11.85 | 0.68 | 8 | 10.8 | 12.8 | 1.885 | 3.702 |  | 0.258 | Martine Carrier | 2023-08-04
T | T338 | Sulfur - ASTM D7039 | mg/kg |  |  | 12.2 | 1.34 | 4 | 10.3 | 13.4 | 3.723 | 2.851 |  | 0.412 | Martine Carrier | 2023-08-04
T | T338 | Sulfur, Mercaptan - ASTM D3227 | Mass % | <0.0003 |  | 0.00015 | 0.0001 | 4 | 0 | 0.0002 | 0 | 0 |  | 1.098 | Martine Carrier | 2023-08-04
T | T338 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 6 | 0 | 0 | 0 |  |  |  | Martine Carrier | 2023-08-04
T | T338 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 5.4712 | 0.0333 | 5 | 5.435 | 5.512 | 0.092 | 0.031 |  | 0.367 | Martine Carrier | 2023-08-04
T | T338 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 5.4387 | 0.0145 | 3 | 5.424 | 5.453 | 0.04 | 0.149 |  | 0.285 | Martine Carrier | 2023-08-04
T | T339 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 147.2 | 1.17 | 143.66 | 3.02 | 22 | 137.1 | 147.8 | 8.379 | 7.901 | 1.125 - Pass | 0.849 | Eric Salvail | 2023-11-10
T | T339 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 167.7 | 0.96 | 167.1 | 0.63 | 22 | 165.4 | 168.2 | 1.736 | 3.676 | 0.223 - Pass | 0.351 | Eric Salvail | 2023-11-10
T | T339 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 204.4 | 0.04 | 204.37 | 0.71 | 22 | 203.3 | 205.7 | 1.963 | 3 | 0.428 - Pass | 0.367 | Eric Salvail | 2023-11-10
T | T339 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 257.5 | -0.57 | 258.13 | 1.11 | 22 | 255.1 | 260.6 | 3.069 | 3.872 | 0.628 - Pass | 0.375 | Eric Salvail | 2023-11-10
T | T339 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 286.9 | 1.82 | 284.1 | 1.54 | 22 | 281.9 | 288.6 | 4.279 | 7.1 | 0.363 - Pass | 0.799 | Eric Salvail | 2023-11-10
T | T339 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.6 | 0.55 | 98.38 | 0.39 | 22 | 97.7 | 99.4 | 1.094 |  |  | 0.327 | Eric Salvail | 2023-11-10
T | T339 | Residue - ASTM D86 (Automatic) | Vol. % | 1.2 | 0 | 1.2 | 0.09 | 21 | 1 | 1.4 | 0.263 |  |  | 3.065 | Eric Salvail | 2023-11-10
T | T339 | Loss - ASTM D86 (Automatic) | Vol. % | 0.2 | -0.82 | 0.45 | 0.3 | 22 | 0 | 1.1 | 0.841 |  |  | 0.47 | Eric Salvail | 2023-11-10
T | T339 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0014 | 0.0008 | 17 | 0 | 0.003 | 0.002 | 0.001 | 2.125 - Fail | 2.193 | Eric Salvail | 2023-11-10
T | T339 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 17.05 | 0.49 | 11 | 16.1 | 17.7 | 1.348 | 1.844 |  | 0.368 | Eric Salvail | 2023-11-10
T | T339 | Aromatics - ASTM D1319 | Vol. % |  |  | 16.79 | 0.58 | 10 | 16.2 | 17.8 | 1.607 |  |  | 0.415 | Eric Salvail | 2023-11-10
T | T339 | Olefins - ASTM D1319 | Vol. % |  |  | 0.9 | 0.15 | 9 | 0.6 | 1.1 | 0.416 |  |  | 0.408 | Eric Salvail | 2023-11-10
T | T339 | Density, at 15°C - ASTM D4052 | kg/m³ | 810.4 | 0.75 | 810.34 | 0.08 | 22 | 810.2 | 810.5 | 0.235 | 0.5 | 0.221 - Pass | 1.361 | Eric Salvail | 2023-11-10
T | T339 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 270 | 0.39 | 261.8 | 21.2 | 14 | 215 | 310 | 58.712 |  |  | 0.59 | Eric Salvail | 2023-11-10
T | T339 | Existent Gum Content - IP540 | mg/100mL |  |  | 2.8 | 2 | 15 | 1 | 7.4 | 5.634 | 4.672 | 1.454 - Pass | 1.298 | Eric Salvail | 2023-11-10
T | T339 | Flash Point - ASTM D56 | °C |  |  | 41.32 | 1 | 21 | 40 | 43.5 | 2.76 | 4.3 | 0.412 - Pass | 0.502 | Eric Salvail | 2023-11-10
T | T339 | Freezing Point - ASTM D5972 | °C | -49.3 | 0.67 | -49.83 | 0.79 | 22 | -51.6 | -48.1 | 2.198 | 0.8 | 7.552 - Fail | 0.36 | Eric Salvail | 2023-11-10
T | T339 | Micro Separometer - ASTM D3948 |  |  |  | 86.6 | 7.9 | 11 | 72 | 100 | 21.785 | 13.63 |  | 0.383 | Eric Salvail | 2023-11-10
T | T339 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 91 | -0.24 | 91.4 | 1.5 | 16 | 89 | 94 | 4.276 | 12.929 | 0.109 - Pass | 0.541 | Eric Salvail | 2023-11-10
T | T339 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.435 | 0.016 | 15 | 0.4 | 0.46 | 0.043 | 0.057 | 0.569 - Pass | 0.395 | Eric Salvail | 2023-11-10
T | T339 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.2114 | 0.01 | 15 | 43.19 | 43.227 | 0.028 | 0.046 | 0.361 - Pass | 0.292 | Eric Salvail | 2023-11-10
T | T339 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 43.2055 | 0.015 | 4 | 43.188 | 43.222 | 0.042 | 0.035 |  | 0.231 | Eric Salvail | 2023-11-10
T | T339 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 22.36 | 1.14 | 20 | 20 | 24.2 | 3.158 | 3.592 | 0.773 - Pass | 0.323 | Eric Salvail | 2023-11-10
T | T339 | Sulfur - ASTM D4294 | mg/kg |  |  | 35 | 21.74 | 4 | 11 | 58 | 60.222 | 18.976 |  | 0.279 | Eric Salvail | 2023-11-10
T | T339 | Sulfur - ASTM D5453 | mg/kg | 9.4 | 0.42 | 9.13 | 0.63 | 15 | 7.7 | 10 | 1.738 | 3.046 | 0.326 - Pass | 0.306 | Eric Salvail | 2023-11-10
T | T339 | Sulfur - ASTM D7039 | mg/kg |  |  | 9.42 | 0.5 | 6 | 8.7 | 10.2 | 1.373 | 2.479 |  | 0.249 | Eric Salvail | 2023-11-10
T | T339 | Sulfur - All Methods | Mass % |  |  | 2.5702 | 4.4425 | 3 | 0.0048 | 7.7 | 12.306 |  |  | 0.732 | Eric Salvail | 2023-11-10
T | T339 | Sulfur, Mercaptan - ASTM D3227 | Mass % |  |  | 0.00017 | 9e-05 | 9 | 0 | 0.0003 | 0 | 0 |  | 0.792 | Eric Salvail | 2023-11-10
T | T339 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0.01 | 0.03 | 17 | 0 | 0.1 | 0.092 |  |  | 5.565 | Eric Salvail | 2023-11-10
T | T339 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.6173 | 0.0133 | 11 | 4.585 | 4.637 | 0.037 | 0.025 |  | 0.546 | Eric Salvail | 2023-11-10
T | T339 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.6345 | 0.0454 | 8 | 4.586 | 4.728 | 0.126 | 0.149 |  | 0.579 | Eric Salvail | 2023-11-10
T | T339 | Total Aromatics - ASTM D6379
T | T340 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 169.7 | 1.17 | 167.52 | 1.86 | 19 | 164 | 170.6 | 5.157 | 9.213 | 0.313 - Pass | 0.272 | Eric Salvail | 2023-11-24
T | T340 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 180.3 | 0.54 | 180.07 | 0.42 | 19 | 179.1 | 181 | 1.16 | 3.962 | 0.086 - Pass | 0.59 | Eric Salvail | 2023-11-24
T | T340 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 196.2 | 1.71 | 195.64 | 0.33 | 19 | 194.9 | 196.2 | 0.906 | 3 | 0.091 - Pass | 0.295 | Eric Salvail | 2023-11-24
T | T340 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 234.6 | 1.04 | 233.58 | 0.98 | 19 | 231.6 | 235.6 | 2.718 | 3.504 | 0.602 - Pass | 0.238 | Eric Salvail | 2023-11-24
T | T340 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 275.1 | 0.25 | 274.53 | 2.3 | 19 | 270.6 | 280.1 | 6.36 | 7.1 | 0.802 - Pass | 0.394 | Eric Salvail | 2023-11-24
T | T340 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | 0.02 | 98.19 | 0.31 | 19 | 97.6 | 98.7 | 0.864 |  |  | 0.546 | Eric Salvail | 2023-11-24
T | T340 | Residue - ASTM D86 (Automatic) | Vol. % | 0.9 | -4.87 | 1.22 | 0.07 | 17 | 1.2 | 1.4 | 0.184 |  |  | 5.565 | Eric Salvail | 2023-11-24 |  | recheck demandé
T | T340 | Loss - ASTM D86 (Automatic) | Vol. % | 0.9 | 1.01 | 0.57 | 0.32 | 19 | 0.1 | 1.2 | 0.894 |  |  | 0.871 | Eric Salvail | 2023-11-24
T | T340 | Acid Number - ASTM D3242 | mg KOH/g |  |  | 0.0011 | 0.0004 | 13 | 0.0004 | 0.002 | 0.001 | 0.001 | 0.774 - Pass | 2.757 | Eric Salvail | 2023-11-24
T | T340 | Total Aromatics - ASTM D6379 | Vol. % |  |  | 17.84 | 0.57 | 10 | 17 | 18.5 | 1.579 | 1.9 |  | 0.503 | Eric Salvail | 2023-11-24
T | T340 | Aromatics - ASTM D1319 | Vol. % |  |  | 17.91 | 0.27 | 8 | 17.5 | 18.3 | 0.747 |  |  | 0.228 | Eric Salvail | 2023-11-24
T | T340 | Olefins - ASTM D1319 | Vol. % |  |  | 0.92 | 0.37 | 8 | 0.4 | 1.4 | 1.02 |  |  | 0.385 | Eric Salvail | 2023-11-24
T | T340 | Density, at 15°C - ASTM D4052 | kg/m³ | 812.3 | -0.27 | 812.32 | 0.06 | 20 | 812.2 | 812.4 | 0.162 | 0.5 | 0.105 - Pass | 2.339 | Eric Salvail | 2023-11-24
T | T340 | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 580 | 0.16 | 571.6 | 51.6 | 14 | 480 | 697 | 142.891 |  |  | 0.435 | Eric Salvail | 2023-11-24
T | T340 | Existent Gum Content - IP540 | mg/100mL |  |  | 0.6 | 0.8 | 7 | 0 | 2 | 2.179 | 1.79 |  | 0.907 | Eric Salvail | 2023-11-24
T | T340 | Flash Point - ASTM D56 | °C |  |  | 53.08 | 1.05 | 18 | 51 | 55.2 | 2.906 | 4.3 | 0.457 - Pass | 0.285 | Eric Salvail | 2023-11-24
T | T340 | Freezing Point - ASTM D5972 | °C | -69.9 | 0.29 | -70.07 | 0.57 | 16 | -71 | -69.1 | 1.586 | 0.8 | 3.928 - Fail | 0.5 | Eric Salvail | 2023-11-24
T | T340 | Micro Separometer - ASTM D3948 |  |  |  | 74.7 | 7.6 | 7 | 66 | 86 | 21.139 | 18.606 |  | 0.295 | Eric Salvail | 2023-11-24
T | T340 | Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP rating | 85 | -0.58 | 86.3 | 2.3 | 16 | 83 | 90 | 6.294 | 15.768 | 0.159 - Pass | 0.537 | Eric Salvail | 2023-11-24
T | T340 | Naphthalenes - ASTM D1840 | Vol. % |  |  | 0.48 | 0.02 | 14 | 0.43 | 0.5 | 0.055 | 0.061 | 0.838 - Pass | 1.312 | Eric Salvail | 2023-11-24
T | T340 | Net Heat of Combustion - ASTM D3338 | MJ/kg |  |  | 43.1327 | 0.0124 | 12 | 43.116 | 43.158 | 0.034 | 0.046 |  | 0.181 | Eric Salvail | 2023-11-24
T | T340 | Net Heat of Combustion - ASTM D4529 | MJ/kg |  |  | 43.1155 | 0.0112 | 4 | 43.108 | 43.132 | 0.031 | 0.035 |  | 0.68 | Eric Salvail | 2023-11-24
T | T340 | Smoke Point, corrected - ASTM D1322 | mm |  |  | 21.59 | 0.94 | 16 | 20 | 23 | 2.6 | 3.519 | 0.546 - Pass | 0.289 | Eric Salvail | 2023-11-24
T | T340 | Sulfur - ASTM D4294 | mg/kg |  |  | 514.4 | 19.69 | 6 | 491 | 535 | 54.54 | 107.3 |  | 0.494 | Eric Salvail | 2023-11-24
T | T340 | Sulfur - ASTM D5453 | mg/kg | 485 | 0.15 | 480.32 | 32.01 | 10 | 428.9 | 545.1 | 88.681 | 60.857 |  | 0.662 | Eric Salvail | 2023-11-24
T | T340 | Sulfur - ASTM D7039 | mg/kg |  |  | 496.22 | 10.55 | 5 | 480.7 | 507 | 29.22 | 21.084 |  | 0.36 | Eric Salvail | 2023-11-24
T | T340 | Sulfur - All Methods | Mass % |  |  | 88.3607 | 197.4668 | 5 | 0.045 | 441.6 | 546.983 |  |  | 1.498 | Eric Salvail | 2023-11-24
T | T340 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0003 | -0.19 | 0.00032 | 0.00011 | 16 | 0 | 0.0005 | 0 | 0 | 0.960 - Pass | 1.032 | Eric Salvail | 2023-11-24
T | T340 | Pressure Drop - ASTM D3241 | mm Hg |  |  | 0 | 0 | 13 | 0 | 0 | 0 |  |  |  | Eric Salvail | 2023-11-24
T | T340 | Viscosity at -20°C - ASTM D445 | cSt. |  |  | 4.2071 | 0.0191 | 8 | 4.17 | 4.232 | 0.053 | 0.022 |  | 0.456 | Eric Salvail | 2023-11-24
T | T340 | Kinematic Viscosity at -20°C - ASTM D7042 | mm2/s |  |  | 4.205 | 0.0236 | 8 | 4.173 | 4.243 | 0.065 | 0.149 |  | 0.386 | Eric Salvail | 2023-11-24
P | P119 | Composition |  | 0 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Methane | [13 empty] | 2024-04-25
P | P119 | Ethane |  | 0 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Propene | Vol % | 0.002 | 1.66 | 0.0005 | 0.0009 | 12 | 0 | 0.002 | 0.003 |  |  | 2.944 | Eric Salvail | 2024-04-25
P | P119 | Propane | Vol % | 1.677 | 2.19 | 1.4795 | 0.0901 | 13 | 1.33 | 1.677 | 0.25 |  |  | 0.628 | Eric Salvail | 2024-04-25
P | P119 | C3 | Vol % | 1.679 | 2.1 | 1.4926 | 0.089 | 12 | 1.33 | 1.679 | 0.246 |  |  | 0.546 | Eric Salvail | 2024-04-25
P | P119 | Isobutane | Vol % | 67.418 | 6.19 | 63.9093 | 0.567 | 13 | 62.84 | 64.7 | 1.57 |  |  | 0.29 | Eric Salvail | 2024-04-25
P | P119 | n-Butane | Vol % | 30.532 | -5.45 | 34.1547 | 0.6644 | 13 | 32.93 | 35.18 | 1.841 |  |  | 0.161 | Eric Salvail | 2024-04-25
P | P119 | Density at 15°C | kg/m³ | 568.7 | -2.18 | 569.47 | 0.35 | 21 | 568.7 | 570 | 0.981 |  |  | 0.215 | Eric Salvail | 2024-04-25
P | P119 | Density Method Used |  | D1838 |  |  |  | 4 | [6 empty] | Eric Salvail | 2024-04-25
P | P119 | Sulfur | ppmw | 2.3 | -1.49 | 4 | 1.2 | 23 | 1.7 | 6 | 3.189 |  |  | 0.24 | Eric Salvail | 2024-04-25
P | P119 | Sulfur Method Used |  | D5504 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Vapor Pressure at 37.8°C | kPa | 368 | 2.78 | 361 | 2.5 | 21 | 358 | 368 | 6.979 |  |  | 1.2 | Eric Salvail | 2024-04-25
P | P119 | VP Method Used |  | D1838 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Butane and Heavier | Vol. % | 98.321 | -2.59 | 98.5198 | 0.0769 | 16 | 98.321 | 98.67 | 0.213 |  |  | 0.662 | Eric Salvail | 2024-04-25
P | P119 | Composition Method (eg. CGSB 14.3) |  | D2163 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Injection Method (eg. LSV, GSV) |  | GSV | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Column Type (eg. Al Plot or Hi Res Single Column) | [12 empty] | Eric Salvail | 2024-04-25
P | P119 | Column Length (eg. 100 m) | [12 empty] | Eric Salvail | 2024-04-25
P | P119 | Calibration Type |  | external | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Detector Type (eg. TCD, FID) |  | FID | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Composition |  | 0 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Methane | [13 empty] | 2024-04-25
P | P119 | Ethane |  | 0 | [10 empty] | Eric Salvail | 2024-04-25
P | P119 | Propene | Vol % | 0.002 | 2.01 | 0.0004 | 0.0008 | 21 | 0 | 0.002 | 0.002 | 0.002 | 1.995 - Fail | 5.787 | Eric Salvail | 2024-04-25
P | P119 | Propane | Vol % | 1.677 | 4.36 | 1.4874 | 0.0435 | 26 | 1.37 | 1.566 | 0.121 | 0.23 | 0.275 - Pass | 0.723 | Eric Salvail | 2024-04-25
P | P119 | C3 | Vol % | 1.679 | 4.67 | 1.4937 | 0.0397 | 23 | 1.393 | 1.566 | 0.11 |  |  | 0.362 | Eric Salvail | 2024-04-25
P | P119 | Isobutane | Vol % | 67.418 | 6.63 | 63.7948 | 0.5465 | 28 | 62.47 | 64.7 | 1.514 | 0.87 | 3.029 - Fail | 0.488 | Eric Salvail | 2024-04-25
P | P119 | n-Butane | Vol % | 30.532 | -6.8 | 34.2276 | 0.5431 | 28 | 32.93 | 35.51 | 1.504 | 0.675 | 4.972 - Fail | 0.469 | Eric Salvail | 2024-04-25
F | F316 | Boiling Range Distribution | °C | 116 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-04-30
F | F316 | IBP | [13 empty] | 2024-04-30
F | F316 | 5% | °C | 151.8 | 0.49 | 151.2 | 1.22 | 7 | 149.1 | 153 | 3.37 | 3.768 |  | 0.249 | Eric Salvail | 2024-04-30
F | F316 | 10% | °C | 167.9 | 0.3 | 167.47 | 1.43 | 7 | 165.4 | 170 | 3.955 | 4.012 |  | 0.311 | Eric Salvail | 2024-04-30
F | F316 | 20% | °C | 189.8 | 0.24 | 189.36 | 1.81 | 7 | 186.5 | 192.5 | 5.027 | 4.34 |  | 0.446 | Eric Salvail | 2024-04-30
F | F316 | 30% | °C | 208.1 | 0.31 | 207.4 | 2.24 | 7 | 203.7 | 211 | 6.21 | 3.996 |  | 0.284 | Eric Salvail | 2024-04-30
F | F316 | 40% | °C | 226.2 | 0.25 | 225.57 | 2.47 | 7 | 221.4 | 229.5 | 6.835 | 4.3 |  | 0.328 | Eric Salvail | 2024-04-30
F | F316 | 50% | °C | 245 | 0.37 | 243.99 | 2.72 | 7 | 239.6 | 248 | 7.529 | 4.3 |  | 0.149 | Eric Salvail | 2024-04-30
F | F316 | 60% | °C | 263.3 | 0.32 | 262.43 | 2.71 | 7 | 257.9 | 266.5 | 7.514 | 4.3 |  | 0.192 | Eric Salvail | 2024-04-30
F | F316 | 70% | °C | 281.1 | 0.21 | 280.54 | 2.71 | 7 | 275.8 | 284.5 | 7.504 | 4.3 |  | 0.306 | Eric Salvail | 2024-04-30
F | F316 | 80% | °C | 301 | 0.18 | 300.54 | 2.49 | 7 | 295.8 | 303.5 | 6.903 | 4.3 |  | 0.33 | Eric Salvail | 2024-04-30
F | F316 | 90% | °C | 328.3 | 0.14 | 327.93 | 2.64 | 7 | 322.7 | 330.5 | 7.325 | 4.3 |  | 0.473 | Eric Salvail | 2024-04-30
F | F316 | 95% | °C | 351.1 | 0.1 | 350.83 | 2.62 | 7 | 345.5 | 353 | 7.248 | 5 |  | 0.685 | Eric Salvail | 2024-04-30
F | F316 | Final Boiling Point | °C | 392.7 | -0.17 | 393.21 | 3.11 | 7 | 389.2 | 399 | 8.617 | 11.8 |  | 0.434 | Eric Salvail | 2024-04-30
F | F316 | Distillation Range | °C | 164 | 1.67 | 159.13 | 2.91 | 14 | 154.2 | 164 | 8.064 | 8.752 | 0.849 - Pass | 0.676 | Eric Salvail | 2024-04-30
F | F316 | (corrected, 760 mm Hg) | [13 empty] | 2024-04-30
F | F316 | IBP | [13 empty] | 2024-04-30
F | F316 | 10% Recovered | °C | 188.6 | 0.84 | 187.87 | 0.86 | 14 | 186.7 | 189.8 | 2.391 | 4.133 | 0.335 - Pass | 0.484 | Eric Salvail | 2024-04-30
F | F316 | 50% Recovered | °C | 242.1 | 0.92 | 241.34 | 0.83 | 14 | 239.3 | 242.7 | 2.311 | 3 | 0.594 - Pass | 0.364 | Eric Salvail | 2024-04-30
F | F316 | 90% Recovered | °C | 310.8 | 0.18 | 310.59 | 1.16 | 14 | 308.5 | 312.6 | 3.218 | 4.659 | 0.477 - Pass | 0.6 | Eric Salvail | 2024-04-30
F | F316 | End Point | °C | 346.1 | 1.03 | 344.2 | 1.84 | 14 | 340.1 | 347.5 | 5.103 | 7.1 | 0.517 - Pass | 0.299 | Eric Salvail | 2024-04-30
F | F316 | Recovery | Vol. % | 98.3 | 0.27 | 98.22 | 0.29 | 14 | 97.6 | 98.8 | 0.814 |  |  | 0.379 | Eric Salvail | 2024-04-30
F | F316 | Residue | Vol. % | 1.4 | 0.3 | 1.34 | 0.19 | 14 | 0.9 | 1.7 | 0.529 |  |  | 2.12 | Eric Salvail | 2024-04-30
F | F316 | Loss | Vol. % | 0.3 | -0.68 | 0.44 | 0.2 | 14 | 0.1 | 0.7 | 0.55 |  |  | 0.461 | Eric Salvail | 2024-04-30
F | F316 | Acid Number | mg KOH/g | 0.012 | 1.65 | 0.004 | 0.005 | 10 | 0 | 0.012 | 0.013 | 0.005 |  | 0.763 | Eric Salvail | 2024-04-30
F | F316 | Carbon Residue, 10% Bottoms, Micro |  | 0 | [10 empty] | Eric Salvail | 2024-04-30
F | F316 | Cloud Point | °C |  |  | -34.4 | 1.9 | 8 | -36 | -31 | 5.326 | 4 |  | 0.725 | Eric Salvail | 2024-04-30
F | F316 | Cloud Point (Report to one decimal) | °C | -34.7 | 0.15 | -35.02 | 2.05 | 13 | -37.9 | -31.4 | 5.673 | 2.5 | 5.150 - Fail | 0.265 | Eric Salvail | 2024-04-30
F | F316 | Density, at 15°C | kg/m³ | 838.4 | -1.56 | 838.59 | 0.12 | 16 | 838.3 | 838.8 | 0.334 | 0.5 | 0.445 - Pass | 0.992 | Eric Salvail | 2024-04-30
F | F316 | Flash Point (corrected, 760 mm Hg) | °C | 54.4 | -1.35 | 56.06 | 1.23 | 16 | 54.4 | 58.5 | 3.406 | 3.98 | 0.732 - Pass | 0.417 | Eric Salvail | 2024-04-30
F | F316 | Nitrogen | ppm wt/wt |  |  | 1.94 | 0.14 | 7 | 1.7 | 2.1 | 0.387 | 1.139 |  | 0.459 | Eric Salvail | 2024-04-30
F | F316 | Pour Point (Must be divisible by 3) | °C |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-04-30
F | F316 | Pour Point (Report to nearest degree) | °C | -55 | 0.33 | -56.7 | 5.2 | 11 | -69 | -50 | 14.398 | 6.3 |  | 0.423 | Eric Salvail | 2024-04-30
F | F316 | Sulfur | mg/kg | 4.6 | 0 | 4.6 | 0.25 | 13 | 4.1 | 5.1 | 0.697 | 1.821 | 0.147 - Pass | 0.334 | Eric Salvail | 2024-04-30
F | F316 | Sulfur | mg/kg |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-04-30
F | F316 | Sulfur, Mercaptan | Mass % | 0 | -0.96 | 0.00014 | 0.00015 | 9 | 0 | 0.0004 | 0 | 0 |  | 0.714 | Eric Salvail | 2024-04-30
F | F316 | Viscosity at 40°C | cSt | 2.131 | 0.8 | 2.1262 | 0.0059 | 13 | 2.118 | 2.138 | 0.016 | 0.026 | 0.411 - Pass | 0.418 | Eric Salvail | 2024-04-30
F | F316 | Dynamic Viscosity and Density by Stabinger Viscometer | mm2/s | 2.126 | 0.26 | 2.1204 | 0.0215 | 9 | 2.077 | 2.141 | 0.059 | 0.012 |  | 0.711 | Eric Salvail | 2024-04-30
F | F316 | Kinematic Viscosity at 40°C
D | D338r | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 82.9 | -1.41 | 88.39 | 3.89 | 10 | 82.9 | 96.9 | 10.777 | 5.834 |  | 0.376 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 145.7 | 0.2 | 145.49 | 1.06 | 10 | 143.1 | 146.9 | 2.939 | 3.682 |  | 0.528 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 163.1 | -0.15 | 163.2 | 0.66 | 10 | 162 | 164 | 1.828 | 3.948 |  | 0.466 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 183.5 | -0.19 | 183.62 | 0.64 | 10 | 182.4 | 184.5 | 1.761 | 4.254 |  | 0.319 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 202.9 | -0.38 | 203.13 | 0.6 | 10 | 201.8 | 204 | 1.662 | 3.941 |  | 0.474 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 220.8 | 0.02 | 220.79 | 0.62 | 10 | 219.6 | 222 | 1.72 | 4.3 |  | 0.376 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 238.1 | -0.06 | 238.14 | 0.62 | 10 | 237 | 239 | 1.723 | 4.3 |  | 0.218 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 260.1 | 0.37 | 259.83 | 0.73 | 10 | 258.7 | 261 | 2.019 | 4.3 |  | 0.289 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 288.3 | 1.49 | 287.47 | 0.56 | 11 | 286.6 | 288.3 | 1.543 | 4.3 |  | 0.378 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 304.4 | 0.64 | 303.92 | 0.75 | 11 | 302.7 | 305.1 | 2.072 | 4.3 |  | 0.204 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 314.5 | 0.4 | 314.14 | 0.9 | 10 | 312.8 | 316 | 2.492 | 4.3 |  | 0.399 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 352.2 | -0.45 | 352.58 | 0.84 | 10 | 351.6 | 354.5 | 2.328 | 5 |  | 0.755 | Eric Salvail | 2024-06-03
D | D338r | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 406.7 | 0.51 | 399.44 | 14.29 | 11 | 363.8 | 414.5 | 39.593 | 11.8 |  | 0.764 | Eric Salvail | 2024-06-03
D | D338r | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 146.5 | 1.38 | 142.67 | 2.78 | 31 | 137.6 | 148.6 | 7.692 | 7.847 | 0.961 - Pass | 0.326 | Eric Salvail | 2024-06-03
D | D338r | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 182.8 | 1.21 | 181.65 | 0.95 | 31 | 179.3 | 183.3 | 2.637 | 3.996 | 0.436 - Pass | 0.456 | Eric Salvail | 2024-06-03
D | D338r | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 241.1 | 1.69 | 239.75 | 0.8 | 31 | 238.2 | 241.1 | 2.211 | 3 | 0.543 - Pass | 0.377 | Eric Salvail | 2024-06-03
D | D338r | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 302.8 | -0.08 | 302.87 | 0.84 | 31 | 301.1 | 304.6 | 2.331 | 4.543 | 0.263 - Pass | 0.293 | Eric Salvail | 2024-06-03
D | D338r | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 331.8 | 0.43 | 330.76 | 2.39 | 31 | 326 | 335.9 | 6.619 | 7.1 | 0.869 - Pass | 0.244 | Eric Salvail | 2024-06-03
D | D338r | Recovery - ASTM D86 (Automatic) | Vol. % | 98.2 | 0.82 | 98.01 | 0.23 | 31 | 97.4 | 98.5 | 0.645 |  |  | 0.586 | Eric Salvail | 2024-06-03
D | D338r | Residue - ASTM D86 (Automatic) | Vol. % | 1.3 | -0.34 | 1.32 | 0.05 | 25 | 1.3 | 1.5 | 0.131 |  |  | 6.13 | Eric Salvail | 2024-06-03
D | D338r | Loss - ASTM D86 (Automatic) | Vol. % | 0.5 | -0.52 | 0.63 | 0.26 | 31 | 0.1 | 1.3 | 0.709 |  |  | 0.973 | Eric Salvail | 2024-06-03
D | D338r | Aniline Point - ASTM D611 | °C | 68.35 | 1.33 | 68.044 | 0.23 | 9 | 67.75 | 68.4 | 0.636 | 0.5 |  | 0.278 | Eric Salvail | 2024-06-03
D | D338r | Cetane Index, Calculated - ASTM D976 |  |  |  | 54.92 | 0.26 | 21 | 54.4 | 55.3 | 0.71 |  |  | 0.649 | Eric Salvail | 2024-06-03
D | D338r | Cetane Index, Calculated - ASTM D4737 B |  |  |  | 54.11 | 0.16 | 13 | 53.7 | 54.3 | 0.43 |  |  | 0.776 | Eric Salvail | 2024-06-03
D | D338r | Cetane Index, Calculated - ASTM D4737 A |  |  |  | 55.29 | 0.25 | 21 | 54.8 | 55.7 | 0.68 |  |  | 0.565 | Eric Salvail | 2024-06-03
D | D338r | Cetane Number - ASTM D613 | [5 empty] | 6 | [6 empty] | Eric Salvail | 2024-06-03
D | D338r | Cloud Point - ASTM D2500 | °C |  |  | -29.8 | 1.4 | 10 | -32 | -28 | 3.874 | 4 |  | 0.566 | Eric Salvail | 2024-06-03
D | D338r | Cloud Point - ASTM D5773 | °C | -28.6 | 1.03 | -29.58 | 0.95 | 31 | -31.7 | -28 | 2.633 | 2.5 | 1.110 - Pass | 0.257 | Eric Salvail | 2024-06-03
D | D338r | Copper Strip Corrosion, 3h at 50°C | [12 empty] | Eric Salvail | 2024-06-03
D | D338r | Density, at 15°C - ASTM D4052 | kg/m³ | 809.2 | 1.56 | 809.06 | 0.09 | 35 | 808.8 | 809.3 | 0.243 | 0.5 | 0.236 - Pass | 2.596 | Eric Salvail | 2024-06-03
D | D338r | Electrical Conductivity @ 20°C, ASTM D2624 | pS/m | 241 | -1.1 | 277.6 | 33.3 | 28 | 203 | 368 | 92.106 | 34.48 | 7.136 - Fail | 0.441 | Eric Salvail | 2024-06-03
D | D338r | Flash Point - ASTM D93 | °C | 42.3 | -2.57 | 45.55 | 1.27 | 31 | 42.3 | 48 | 3.504 | 3.234 | 1.174 - Pass | 0.449 | Eric Salvail | 2024-06-03 |  | recheck demandé -2.57Z
D | D338r | HFRR, Major Axis - ASTM D6079 | µm | 336 | -0.96 | 380.4 | 46.2 | 11 | 336 | 456 | 127.854 |  |  | 0.688 | Eric Salvail | 2024-06-03
D | D338r | HFRR, Minor Axis - ASTM D6079 | µm | 322 | 0.58 | 310.8 | 19.4 | 11 | 283 | 349 | 53.852 |  |  | 0.23 | Eric Salvail | 2024-06-03
D | D338r | HFRR, Wear Scar Diameter - ASTM D6079 | µm | 329 | -0.68 | 345.9 | 25 | 12 | 320 | 390 | 69.149 | 80 |  | 0.497 | Eric Salvail | 2024-06-03
D | D338r | LTFT, Minimum Pass Temperature | [12 empty] | Eric Salvail | 2024-06-03
D | D338r | Pour Point - ASTM D97 | °C |  |  | -35.4 | 3.4 | 10 | -42 | -30 | 9.434 | 9 |  | 0.473 | Eric Salvail | 2024-06-03
D | D338r | Pour Point - ASTM D5949 | °C | -34 | 1.47 | -37.1 | 2.1 | 29 | -42 | -33 | 5.845 | 6.3 | 0.861 - Pass | 1.374 | Eric Salvail | 2024-06-03
D | D338r | Sulfur - ASTM D5453 | mg/kg | 0.4 | -0.34 | 0.5 | 0.29 | 23 | 0 | 1.1 | 0.805 | 0.345 | 5.460 - Fail | 0.696 | Eric Salvail | 2024-06-03
D | D338r | Sulfur - ASTM D7039 | mg/kg |  |  | 0.67 | 0.21 | 9 | 0.4 | 1 | 0.571 | 0.593 |  | 0.392 | Eric Salvail | 2024-06-03
D | D338r | Viscosity at 40°C - ASTM D445 | cSt | 1.867 | -0.3 | 1.8707 | 0.0124 | 23 | 1.842 | 1.901 | 0.034 | 0.042 | 0.672 - Pass | 1.466 | Eric Salvail | 2024-06-03
D | D338r | Water Content - ASTM D6304, Procedure A | mg/kg | 37 | 0.84 | 31.4 | 6.7 | 22 | 16 | 39 | 18.484 | 3.354 | 30.376 - Fail | 0.815 | Eric Salvail | 2024-06-03
D | D338r | Dynamic Viscosity at 40°C - ASTM D7042 | mPa•s |  |  | 1.57 | 0.1759 | 8 | 1.46 | 1.873 | 0.487 | 0.045 |  | 1.672 | Eric Salvail | 2024-06-03
D | D338r | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s | 1.874 | 0.72 | 1.8634 | 0.0147 | 15 | 1.841 | 1.882 | 0.041 | 0.055 | 0.553 - Pass | 1.103 | Eric Salvail | 2024-06-03
D | D338r | Density at 40°C - ASTM D7042 | kg/m³ |  |  | 791 | 0.4 | 7 | 790.5 | 791.6 | 1.024 | -5529.79 |  | 0.195 | Eric Salvail | 2024-06-03
D | D338r | Density at 15°C - ASTM D7042 | kg/m³ | 809.4 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
D | D338r | Kinematic Viscosity at 0°C, ASTM D7042 | mm2/s |  |  |  |  | 4 | [6 empty] | Eric Salvail | 2024-06-03
D | D338r | Biodiesel Content - ASTM D7371 | Volume % |  |  | 3.595 | 1.121 | 10 | 1.64 | 4.7 | 3.104 | 0.882 |  | 0.827 | Eric Salvail | 2024-06-03
D- | D-1203 | Test Equipment
D- | D-1203 | Cetane Number - ASTM D613 |  |  |  | 53.79 | 1.22 | 14 | 51.8 | 55.6 | 3.388 | 4.524 | 0.561 - Pass | 0.641
D- | D-1203 | Test Conditions
D- | D-1203 | Setpoint Temperature
D- | D-1203 | Test 1 Charge Air Test Temperature
D- | D-1203 | Test 2 Charge Air Test Temperature
D- | D-1203 | Test Results
D- | D-1203 | Test 1 Ignition Delay
D- | D-1203 | Test 1 Derived Cetane Number - ASTM D6890 | [5 empty] | 6
D- | D-1203 | Test 1 Charge air pressure
D- | D-1203 | Test 1 Injection actuator air pressure
D- | D-1203 | Test 2 Ignition Delay
D- | D-1203 | Test 2 Derived Cetane Number
D- | D-1203 | Test 2 Charge air pressure
D- | D-1203 | Test 2 Injection actuator air pressure
D- | D-1203 | Heptane Check Ignition Delay
D- | D-1203 | QC Sample
D- | D-1203 | Lot Number
D- | D-1203 | ARV
D- | D-1203 | Derived Cetane Number
D- | D-1203 | TALM Precision Package
D- | D-1203 | Average MCH Ignition Delay at last calibration
D- | D-1203 | Date of last calibration
D- | D-1203 | Test 1 Wall Temperature
D- | D-1203 | Test 2 Wall Temperature
D- | D-1203 | Test 1 Ignition Delay
D- | D-1203 | Test 1 Combustion Delay
D- | D-1203 | Test 1 Derived Cetane Number - ASTM D7668 | [5 empty] | 3
D- | D-1203 | Test 1 Charge air pressure
D- | D-1203 | Test 2 Ignition Delay
D- | D-1203 | Test 2 Combustion Delay
D- | D-1203 | Test 2 Derived Cetane Number
D- | D-1203 | Test 2 Charge air pressure
D- | D-1203 | CRM Ignition Delay
D- | D-1203 | CRM Combustion Delay
D- | D-1203 | QC Sample Lot Number
D- | D-1203 | QC Sample ARV
D- | D-1203 | QC Sample Derived Cetane Number
D- | D-1203 | Average MCH Ignition Delay at last calibration
D- | D-1203 | Average MCH Combustion Delay at last calibration
D- | D-1203 | Date of last calibration
D- | D-1203 | Gravity, API at 60 °F - All Methods |  |  |  | 39.83 | 0.03 | 11 | 39.8 | 39.866 | 0.071 |  |  | 0.583
D- | D-1203 | Gravity Method Used
D- | D-1203 | Sulfur - All Methods | mg/kg |  |  | 0.6 | 0.5 | 8 | 0.15 | 1.6 | 1.273 |  |  | 0.885
D- | D-1203 | Sulfur Method Used
D- | D-1203 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 2.5196 | 0.0056 | 9 | 2.512 | 2.528 | 0.016 | 0.056 |  | 0.332
D- | D-1203 | Aniline Point - ASTM D611 | °C |  |  |  |  | 6
D- | D-1203 | Total Aromatics - All Methods | Vol. % |  |  |  |  | 5
D- | D-1203 | Total Aromatics Method Used
D- | D-1203 | Aromatics - ASTM D1319 | Vol. % |  |  |  |  | 4
D- | D-1203 | Olefins - ASTM D1319 | Vol. % |  |  |  |  | 4
D- | D-1203 | Saturates - ASTM D1319 | Vol. % |  |  |  |  | 4
D- | D-1203 | Flash Point - ASTM D93 | °F |  |  | 186.8 | 1.9 | 10 | 184.5 | 190 | 5.397 |  |  | 0.431
D- | D-1203 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F |  |  | 401.27 | 5.67 | 11 | 392 | 409 | 15.698 |  |  | 0.272
D- | D-1203 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F |  |  | 442.2 | 1.48 | 11 | 440 | 444.3 | 4.107 |  |  | 0.378
D- | D-1203 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F |  |  | 480.84 | 0.91 | 11 | 479.7 | 482.5 | 2.515 |  |  | 0.332
D- | D-1203 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F |  |  | 567 | 1.85 | 11 | 564.3 | 570 | 5.116 |  |  | 0.357
D- | D-1203 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F |  |  | 619.71 | 3.09 | 11 | 615.1 | 625.6 | 8.563 |  |  | 0.218
D- | D-1203 | Recovery - ASTM D86 | Vol. % |  |  | 98.31 | 0.27 | 11 | 98 | 99 | 0.748 |  |  | 0.746
D- | D-1203 | Residue - ASTM D86 | Vol. % |  |  | 1.25 | 0.26 | 11 | 0.7 | 1.7 | 0.716 |  |  | 0.841
D- | D-1203 | Loss - ASTM D86 | Vol. % |  |  | 0.44 | 0.33 | 11 | 0 | 1.1 | 0.921 |  |  | 0.388
D- | D-1204 | Cetane Number - ASTM D613 |  |  |  | 50.17 | 0.92 | 15 | 48.2 | 51.4 | 2.559 | 4.072 | 0.395 - Pass | 0.555 | Eric Salvail | 2024-04-17
D- | D-1204 | Test 1 Charge Air Test Temperature |  | 545.4 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 2 Charge Air Test Temperature |  | 545.4 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 1 Derived Cetane Number - ASTM D6890 |  | 48.4 | -0.52 | 48.83 | 0.82 | 7 | 47.7 | 50 | 2.258 | 2.573 |  | 0.461 | Eric Salvail | 2024-04-17
D- | D-1204 | Test 1 Charge air pressure |  | 309.3 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 1 Injection actuator air pressure |  | 174.4 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 2 Ignition Delay |  | 4.254 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 2 Derived Cetane Number |  | 48.3 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 2 Charge air pressure |  | 309.3 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Test 2 Injection actuator air pressure |  | 174.4 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Heptane Check Ignition Delay |  | 3.757 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | ARV |  | 49.79 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Derived Cetane Number |  | 48.18 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | TALM Precision Package |  | Yes | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Average MCH Ignition Delay at last calibration |  | 10.597 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Date of last calibration |  | 2024-01-10 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Gravity, API at 60 °F - All Methods |  |  |  | 34.6 | 0.01 | 10 | 34.58 | 34.62 | 0.029 |  |  | 1.03 | Eric Salvail | 2024-04-17
D- | D-1204 | Gravity Method Used | [12 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Sulfur - All Methods | mg/kg |  |  | 6.6 | 0.8 | 10 | 5.8 | 8.5 | 2.19 |  |  | 0.76 | Eric Salvail | 2024-04-17
D- | D-1204 | Sulfur Method Used |  | d5453 | [10 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Viscosity at 40°C - ASTM D445 | cSt | 3.774 | 1.91 | 3.7635 | 0.0055 | 8 | 3.758 | 3.774 | 0.015 | 0.084 |  | 0.394 | Eric Salvail | 2024-04-17 |  | recheck demandé
D- | D-1204 | Aniline Point - ASTM D611 | °C | 69.85 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Total Aromatics - All Methods | Vol. % |  |  |  |  | 3 | [6 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Total Aromatics Method Used | [12 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Aromatics - ASTM D1319 | Vol. % |  |  |  |  | 3 | [6 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Olefins - ASTM D1319 | Vol. % |  |  |  |  | 3 | [6 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Saturates - ASTM D1319 | Vol. % |  |  |  |  | 3 | [6 empty] | Eric Salvail | 2024-04-17
D- | D-1204 | Flash Point - ASTM D93 | °F |  |  | 159.2 | 2.7 | 10 | 155 | 165 | 7.478 |  |  | 0.614 | Eric Salvail | 2024-04-17
D- | D-1204 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 341.2 | -0.03 | 341.56 | 10.39 | 11 | 319.8 | 360 | 28.788 |  |  | 0.521 | Eric Salvail | 2024-04-17
D- | D-1204 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 474.6 | 1.54 | 469.94 | 3.02 | 10 | 465.6 | 474.6 | 8.364 |  |  | 0.4 | Eric Salvail | 2024-04-17
D- | D-1204 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 554.1 | 1.07 | 553.12 | 0.92 | 11 | 551.7 | 554.5 | 2.535 |  |  | 0.233 | Eric Salvail | 2024-04-17
D- | D-1204 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 638.1 | 0.3 | 637.35 | 2.5 | 11 | 632.5 | 641 | 6.933 |  |  | 0.212 | Eric Salvail | 2024-04-17
D- | D-1204 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 705.5 | 0.62 | 702.96 | 4.11 | 11 | 692.5 | 708 | 11.396 |  |  | 0.674 | Eric Salvail | 2024-04-17
D- | D-1204 | Recovery - ASTM D86 | Vol. % | 98.1 | 0.49 | 97.96 | 0.28 | 11 | 97.4 | 98.3 | 0.766 |  |  | 0.424 | Eric Salvail | 2024-04-17
D- | D-1204 | Residue - ASTM D86 | Vol. % | 1.1 | -0.7 | 1.32 | 0.31 | 11 | 0.7 | 1.9 | 0.866 |  |  | 0.644 | Eric Salvail | 2024-04-17
D- | D-1204 | Loss - ASTM D86 | Vol. % | 0.8 | 0.31 | 0.72 | 0.27 | 11 | 0.3 | 1.2 | 0.741 |  |  | 0.358 | Eric Salvail | 2024-04-17
D- | D-1205 | Cetane Number - ASTM D613 |  |  |  | 52.79 | 2.3 | 12 | 48.9 | 56.3 | 6.372 | 4.399 |  | 0.527 | Eric Salvail | 2024-04-15
D- | D-1205 | Test 1 Charge Air Test Temperature |  | 545.5 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 2 Charge Air Test Temperature |  | 545.6 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 1 Derived Cetane Number - ASTM D6890 |  | w |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 1 Charge air pressure |  | 309.4 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 1 Injection actuator air pressure |  | 174.7 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 2 Ignition Delay |  | 3.778 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 2 Derived Cetane Number |  | w | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 2 Charge air pressure |  | 309.4 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Test 2 Injection actuator air pressure |  | 174.6 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Heptane Check Ignition Delay |  | 3.761 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | QC Sample |  | cfh02 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | ARV |  | 49.79 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Derived Cetane Number |  | 48.83 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | TALM Precision Package |  | yes | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Average MCH Ignition Delay at last calibration | [12 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Date of last calibration | [12 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Gravity, API at 60 °F - All Methods |  |  |  | 28.79 | 0.01 | 9 | 28.77 | 28.8 | 0.03 |  |  | 1.028 | Eric Salvail | 2024-04-15
D- | D-1205 | Gravity Method Used | [12 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Sulfur - All Methods | mg/kg |  |  | 2 | 0.5 | 9 | 1.33 | 2.5 | 1.286 |  |  | 0.734 | Eric Salvail | 2024-04-15
D- | D-1205 | Sulfur Method Used | [12 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Viscosity at 40°C - ASTM D445 | cSt |  |  | 4.4421 | 0.0259 | 10 | 4.406 | 4.495 | 0.072 | 0.1 |  | 0.429 | Eric Salvail | 2024-04-15
D- | D-1205 | Aniline Point |  | <22.00 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Flash Point - ASTM D93 | °F | >120 |  | 276.6 | 11.5 | 7 | 259 | 290 | 31.819 |  |  | 0.368 | Eric Salvail | 2024-04-15
D- | D-1205 | Distillation, IBP Recovered - ASTM D86 (Automatic) | [13 empty] | 2024-04-15
D- | D-1205 | Distillation, 10% Recovered - ASTM D86 (Automatic) |  | 635 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | DIstillation, 50% Recovered - ASTM D86 (Automatic) |  | 638.5 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Distillation, 90% Recovered - ASTM D86 (Automatic) |  | 655.1 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Distillation, FBP Recovered - ASTM D86 (Automatic) |  | 672.7 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Recovery - ASTM D86 |  | 98 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Residue - ASTM D86 |  | 1.3 | [10 empty] | Eric Salvail | 2024-04-15
D- | D-1205 | Loss - ASTM D86 |  | 0.7 | [10 empty] | Eric Salvail | 2024-04-15
F | F317 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 127.4 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 177.5 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 202.9 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 234.2 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 256.6 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 274.2 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 289.4 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 303.1 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 316.3 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 332.4 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 354.2 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 369.9 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 407.9 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 170.9 | 1.44 | 166.08 | 3.35 | 13 | 161.7 | 170.9 | 9.278 | 9.135 | 1.032 - Pass | 0.502 | Eric Salvail | 2024-06-03
F | F317 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 219.6 | 1.28 | 217.05 | 1.98 | 13 | 214 | 221.5 | 5.492 | 4.775 | 1.323 - Pass | 0.294 | Eric Salvail | 2024-06-03
F | F317 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 282.2 | 1.5 | 281.05 | 0.77 | 13 | 279.9 | 282.3 | 2.12 | 3 | 0.500 - Pass | 0.313 | Eric Salvail | 2024-06-03
F | F317 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 333.8 | 0.97 | 332.42 | 1.43 | 13 | 329.6 | 334.4 | 3.964 | 4.986 | 0.632 - Pass | 0.197 | Eric Salvail | 2024-06-03
F | F317 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 357.9 | 0.73 | 356.64 | 1.73 | 12 | 353.7 | 359 | 4.801 | 7.1 |  | 0.399 | Eric Salvail | 2024-06-03
F | F317 | Recovery - ASTM D86 (Automatic) | Vol. % | 98 | -0.2 | 98.12 | 0.63 | 13 | 97.2 | 99.5 | 1.741 |  |  | 0.343 | Eric Salvail | 2024-06-03
F | F317 | Residue - ASTM D86 (Automatic) | Vol. % | 1.4 | 0.07 | 1.39 | 0.12 | 12 | 1.1 | 1.6 | 0.323 |  |  | 1.594 | Eric Salvail | 2024-06-03
F | F317 | Loss - ASTM D86 (Automatic) | Vol. % | 0.6 | 0.12 | 0.55 | 0.47 | 13 | 0 | 1.4 | 1.292 |  |  | 0.675 | Eric Salvail | 2024-06-03
F | F317 | Acid Number - ASTM D664 | mg KOH/g | 0.04 | 0.61 | 0.03 | 0.016 | 9 | 0.01 | 0.06 | 0.044 | 0.023 |  | 0.324 | Eric Salvail | 2024-06-03
F | F317 | Carbon Residue, 10% Bottoms, Micro |  | 0 | [10 empty] | Eric Salvail | 2024-06-03
F | F317 | Cloud Point - ASTM D2500 | °C |  |  | -19.1 | 1.2 | 7 | -20 | -17 | 3.366 | 4 |  | 0.877 | Eric Salvail | 2024-06-03
F | F317 | Cloud Point - ASTM D5773 | °C | -19 | 0.49 | -19.59 | 1.21 | 12 | -21.1 | -17 | 3.358 | 2.5 |  | 0.606 | Eric Salvail | 2024-06-03
F | F317 | Density, at 15°C - ASTM D4052 | kg/m³ | 860.9 | 1.91 | 860.54 | 0.19 | 15 | 860.3 | 861 | 0.521 | 0.5 | 1.087 - Pass | 1.552 | Eric Salvail | 2024-06-03
F | F317 | Flash Point - ASTM D93 | °C | 65.3 | -0.9 | 66.35 | 1.17 | 14 | 63.5 | 68 | 3.228 | 4.711 | 0.470 - Pass | 0.472 | Eric Salvail | 2024-06-03
F | F317 | Nitrogen - ASTM D4629 | ppm wt/wt | 20 | 0.39 | 19.44 | 1.42 | 9 | 17 | 22 | 3.944 | 3.73 |  | 0.48 | Eric Salvail | 2024-06-03
F | F317 | Pour Point - ASTM D97 | °C |  |  | -28.3 | 2.9 | 7 | -33 | -24 | 8.11 | 9 |  | 0.402 | Eric Salvail | 2024-06-03
F | F317 | Pour Point - ASTM D5949 | °C | -23 | 0.88 | -27.8 | 5.5 | 11 | -37 | -23 | 15.213 | 6.3 |  | 0.975 | Eric Salvail | 2024-06-03
F | F317 | Sulfur - ASTM D5453 | mg/kg | 6.8 | -1 | 7.47 | 0.67 | 13 | 6.3 | 9.1 | 1.859 | 2.619 | 0.504 - Pass | 0.533 | Eric Salvail | 2024-06-03
F | F317 | Sulfur - ASTM D7039 | mg/kg |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-06-03
F | F317 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0012 | 6.89 | 0.00018 | 0.00015 | 8 | 0 | 0.0004 | 0 | 0 |  | 0.351 | Eric Salvail | 2024-06-03 |  | rechek demandé 6.89Z
F | F317 | Viscosity at 40°C - ASTM D445 | cSt | 3.375 | 0.54 | 3.3673 | 0.0142 | 10 | 3.335 | 3.388 | 0.039 | 0.036 |  | 0.519 | Eric Salvail | 2024-06-03
F | F317 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s | 3.383 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-03
M | M978 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 26.4 | 1.12 | 25.33 | 0.95 | 23 | 23.1 | 26.8 | 2.639 | 4.7 | 0.315 - Pass | 0.463 | Eric Salvail | 2024-05-31
M | M978 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 37.9 | -1.24 | 39.35 | 1.17 | 24 | 37.9 | 42.7 | 3.232 |  |  | 1.273 | Eric Salvail | 2024-05-31
M | M978 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 66.6 | -0.52 | 66.87 | 0.52 | 24 | 66 | 68 | 1.435 |  |  | 0.572 | Eric Salvail | 2024-05-31
M | M978 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 155.3 | 0.05 | 155.25 | 1.06 | 22 | 153.3 | 158 | 2.94 |  |  | 0.374 | Eric Salvail | 2024-05-31
M | M978 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 238.9 | 0.5 | 234.53 | 8.81 | 24 | 217.5 | 246 | 24.399 | 7.1 | 11.809 - Fail | 0.448 | Eric Salvail | 2024-05-31
M | M978 | Recovery - ASTM D86 (Automatic) | Vol.% | 93.7 | -1.92 | 95.43 | 0.9 | 23 | 93.7 | 96.9 | 2.504 |  |  | 0.324 | Eric Salvail | 2024-05-31
M | M978 | Residue - ASTM D86 (Automatic) | Vol.% | 1.4 | 0.73 | 1.22 | 0.25 | 24 | 0.8 | 1.9 | 0.683 |  |  | 2.522 | Eric Salvail | 2024-05-31
M | M978 | Loss - ASTM D86 (Automatic) | Vol.% | 4.9 | 1.69 | 3.38 | 0.9 | 24 | 2 | 4.9 | 2.49 |  |  | 0.421 | Eric Salvail | 2024-05-31
M | M978 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 61.4 | 0.14 | 61.34 | 0.44 | 22 | 60.5 | 62 | 1.208 |  |  | 0.318 | Eric Salvail | 2024-05-31
M | M978 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 87.3 | -1.18 | 87.83 | 0.45 | 23 | 86.7 | 88.5 | 1.257 |  |  | 0.511 | Eric Salvail | 2024-05-31
M | M978 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 17.31 | 1.94 | 7 | 15.5 | 21.4 | 5.36 |  |  | 0.918 | Eric Salvail | 2024-05-31
M | M978 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 12.5 | 1.25 | 7 | 10.5 | 14.6 | 3.452 |  |  | 0.322 | Eric Salvail | 2024-05-31
M | M978 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 19 | 1.86 | 18.14 | 0.46 | 21 | 17.1 | 19.1 | 1.274 | 2.706 | 0.222 - Pass | 0.398 | Eric Salvail | 2024-05-31
M | M978 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 12.7 | -0.3 | 12.86 | 0.54 | 21 | 11.5 | 13.7 | 1.506 | 1.99 | 0.572 - Pass | 0.527 | Eric Salvail | 2024-05-31
M | M978 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.916 | 1.67 | 0.873 | 0.0257 | 22 | 0.827 | 0.929 | 0.071 | 0.12 | 0.348 - Pass | 0.76 | Eric Salvail | 2024-05-31
M | M978 | Benzene Emissions Number, Winter - Regulation SOR/97-493 |  |  |  | 36.9 | 5.5 | 13 | 23 | 41 | 15.149 |  |  | 1.914 | Eric Salvail | 2024-05-31
M | M978 | Oxygen Content - All Methods | Mass % | 3.759 | 0.07 | 3.75 | 0.15 | 15 | 3.434 | 4 | 0.422 |  |  | 0.369 | Eric Salvail | 2024-05-31
M | M978 | Ethanol - All Methods | Vol. % | 9.8329 | -0.34 | 10.088 | 0.746 | 21 | 8.8 | 12.2 | 2.065 |  |  | 0.522 | Eric Salvail | 2024-05-31
M | M978 | Methanol |  | 0 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Oxygenates Method Used |  | cgsb3.014.3 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | % Corrosion | [12 empty] | Eric Salvail | 2024-05-31
M | M978 | Corrosion, Silver Wool | [12 empty] | Eric Salvail | 2024-05-31
M | M978 | Density, at 15°C - ASTM D4052 | kg/m³ | 725 | 1.69 | 724.35 | 0.39 | 24 | 723.8 | 725.3 | 1.071 | 2.758 | 0.151 - Pass | 0.952 | Eric Salvail | 2024-05-31
M | M978 | Driveablity Index (ethanol corrected) - CAN/CGSB-3.511 | °C |  |  | 428 | 3.1 | 15 | 423 | 435 | 8.633 |  |  | 0.224 | Eric Salvail | 2024-05-31
M | M978 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 14.57 | 3.39 | 19 | 6 | 21.5 | 9.402 | 5.63 | 2.789 - Fail | 0.474 | Eric Salvail | 2024-05-31
M | M978 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.54 | 0.56 | 13 | 0 | 1.5 | 1.545 | 2.071 | 0.556 - Pass | 0.742 | Eric Salvail | 2024-05-31
M | M978 | Induction Period at 100°C, ASTM D525 | min | 999 |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-05-31
M | M978 | Sulfur - ASTM D5453 | mg/kg |  |  | 6.08 | 1.1 | 23 | 3.9 | 8.6 | 3.037 | 2.245 | 1.830 - Pass | 0.242 | Eric Salvail | 2024-05-31
M | M978 | Sulfur - ASTM D7039 | mg/kg | 4.7 | -1.79 | 5.94 | 0.69 | 7 | 4.7 | 6.8 | 1.925 | 1.933 |  | 0.241 | Eric Salvail | 2024-05-31
M | M978 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0002 | -0.26 | 0.00025 | 0.00019 | 10 | 0 | 0.0007 | 0.001 | 0 |  | 0.897 | Eric Salvail | 2024-05-31
M | M978 | Vapor Pressure - ASTM D5191 | kPa | 106 | 1.09 | 105.11 | 0.81 | 24 | 103.3 | 106.2 | 2.256 |  |  | 0.436 | Eric Salvail | 2024-05-31
M | M978 | Container Size |  | 4l | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Octane Number, Research - ASTM D2699 |  | 91.5 | 2.33 | 90.75 | 0.32 | 20 | 90.3 | 91.5 | 0.892 | 0.7 | 1.624 - Pass | 0.605 | Eric Salvail | 2024-05-31 |  | recheck demandé 2.33Z
M | M978 | Octane Number, Motor - ASTM D2700 |  | 82.3 | 0.71 | 82.1 | 0.28 | 19 | 81.5 | 82.6 | 0.783 | 0.9 | 0.758 - Pass | 0.451 | Eric Salvail | 2024-05-31
M | M978 | RON Method Used |  | cr | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Barometric Pressure |  | 29.76 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | IAT, (ACTUAL) |  | 157 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Cylinder Height Reading (Sample) |  | 747 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | TSF Rating (ACTUAL) |  | 93.4 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | MON Method Used |  | cr | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Barometric Pressure (at time of TSF rating) |  | 29.78 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | IAT |  | 100 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | MAT |  | 314 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | TSF Rating (ACTUAL) |  | 85.2 | [10 empty] | Eric Salvail | 2024-05-31
M | M978 | Cylinder Height Reading (Sample) |  | 599 | [10 empty] | Eric Salvail | 2024-05-31
C | C544 | Octane Number, Research - ASTM D2699 |  | 91 | 0.07 | 90.98 | 0.26 | 17 | 90.5 | 91.3 | 0.715 | 0.7 | 1.042 - Pass | 0.556 | Martine Carrier | 2024-05-13
C | C544 | RON Method Used |  | CR | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | Barometric Pressure |  | 29.78 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | IAT, (ACTUAL) |  | 163 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | Cylinder Height Reading (Sample) |  | 740 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | TSF Rating (ACTUAL) |  | 93.4 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | Octane Number, Motor - ASTM D2700 |  | 83.6 | 0.3 | 83.51 | 0.31 | 17 | 83.1 | 84.3 | 0.868 | 0.9 | 0.929 - Pass | 0.632 | Martine Carrier | 2024-05-13 |  | comment test
C | C544 | MON Method Used |  | CR | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | Barometric Pressure (at time of TSF rating) |  | 29.73 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | IAT |  | 100 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | MAT |  | 318 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | Cylinder Height Reading (Sample) |  | 622 | [10 empty] | Martine Carrier | 2024-05-13
C | C544 | TSF Rating (ACTUAL) |  | 85.2 | [10 empty] | Martine Carrier | 2024-05-13
M | M977 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 29.3 | -0.44 | 30.06 | 1.71 | 29 | 26.8 | 33.8 | 4.729 | 4.7 | 1.012 - Pass | 0.498 | Eric Salvail | 2024-04-12
M | M977 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 49.8 | -1.48 | 51.72 | 1.3 | 29 | 49.8 | 55.7 | 3.599 |  |  | 0.461 | Eric Salvail | 2024-04-12
M | M977 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 105.3 | -1.8 | 106.43 | 0.63 | 28 | 105.2 | 107.4 | 1.739 |  |  | 0.324 | Eric Salvail | 2024-04-12
M | M977 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 157.9 | -1.21 | 158.78 | 0.73 | 29 | 156.6 | 159.7 | 2.009 |  |  | 1.223 | Eric Salvail | 2024-04-12
M | M977 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 201.7 | 2.56 | 196.52 | 2.02 | 29 | 193.1 | 201.7 | 5.602 | 7.1 | 0.622 - Pass | 0.842 | Eric Salvail | 2024-04-12
M | M977 | Recovery - ASTM D86 (Automatic) | Vol.% | 98 | 0.33 | 97.85 | 0.46 | 27 | 96.9 | 98.9 | 1.261 |  |  | 0.294 | Eric Salvail | 2024-04-12
M | M977 | Residue - ASTM D86 (Automatic) | Vol.% | 0.6 | -11.5 | 1 | 0.04 | 25 | 0.9 | 1.1 | 0.097 |  |  | 6.933 | Eric Salvail | 2024-04-12
M | M977 | Loss - ASTM D86 (Automatic) | Vol.% | 1.4 | 0.53 | 1.16 | 0.45 | 28 | 0.3 | 2.1 | 1.249 |  |  | 0.344 | Eric Salvail | 2024-04-12
M | M977 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 41.7 | 1.48 | 41.03 | 0.46 | 27 | 40.2 | 41.7 | 1.264 |  |  | 0.69 | Eric Salvail | 2024-04-12
M | M977 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 84.1 | 4.98 | 82.83 | 0.26 | 26 | 82.3 | 83.3 | 0.708 |  |  | 0.396 | Eric Salvail | 2024-04-12
M | M977 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  | 26.83 | 1.16 | 8 | 25.2 | 29.1 | 3.216 |  |  | 0.615 | Eric Salvail | 2024-04-12
M | M977 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  | 1.04 | 0.31 | 8 | 0.5 | 1.4 | 0.85 |  |  | 0.212 | Eric Salvail | 2024-04-12
M | M977 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 29.5 | 0.46 | 29.27 | 0.51 | 22 | 28.2 | 30.2 | 1.4 | 2.706 | 0.268 - Pass | 0.168 | Eric Salvail | 2024-04-12
M | M977 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 2.3 | 0.14 | 2.24 | 0.41 | 21 | 1.2 | 3.3 | 1.147 | 0.831 | 1.906 - Pass | 0.498 | Eric Salvail | 2024-04-12
M | M977 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.373 | 0.09 | 0.3723 | 0.008 | 23 | 0.353 | 0.388 | 0.022 | 0.051 | 0.187 - Pass | 0.431 | Eric Salvail | 2024-04-12
M | M977 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 35.3 | 0.6 | 14 | 34 | 36 | 1.705 |  |  | 1.242 | Eric Salvail | 2024-04-12
M | M977 | Oxygen Content - All Methods |  | 0 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Ethanol - All Methods |  | 0 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Methanol - All Methods |  | 0 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Oxygenates Method Used | [12 empty] | Eric Salvail | 2024-04-12
M | M977 | Copper Strip Corrosion, 3h at 50°C - ASTM D130 | [12 empty] | Eric Salvail | 2024-04-12
M | M977 | Corrosive Properties | [12 empty] | Eric Salvail | 2024-04-12
M | M977 | Rating - NACE TM0172 | [13 empty] | 2024-04-12
M | M977 | % Corrosion - NACE TM0172 | [12 empty] | Eric Salvail | 2024-04-12
M | M977 | Corrosion, Silver Wool - CAN/CGSB-3.0 No. 60.32 | [12 empty] | Eric Salvail | 2024-04-12
M | M977 | Density, at 15°C - ASTM D4052 | kg/m³ | 744.2 | 1.59 | 743.37 | 0.52 | 27 | 742.9 | 745 | 1.449 | 2.159 | 0.451 - Pass | 2.118 | Eric Salvail | 2024-04-12
M | M977 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 555.1 | 2.9 | 18 | 552 | 562 | 8.062 |  |  | 1.004 | Eric Salvail | 2024-04-12
M | M977 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 10.03 | 3.24 | 19 | 1.4 | 14 | 8.987 | 4.849 | 3.435 - Fail | 0.731 | Eric Salvail | 2024-04-12
M | M977 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 1.02 | 0.82 | 12 | 0 | 3 | 2.265 | 2.506 |  | 0.494 | Eric Salvail | 2024-04-12
M | M977 | Induction Period at 100°C, ASTM D525 | min | 999 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-04-12
M | M977 | Sulfur - ASTM D5453 | mg/kg | 18 | 2.18 | 15.39 | 1.2 | 26 | 13 | 18 | 3.323 | 4.504 | 0.544 - Pass | 0.429 | Eric Salvail | 2024-04-12
M | M977 | Sulfur - ASTM D7039 | mg/kg |  |  | 14.98 | 1.76 | 8 | 12.2 | 17.1 | 4.865 | 3.184 |  | 0.61 | Eric Salvail | 2024-04-12
M | M977 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0014 | -0.15 | 0.00143 | 0.00019 | 18 | 0.0009 | 0.0017 | 0.001 | 0 | 2.032 - Fail | 0.964 | Eric Salvail | 2024-04-12
M | M977 | Vapor Pressure - ASTM D5191 | kPa | 62.72 | -0.31 | 62.91 | 0.6 | 27 | 61.6 | 64.7 | 1.663 |  |  | 0.577 | Eric Salvail | 2024-04-12
M | M977 | Container Size - ASMT D5191 |  | 4L | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Octane Number, Research - ASTM D2699 |  | 85.7 | 2.34 | 85.2 | 0.21 | 24 | 84.9 | 85.7 | 0.586 | 0.7 | 0.701 - Pass | 0.683 | Eric Salvail | 2024-04-12
M | M977 | Octane Number, Motor - ASTM D2700 |  | 79.5 | -0.8 | 79.82 | 0.41 | 25 | 79.2 | 81 | 1.123 | 0.9 | 1.557 - Pass | 0.467 | Eric Salvail | 2024-04-12
M | M977 | RON Method Used - ASTM D2699 |  | cr | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Barometric Pressure - RON |  | 24.9 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | IAT, (ACTUAL) - RON |  | 114 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Cylinder Height Reading (Sample) - RON |  | 670 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | TSF Rating (ACTUAL) - RON |  | 93.4 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | MON Method Used - ASTM D2700 |  | cr | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Barometric Pressure (at time of TSF rating) - MON |  | 29.4 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | IAT - MON |  | 100 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | MAT - MON |  | 253 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | TSF Rating (ACTUAL) - MON |  | 85.2 | [10 empty] | Eric Salvail | 2024-04-12
M | M977 | Cylinder Height Reading (Sample) - MON |  | 551 | [10 empty] | Eric Salvail | 2024-04-12
M | M979 | Distillation, IBP Evaporated - ASTM D86 (Automatic) | °C | 33.9 | 0.77 | 32.82 | 1.4 | 23 | 29.5 | 35.2 | 3.884 | 4.7 | 0.683 - Pass | 0.719 | Eric Salvail | 2024-06-21
M | M979 | Distillation, 10% Evaporated - ASTM D86 (Automatic) | °C | 49.3 | -0.83 | 49.74 | 0.53 | 22 | 48.7 | 51 | 1.47 |  |  | 0.428 | Eric Salvail | 2024-06-21
M | M979 | Distillation, 50% Evaporated - ASTM D86 (Automatic) | °C | 71 | -0.6 | 71.68 | 1.14 | 22 | 69.1 | 74.7 | 3.147 |  |  | 0.524 | Eric Salvail | 2024-06-21
M | M979 | Distillation, 90% Evaporated - ASTM D86 (Automatic) | °C | 163.6 | -0.51 | 163.98 | 0.74 | 22 | 161.8 | 165.1 | 2.05 |  |  | 0.784 | Eric Salvail | 2024-06-21
M | M979 | Distillation, FBP Evaporated - ASTM D86 (Automatic) | °C | 212.6 | 0.21 | 212.12 | 2.33 | 23 | 205.4 | 216.2 | 6.456 | 7.1 | 0.827 - Pass | 0.362 | Eric Salvail | 2024-06-21
M | M979 | Recovery - ASTM D86 (Automatic) | Vol.% | 96.9 | -1.99 | 97.92 | 0.51 | 22 | 96.9 | 98.8 | 1.419 |  |  | 0.204 | Eric Salvail | 2024-06-21
M | M979 | Residue - ASTM D86 (Automatic) | Vol.% | 1 | -1.63 | 1.09 | 0.06 | 20 | 1 | 1.2 | 0.153 |  |  | 3.013 | Eric Salvail | 2024-06-21
M | M979 | Loss - ASTM D86 (Automatic) | Vol.% | 2.1 | 2.12 | 0.99 | 0.52 | 23 | 0 | 2.1 | 1.449 |  |  | 0.214 | Eric Salvail | 2024-06-21
M | M979 | Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) |  | 54.2 | -5.54 | 56.31 | 0.38 | 21 | 55.2 | 57.1 | 1.054 |  |  | 0.548 | Eric Salvail | 2024-06-21 |  | recheck demandé 2024.07.03.001
M | M979 | Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) |  | 82.6 | -6.38 | 84.65 | 0.32 | 21 | 84.2 | 85.5 | 0.892 |  |  | 1.018 | Eric Salvail | 2024-06-21 |  | recheck demandé
M | M979 | Aromatics (corrected) - ASTM D1319 | Vol. % |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-06-21
M | M979 | Olefins (corrected) - ASTM D1319 | Vol. % |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-06-21
M | M979 | Aromatics - CAN/CGSB-3.0 No. 14.3 | Vol. % | 22.5 | 1.55 | 21.29 | 0.78 | 20 | 19.6 | 22.8 | 2.174 | 2.706 | 0.646 - Pass | 0.697 | Eric Salvail | 2024-06-21
M | M979 | Olefins - CAN/CGSB-3.0 No. 14.3 | Vol. % | 9.7 | -0.24 | 9.77 | 0.29 | 20 | 9.1 | 10.2 | 0.794 | 1.735 | 0.210 - Pass | 0.41 | Eric Salvail | 2024-06-21
M | M979 | Benzene - CAN/CGSB-3.0 No. 14.3 | Vol % | 0.444 | 2.22 | 0.3904 | 0.0242 | 21 | 0.357 | 0.444 | 0.067 | 0.054 | 1.547 - Pass | 0.77 | Eric Salvail | 2024-06-21
M | M979 | Benzene Emissions Number, Summer - Regulation SOR/97-493 |  |  |  | 25.3 | 0.6 | 13 | 24 | 26 | 1.66 |  |  | 1.264 | Eric Salvail | 2024-06-21
M | M979 | Oxygen Content - All Methods | Mass % | 3.7087 | -0.05 | 3.71 | 0.12 | 14 | 3.445 | 3.9 | 0.339 |  |  | 0.341 | Eric Salvail | 2024-06-21
M | M979 | Ethanol - All Methods | Vol. % | 9.9397 | -0.27 | 10.019 | 0.288 | 17 | 9.27 | 10.39 | 0.798 |  |  | 0.493 | Eric Salvail | 2024-06-21
M | M979 | Methanol - All Methods | [12 empty] | Eric Salvail | 2024-06-21
M | M979 | Oxygenates Method Used |  | cgsb3.014.3 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Density, at 15°C - ASTM D4052 | kg/m³ | 741.4 | 0.1 | 741.34 | 0.61 | 24 | 740.7 | 743.2 | 1.68 | 2.223 | 0.571 - Pass | 1.191 | Eric Salvail | 2024-06-21
M | M979 | Driveability Index - CAN/CGSB-3.5 | °C |  |  | 464.5 | 6.9 | 17 | 453 | 477 | 19.144 |  |  | 0.446 | Eric Salvail | 2024-06-21
M | M979 | Gum, Unwashed Content - ASTM D381 | mg/100mL |  |  | 6.46 | 0.97 | 16 | 5 | 8 | 2.701 | 4.067 | 0.441 - Pass | 0.271 | Eric Salvail | 2024-06-21
M | M979 | Gum, Solvent Washed Content - ASTM D381 | mg/100mL |  |  | 0.55 | 0.47 | 8 | 0 | 1.4 | 1.308 | 2.085 |  | 0.377 | Eric Salvail | 2024-06-21
M | M979 | Induction Period at 100°C, ASTM D525 | min |  |  |  |  | 3 | [6 empty] | Eric Salvail | 2024-06-21
M | M979 | Sulfur - ASTM D5453 | mg/kg | 4.9 | -0.09 | 4.95 | 0.53 | 21 | 4.3 | 6.3 | 1.473 | 1.923 | 0.586 - Pass | 0.653 | Eric Salvail | 2024-06-21
M | M979 | Sulfur - ASTM D7039 | mg/kg |  |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
M | M979 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0.0001 | -0.7 | 0.00018 | 0.00012 | 12 | 0 | 0.0004 | 0 | 0 |  | 0.626 | Eric Salvail | 2024-06-21
M | M979 | Vapor Pressure - ASTM D5191 | kPa | 68.9 | -0.22 | 69.04 | 0.63 | 22 | 68.2 | 70.9 | 1.743 |  |  | 0.614 | Eric Salvail | 2024-06-21
M | M979 | Container Size - ASMT D5191 |  | 4l | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Octane Number, Research - ASTM D2699 |  | 93.4 | 0.37 | 93.33 | 0.2 | 21 | 92.9 | 93.8 | 0.541 | 0.7 | 0.597 - Pass | 0.566 | Eric Salvail | 2024-06-21
M | M979 | Octane Number, Motor - ASTM D2700 |  | 83.9 | 1.34 | 83.44 | 0.34 | 21 | 82.8 | 83.9 | 0.953 | 0.9 | 1.122 - Pass | 0.541 | Eric Salvail | 2024-06-21
M | M979 | RON Method Used - ASTM D2699 |  | cr | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Barometric Pressure - RON |  | 29.9 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | IAT, (ACTUAL) - RON |  | 157 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Cylinder Height Reading (Sample) - RON |  | 778 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | TSF Rating (ACTUAL) - RON |  | 93.4 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | MON Method Used - ASTM D2700 |  | cr | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Barometric Pressure (at time of TSF rating) - MON |  | 29.9 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | IAT - MON |  | 99.8 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | MAT - MON |  | 309 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | TSF Rating (ACTUAL) - MON |  | 85.2 | [10 empty] | Eric Salvail | 2024-06-21
M | M979 | Cylinder Height Reading (Sample) - MON |  | 628 | [10 empty] | Eric Salvail | 2024-06-21
D- | D-1206 | Cetane Number - ASTM D613 |  |  |  | 46.96 | 1.63 | 14 | 43.9 | 50 | 4.52 | 3.671 | 1.516 - Pass | 0.432 | Fowler | 2024-07-02
D- | D-1206 | Setpoint Temperature - ASTM D6890
D- | D-1206 | Test 1 Charge Air Test Temperature - ASTM D6890 |  | 547.4 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 2 Charge Air Test Temperature - ASTM D6890 |  | 547 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test Results |  | 4.402 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 1 Ignition Delay - ASTM D6890
D- | D-1206 | Test 1 Derived Cetane Number - ASTM D6890 |  | 46.85 | 0.38 | 46.63 | 0.58 | 8 | 45.9 | 47.61 | 1.604 | 2.488 |  | 0.295 | Fowler | 2024-07-02
D- | D-1206 | Test 1 Charge air pressure - ASTM D6890 |  | 309.4 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 1 Injection actuator air pressure - ASTM D6890 |  | 174.7 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 2 Ignition Delay - ASTM D6890 |  | 4.421 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 2 Derived Cetane Number - ASTM D6890 |  | 46.67 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 2 Charge air pressure - ASTM D6890 |  | 309.4 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Test 2 Injection actuator air pressure - ASTM D6890 |  | 174.6 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Heptane Check Ignition Delay - ASTM D6890 |  | 3.778 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | QC Sample |  | 17CPCFL01 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Lot Number - ASTM D6890
D- | D-1206 | ARV - ASTM D6890 |  | 41.91 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Derived Cetane Number - ASTM D6890 |  | 42.06 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | TALM Precision Package - ASTM D6890 |  | yes | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Average MCH Ignition Delay at last calibration - ASTM D6890 |  | 10.85 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Date of last calibration - ASTM D6890 |  | April,14th,2023 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Gravity, API at 60 °F - All Methods |  | 36 | -20.77 | 36.1 | 0 | 10 | 36.09 | 36.105 | 0.013 |  |  | 1.836 | Fowler | 2024-07-02
D- | D-1206 | Gravity Method Used |  | ASTMD4052 | [10 empty] | Fowler | 2024-07-02
D- | D-1206 | Sulfur - All Methods | mg/kg |  |  | 7 | 0.5 | 10 | 6 | 7.9 | 1.373 |  |  | 0.61 | Fowler | 2024-07-02
D- | D-1206 | Sulfur Method Used | [12 empty] | Fowler | 2024-07-02
D- | D-1206 | Viscosity at 40°C - ASTM D445 | cSt | 2.603 | 0.85 | 2.5964 | 0.0078 | 11 | 2.585 | 2.608 | 0.022 | 0.058 |  | 0.323 | Fowler | 2024-07-02
D- | D-1206 | Aniline Point - ASTM D611 | °C | 64.25 | -0.97 | 64.687 | 0.452 | 7 | 64.25 | 65.56 | 1.252 | 0.5 |  | 0.57 | Fowler | 2024-07-02
D- | D-1206 | Flash Point - ASTM D93 | °F | 127.8 | 0.83 | 126.5 | 1.6 | 10 | 124 | 129.2 | 4.478 |  |  | 0.205 | Fowler | 2024-07-02
D- | D-1206 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 326.3 | -0.27 | 327.29 | 3.64 | 11 | 321.2 | 333.37 | 10.074 |  |  | 0.327 | Fowler | 2024-07-02
D- | D-1206 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 401.9 | 0.27 | 401.33 | 2.14 | 11 | 397.2 | 404.72 | 5.923 |  |  | 0.261 | Fowler | 2024-07-02
D- | D-1206 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 503.7 | -0.19 | 503.99 | 1.49 | 11 | 502 | 506.84 | 4.118 |  |  | 0.261 | Fowler | 2024-07-02
D- | D-1206 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 624.8 | -1.18 | 627.86 | 2.58 | 11 | 624.8 | 633.12 | 7.155 |  |  | 0.327 | Fowler | 2024-07-02
D- | D-1206 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 684.7 | 1.29 | 681.68 | 2.34 | 11 | 677.1 | 684.7 | 6.487 |  |  | 0.354 | Fowler | 2024-07-02
D- | D-1206 | Recovery - ASTM D86 | Vol. % | 98.9 | 2.25 | 98.18 | 0.32 | 11 | 97.8 | 98.9 | 0.883 |  |  | 0.596 | Fowler | 2024-07-02 |  | 2.25 Z recheck demandé 2024.07.03.002
D- | D-1206 | Residue - ASTM D86 | Vol. % | 0.7 | -1.45 | 1.26 | 0.39 | 11 | 0.5 | 1.9 | 1.075 |  |  | 0.63 | Fowler | 2024-07-02
D- | D-1206 | Loss - ASTM D86 | Vol. % | 0.4 | -0.55 | 0.59 | 0.34 | 11 | 0.2 | 1.4 | 0.955 |  |  | 0.562 | Fowler | 2024-07-02
F | F318 | Simulated Distillation, IBP, °C - ASTM D2887 | °C | 102.1 | -0.48 | 102.79 | 1.42 | 7 | 101.2 | 105.1 | 3.945 | 6.784 |  | 0.259 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 5%, °C - ASTM D2887 | °C | 143.9 | -0.66 | 144.61 | 1.09 | 7 | 143.1 | 146.6 | 3.011 | 3.669 |  | 0.308 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 10%, °C - ASTM D2887 | °C | 162.2 | -0.38 | 162.46 | 0.67 | 7 | 161.6 | 163.6 | 1.864 | 3.937 |  | 0.207 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 20%, °C - ASTM D2887 | °C | 190.2 | 0.38 | 189.79 | 1.09 | 7 | 188.6 | 191.9 | 3.015 | 4.347 |  | 0.435 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 30%, °C - ASTM D2887 | °C | 213.7 | 0.31 | 213.23 | 1.52 | 7 | 211.1 | 215.3 | 4.197 | 4.072 |  | 0.249 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 40%, °C - ASTM D2887 | °C | 233.9 | 0.29 | 233.41 | 1.65 | 7 | 231 | 235.4 | 4.57 | 4.3 |  | 0.253 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 50%, °C - ASTM D2887 | °C | 253.6 | 0.64 | 252.49 | 1.74 | 7 | 249.8 | 254.5 | 4.81 | 4.3 |  | 0.251 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 60%, °C - ASTM D2887 | °C | 271.8 | 0.73 | 270.49 | 1.79 | 7 | 267.3 | 272.2 | 4.961 | 4.3 |  | 0.405 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 70%, °C - ASTM D2887 | °C | 292.3 | 0.5 | 291.16 | 2.29 | 7 | 286.9 | 293.1 | 6.353 | 4.3 |  | 0.505 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 80%, °C - ASTM D2887 | °C | 313.1 | 0.38 | 311.97 | 2.97 | 7 | 306 | 314.3 | 8.237 | 4.3 |  | 0.704 | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 90%, °C - ASTM D2887 | °C | 341.7 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, 95%, °C - ASTM D2887 | °C | 363.5 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Simulated Distillation, FBP, °C - ASTM D2887 | °C | 419.8 |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °C | 150.8 | 0.79 | 148.85 | 2.48 | 13 | 144.4 | 153.7 | 6.882 | 8.187 | 0.707 - Pass | 0.296 | Eric Salvail | 2024-06-21
F | F318 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °C | 182.5 | 1.32 | 181.35 | 0.87 | 13 | 179.9 | 183.3 | 2.411 | 3.99 | 0.365 - Pass | 0.472 | Eric Salvail | 2024-06-21
F | F318 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °C | 250.2 | 0.01 | 250.19 | 0.93 | 13 | 249 | 251.8 | 2.57 | 3 | 0.734 - Pass | 0.353 | Eric Salvail | 2024-06-21
F | F318 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °C | 322.3 | -0.03 | 322.35 | 1.51 | 13 | 320.7 | 326 | 4.188 | 4.835 | 0.750 - Pass | 0.655 | Eric Salvail | 2024-06-21
F | F318 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °C | 360 | 1.66 | 355.9 | 2.47 | 13 | 352.6 | 361.1 | 6.852 | 7.1 | 0.931 - Pass | 0.518 | Eric Salvail | 2024-06-21
F | F318 | Recovery - ASTM D86 (Automatic) | Vol. % | 98.7 | 1.18 | 98.28 | 0.35 | 13 | 97.8 | 98.9 | 0.978 |  |  | 0.345 | Eric Salvail | 2024-06-21
F | F318 | Residue - ASTM D86 (Automatic) | Vol. % | 0.6 | -2.18 | 1.31 | 0.33 | 13 | 0.6 | 1.9 | 0.901 |  |  | 1.006 | Eric Salvail | 2024-06-21 |  | recheck demandé DA-07, 2024.07.08.001
F | F318 | Loss - ASTM D86 (Automatic) | Vol. % | 0.7 | 1.02 | 0.42 | 0.28 | 13 | 0 | 1 | 0.774 |  |  | 0.233 | Eric Salvail | 2024-06-21
F | F318 | Acid Number - ASTM D664 | mg KOH/g | 0.005 | 0.58 | 0.003 | 0.004 | 8 | 0 | 0.01 | 0.01 | 0.003 |  | 0.763 | Eric Salvail | 2024-06-21
F | F318 | Carbon Residue, 10% Bottoms, Micro - ASTM D4530 |  | 0 | [10 empty] | Eric Salvail | 2024-06-21
F | F318 | Cloud Point - ASTM D2500 | °C |  |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Cloud Point - ASTM D5773 | °C | -18.5 | 0.07 | -18.57 | 0.93 | 12 | -19.9 | -16.7 | 2.567 | 2.5 |  | 0.19 | Eric Salvail | 2024-06-21
F | F318 | Density, at 15°C - ASTM D4052 | kg/m³ | 838.4 | -0.44 | 838.44 | 0.09 | 15 | 838.2 | 838.6 | 0.252 | 0.5 | 0.254 - Pass | 1.285 | Eric Salvail | 2024-06-21
F | F318 | Flash Point - ASTM D93 | °C | 46 | -1.92 | 48.69 | 1.41 | 14 | 46 | 50.5 | 3.892 | 3.457 | 1.268 - Pass | 0.361 | Eric Salvail | 2024-06-21 |  | Recheck demandé, PEA-07, 2024.07.08.001
F | F318 | Nitrogen - ASTM D4629 | ppm wt/wt | 8.8 | 0.71 | 8.2 | 0.85 | 9 | 6.9 | 9.5 | 2.346 | 2.392 |  | 0.316 | Eric Salvail | 2024-06-21
F | F318 | Pour Point - ASTM D97 | °C |  |  |  |  | 6 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Pour Point - ASTM D5949 | °C | -29 | 0.22 | -29.6 | 2.9 | 11 | -33 | -23 | 8.054 | 6.3 |  | 0.408 | Eric Salvail | 2024-06-21
F | F318 | Sulfur - ASTM D5453 | mg/kg | 9.1 | 0.55 | 8.87 | 0.43 | 12 | 7.9 | 9.5 | 1.183 | 2.979 |  | 0.401 | Eric Salvail | 2024-06-21
F | F318 | Sulfur - ASTM D7039 | mg/kg |  |  |  |  | 5 | [6 empty] | Eric Salvail | 2024-06-21
F | F318 | Sulfur, Mercaptan - ASTM D3227 | Mass % | 0 | -0.85 | 0.00011 | 0.00013 | 7 | 0 | 0.0003 | 0 | 0 |  | 0.781 | Eric Salvail | 2024-06-21
F | F318 | Viscosity at 40°C - ASTM D445 | cSt | 2.211 | 0.66 | 2.2055 | 0.0083 | 10 | 2.189 | 2.222 | 0.023 | 0.026 |  | 0.538 | Eric Salvail | 2024-06-21
F | F318 | Kinematic Viscosity at 40°C - ASTM D7042 | mm2/s | 2.211 | 0.44 | 2.2064 | 0.0103 | 7 | 2.189 | 2.214 | 0.029 | 0.013 |  | 1.155 | Eric Salvail | 2024-06-21
P | P120 | Composition - Methane - All Methods
P | P120 | Composition - Ethane - All Methods | Vol % |  |  | 0.0718 | 0.0055 | 15 | 0.064 | 0.08 | 0.015 |  |  | 0.492 | Eric Salvail | 2024-06-21
P | P120 | Composition - Propene - All Methods | Vol % |  |  | 0.0334 | 0.004 | 12 | 0.027 | 0.04 | 0.011 |  |  | 0.54 | Eric Salvail | 2024-06-21
P | P120 | Composition - Propane - All Methods | Vol % |  |  | 98.5675 | 0.0828 | 13 | 98.443 | 98.72 | 0.229 |  |  | 0.921 | Eric Salvail | 2024-06-21
P | P120 | Composition - C3 - All Methods | Vol % |  |  | 98.5907 | 0.0844 | 15 | 98.475 | 98.76 | 0.234 |  |  | 0.72 | Eric Salvail | 2024-06-21
P | P120 | Composition - Isobutane - All Methods | Vol % |  |  | 1.3368 | 0.0763 | 16 | 1.16 | 1.451 | 0.211 |  |  | 0.719 | Eric Salvail | 2024-06-21
P | P120 | Composition - n-Butane - All Methods | Vol % |  |  | 0.0019 | 0.0033 | 13 | 0 | 0.01 | 0.009 |  |  | 2.166 | Eric Salvail | 2024-06-21
P | P120 | Density, LPG, at 15°C - All Methods | kg/m³ | 507.8 | 0.13 | 507.75 | 0.39 | 23 | 506.866 | 508.5 | 1.082 |  |  | 0.867 | Eric Salvail | 2024-06-21
P | P120 | Density Method Used |  | d2598 | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Sulfur, LPG - All Methods | ppmw | 9.48 | -2.45 | 17.3 | 3.2 | 27 | 9.48 | 24.4 | 8.785 |  |  | 0.313 | Eric Salvail | 2024-06-21
P | P120 | Sulfur Method Used |  | 5504 | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Vapor Pressure LPG at 37.8°C - All Methods | kPa | 1192 | 0.43 | 1191.5 | 1.1 | 24 | 1189.95 | 1193.2 | 3.12 |  |  | 1.063 | Eric Salvail | 2024-06-21
P | P120 | VP Method Used |  | d2598 | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Butane and Heavier - All Methods | Vol. % | 1.34 | -0.05 | 1.3439 | 0.0772 | 19 | 1.16 | 1.48 | 0.214 |  |  | 0.418 | Eric Salvail | 2024-06-21
P | P120 | Composition Method (eg. CGSB 14.3) |  | d2163 | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Injection Method (eg. LSV, GSV) |  | gsv | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Column Type (eg. Al Plot or Hi Res Single Column) |  | cap | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Column Length (eg. 100 m) |  | nd | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Calibration Type |  | extstd | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Detector Type (eg. TCD, FID) |  | fid | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Composition |  | 0 | [10 empty] | Eric Salvail | 2024-06-21
P | P120 | Methane - ASTM D2163
P | P120 | Composition - Ethane - ASTM D2163 | Vol % | 0.074 | 0.41 | 0.0718 | 0.0053 | 31 | 0.06 | 0.083 | 0.015 | 0.058 | 0.064 - Pass | 0.427 | Eric Salvail | 2024-06-21
P | P120 | Composition - Propene - ASTM D2163 | Vol % | 0.035 | 0.73 | 0.0328 | 0.003 | 31 | 0.027 | 0.04 | 0.008 | 0.03 | 0.078 - Pass | 0.9 | Eric Salvail | 2024-06-21
P | P120 | Composition - Propane - ASTM D2163 | Vol % | 98.761 | 2.71 | 98.5549 | 0.0759 | 31 | 98.4 | 98.761 | 0.21 | 4.33 | 0.002 - Pass | 1.494 | Eric Salvail | 2024-06-21
P | P120 | Composition - C3 - ASTM D2163 | Vol % | 98.796 | 2.52 | 98.5894 | 0.0819 | 27 | 98.43 | 98.796 | 0.227 |  |  | 1.156 | Eric Salvail | 2024-06-21
P | P120 | Composition - Isobutane - ASTM D2163 | Vol % | 1.127 | -2.93 | 1.3363 | 0.0715 | 32 | 1.127 | 1.48 | 0.198 | 0.185 | 1.141 - Pass | 1.61 | Eric Salvail | 2024-06-21
P | P120 | Composition - n-Butane - ASTM D2163 | Vol % | 0.003 | 0.25 | 0.0022 | 0.0033 | 23 | 0 | 0.01 | 0.009 | 0.009 | 1.117 - Pass | 3.052 | Eric Salvail | 2024-06-21
D- | D-1207 | Cetane Number - ASTM D613 |  |  |  | 46.35 | 0.59 | 12 | 45.2 | 47.2 | 1.647 | 3.594 |  | 0.54 | Martine Carrier | 2024-06-11
D- | D-1207 | Setpoint Temperature - ASTM D6890
D- | D-1207 | Test 1 Charge Air Test Temperature - ASTM D6890 |  | 546.9 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 2 Charge Air Test Temperature - ASTM D6890 |  | 546.8 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test Results |  | 4.598 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 1 Ignition Delay - ASTM D6890
D- | D-1207 | Test 1 Derived Cetane Number - ASTM D6890 |  | 45.04 | -0.58 | 45.49 | 0.78 | 8 | 44.4 | 46.56 | 2.162 | 2.445 |  | 0.248 | Martine Carrier | 2024-06-11
D- | D-1207 | Test 1 Charge air pressure - ASTM D6890 |  | 309.4 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 1 Injection actuator air pressure - ASTM D6890 |  | 174.6 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 2 Ignition Delay - ASTM D6890 |  | 4.585 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 2 Derived Cetane Number - ASTM D6890 |  | 45.16 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 2 Charge air pressure - ASTM D6890 |  | 309.4 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Test 2 Injection actuator air pressure - ASTM D6890 |  | 174.7 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Heptane Check Ignition Delay - ASTM D6890 |  | 3.794 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | QC Sample |  | 17CPCFL01 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Lot Number - ASTM D6890
D- | D-1207 | ARV - ASTM D6890 |  | 39.72 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Derived Cetane Number - ASTM D6890 |  | 42.19 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | TALM Precision Package - ASTM D6890 |  | YES | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Average MCH Ignition Delay at last calibration - ASTM D6890 |  | 10.597 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Date of last calibration - ASTM D6890 |  | Jan10th,2024 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Gravity, API at 60 °F - All Methods |  | 35.08 | -6.91 | 35.19 | 0.02 | 9 | 35.16 | 35.2 | 0.043 |  |  | 0.624 | Martine Carrier | 2024-06-11
D- | D-1207 | Gravity Method Used |  | D4052 | [10 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Sulfur - All Methods | mg/kg |  |  | 8 | 0.3 | 10 | 7.64 | 8.8 | 0.955 |  |  | 0.617 | Martine Carrier | 2024-06-11
D- | D-1207 | Sulfur Method Used | [12 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Viscosity at 40°C - ASTM D445 | cSt | 3.013 | 0.4 | 3.0065 | 0.0162 | 11 | 2.986 | 3.046 | 0.045 | 0.067 |  | 0.747 | Martine Carrier | 2024-06-11
D- | D-1207 | Aniline Point - ASTM D611 | °C | 65.75 |  |  |  | 6 | [6 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Total Aromatics - All Methods | Vol. % |  |  |  |  | 4 | [6 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Total Aromatics Method Used | [12 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Aromatics - ASTM D1319 | Vol. % |  |  |  |  | 4 | [6 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Olefins - ASTM D1319 | Vol. % |  |  |  |  | 4 | [6 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Saturates - ASTM D1319 | Vol. % |  |  |  |  | 4 | [6 empty] | Martine Carrier | 2024-06-11
D- | D-1207 | Flash Point - ASTM D93 | °F | 135 | 1.66 | 130.7 | 2.6 | 10 | 127 | 135 | 7.15 |  |  | 0.291 | Martine Carrier | 2024-06-11
D- | D-1207 | Distillation, IBP Recovered - ASTM D86 (Automatic) | °F | 345 | 1.08 | 340.56 | 4.1 | 11 | 332.6 | 345 | 11.358 |  |  | 0.857 | Martine Carrier | 2024-06-11
D- | D-1207 | Distillation, 10% Recovered - ASTM D86 (Automatic) | °F | 421.1 | 0.62 | 418.85 | 3.62 | 11 | 412 | 424.02 | 10.025 |  |  | 0.254 | Martine Carrier | 2024-06-11
D- | D-1207 | DIstillation, 50% Recovered - ASTM D86 (Automatic) | °F | 527.6 | 0.46 | 526.95 | 1.4 | 11 | 523.9 | 528.68 | 3.866 |  |  | 0.482 | Martine Carrier | 2024-06-11
D- | D-1207 | Distillation, 90% Recovered - ASTM D86 (Automatic) | °F | 573 | -22.33 | 629.88 | 2.55 | 9 | 626.3 | 633 | 7.056 |  |  | 0.543 | Martine Carrier | 2024-06-11 |  | reprise demandée DA-08
D- | D-1207 | Distillation, FBP Recovered - ASTM D86 (Automatic) | °F | 682.7 | 0.91 | 678.98 | 4.07 | 11 | 669.9 | 683.5 | 11.266 |  |  | 0.493 | Martine Carrier | 2024-06-11
D- | D-1207 | Recovery - ASTM D86 | Vol. % | 99 | 2.18 | 98.25 | 0.34 | 11 | 97.9 | 99 | 0.946 |  |  | 0.68 | Martine Carrier | 2024-06-11
D- | D-1207 | Residue - ASTM D86 | Vol. % | 0.1 | -9.19 | 1.4 | 0.14 | 9 | 1.2 | 1.7 | 0.392 |  |  | 0.58 | Martine Carrier | 2024-06-11
D- | D-1207 | Loss - ASTM D86 | Vol. % | 0.8 | 0.73 | 0.53 | 0.37 | 11 | 0 | 1.4 | 1.037 |  |  | 0.599 | Martine Carrier | 2024-06-11

## Parametres
Paramètres
Cetane Number - ASTM D613
Test 1 Derived Cetane Number - ASTM D6890
Gravity, API at 60 °F - All Methods
Sulfur - All Methods
Viscosity at 40°C - ASTM D445
Aniline Point - ASTM D611
Total Aromatics - All Methods
Flash Point - ASTM D93
Distillation, IBP Recovered - ASTM D86 (Automatic)
Distillation, 10% Recovered - ASTM D86 (Automatic)
DIstillation, 50% Recovered - ASTM D86 (Automatic)
Distillation, 90% Recovered - ASTM D86 (Automatic)
Distillation, FBP Recovered - ASTM D86 (Automatic)
Recovery - ASTM D86
Residue - ASTM D86
Loss - ASTM D86
Test 1 Derived Cetane Number - ASTM D7668
Aromatics - ASTM D1319
Olefins - ASTM D1319
Saturates - ASTM D1319
Residue - ASTM D86 (Automatic)
Simulated Distillation, IBP, °C - ASTM D2887
Simulated Distillation, 5%, °C - ASTM D2887
Simulated Distillation, 10%, °C - ASTM D2887
Simulated Distillation, 20%, °C - ASTM D2887
Simulated Distillation, 30%, °C - ASTM D2887
Simulated Distillation, 40%, °C - ASTM D2887
Simulated Distillation, 50%, °C - ASTM D2887
Simulated Distillation, 60%, °C - ASTM D2887
Simulated Distillation, 70%, °C - ASTM D2887
Simulated Distillation, 80%, °C - ASTM D2887
Simulated Distillation, 90%, °C - ASTM D2887
Simulated Distillation, 95%, °C - ASTM D2887
Simulated Distillation, FBP, °C - ASTM D2887
Recovery - ASTM D86 (Automatic)
Loss - ASTM D86 (Automatic)
Cetane Index, Calculated - ASTM D976
Cetane Index, Calculated - ASTM D4737 B
Cetane Index, Calculated - ASTM D4737 A
Cloud Point - ASTM D2500
Cloud Point - ASTM D5773
Density, at 15°C - ASTM D4052
Electrical Conductivity @ 20°C, ASTM D2624
HFRR, Major Axis - ASTM D6079
HFRR, Minor Axis - ASTM D6079
HFRR, Wear Scar Diameter - ASTM D6079
Pour Point - ASTM D97
Pour Point - ASTM D5949
Sulfur - ASTM D5453
Sulfur - ASTM D7039
Water Content - ASTM D6304, Procedure A
Dynamic Viscosity at 40°C - ASTM D7042
Kinematic Viscosity at 40°C - ASTM D7042
Density at 40°C - ASTM D7042
Density at 15°C - ASTM D7042
Kinematic Viscosity at 0°C, ASTM D7042
Biodiesel Content - ASTM D7371
Acid Number - ASTM D664
Carbon Residue, Micro Conradson - ASTM D4530
Nitrogen - ASTM D4629
Sulfur, Mercaptan - ASTM D3227
Benzene - CAN/CGSB-3.0 No. 14.3
Distillation, IBP Evaporated - ASTM D86 (Automatic)
Distillation, 10% Evaporated - ASTM D86 (Automatic)
Distillation, 50% Evaporated - ASTM D86 (Automatic)
Distillation, 90% Evaporated - ASTM D86 (Automatic)
Distillation, FBP Evaporated - ASTM D86 (Automatic)
Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic)
Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic)
Olefins (corrected) - ASTM D1319
Aromatics - CAN/CGSB-3.0 No. 14.3
Olefins - CAN/CGSB-3.0 No. 14.3
Benzene Emissions Number, Summer - Regulation SOR/97-493
Oxygen Content - All Methods
Ethanol - All Methods
Driveablity Index (ethanol corrected) - CAN/CGSB-3.511
Gum, Unwashed Content - ASTM D381
Gum, Solvent Washed Content - ASTM D381
Vapor Pressure - ASTM D5191
Octane Number, Research - ASTM D2699
Octane Number, Motor - ASTM D2700
Aromatics (corrected) - ASTM D1319
Driveability Index - CAN/CGSB-3.5
Induction Period at 100°C, ASTM D525
Benzene Emissions Number, Winter - Regulation SOR/97-493
Composition - Ethane - All Methods
Composition - Ethane - ASTM D2163
Composition - Propane - All Methods
Composition - C3 - All Methods
Composition - Isobutane - All Methods
Composition - n-Butane - All Methods
Density, LPG, at 15°C - All Methods

## CORRESPONDANCE
Method | SAMPLEPTID | SAMPLEPTID | SAMPLEPTID | SAMPLEPTID | LV8 test | LV8 param
Cetane Number - ASTM D613
Test 1 Derived Cetane Number - ASTM D6890
Gravity, API at 60 °F - All Methods
Sulfur - All Methods | S-08 | S-09 |  |  | D5453 - Soufre (mg/kg) | S
Viscosity at 40°C - ASTM D445 | VIS-01 | MVV-03 |  |  | D445 - Viscosité à 40°C (cSt) | VIS40
Aniline Point - ASTM D611 | ANIL-1 |  |  |  | D611 - Point d'aniline (°C) | P.ANIL
Total Aromatics - All Methods
Flash Point - ASTM D93 | PEA-04 | PEA-05 | PEA-06 | PEA-07 | D93A - Point d'éclair (°C) | P.ECL
Distillation, IBP Recovered - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | P.I.E.
Distillation, 10% Recovered - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | 10 PCT
DIstillation, 50% Recovered - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | 50 PCT
Distillation, 90% Recovered - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | 90 PCT
Distillation, FBP Recovered - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | P.F.E.
Recovery - ASTM D86 | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | VOLREC
Residue - ASTM D86 | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | VOLRES
Loss - ASTM D86 | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique (°C) | VOLPER
Test 1 Derived Cetane Number - ASTM D7668 |  |  |  |  | D86 - Distillation automatique (°C)
Aromatics - ASTM D1319 |  |  |  |  | D86 - Distillation automatique (°C)
Olefins - ASTM D1319 |  |  |  |  | D86 - Distillation automatique (°C)
Saturates - ASTM D1319 |  |  |  |  | D86 - Distillation automatique (°C)
Simulated Distillation, IBP, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | P.I.E.
Simulated Distillation, 5%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 5 PCT
Simulated Distillation, 10%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 10 PCT
Simulated Distillation, 20%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 20 PCT
Simulated Distillation, 30%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 30 PCT
Simulated Distillation, 40%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 40 PCT
Simulated Distillation, 50%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 50 PCT
Simulated Distillation, 60%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 60 PCT
Simulated Distillation, 70%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 70 PCT
Simulated Distillation, 80%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 80 PCT
Simulated Distillation, 90%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 90 PCT
Simulated Distillation, 95%, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | 95 PCT
Simulated Distillation, FBP, °C - ASTM D2887 | SIMDIS-02 | SIMDIS-03 |  |  | D2887 - Distillation simulée (°C) | P.F.E.
Cetane Index, Calculated - ASTM D976 |  |  |  |  | D2887 - Distillation simulée (°C)
Cetane Index, Calculated - ASTM D4737 B |  |  |  |  | D2887 - Distillation simulée (°C)
Cetane Index, Calculated - ASTM D4737 A |  |  |  |  | D2887 - Distillation simulée (°C)
Cloud Point - ASTM D2500 |  |  |  |  | D2887 - Distillation simulée (°C)
Cloud Point - ASTM D5773 | PEF-01 | PEF-02 | PEF-03 |  | D5773 - Point de trouble (°C) | P.TRO
Density, at 15°C - ASTM D4052 | MV-01 | MVV-03 |  |  | D4052 - Masse volumique (kg/L) | M.VOL
Electrical Conductivity @ 20°C, ASTM D2624 | CONDM-5 | CONDM-6 |  |  | D2624 - Conductivité (pS/m) | COND.
HFRR, Major Axis - ASTM D6079
HFRR, Minor Axis - ASTM D6079
HFRR, Wear Scar Diameter - ASTM D6079 | HFRR-01 |  |  |  | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
Pour Point - ASTM D97
Pour Point - ASTM D5949 | PEF-01 | PEF-02 | PEF-03 |  | D5949 - Point d'écoulement (°C) | P.ECO
Sulfur - ASTM D5453 | S-08 | S-09 |  |  | D5453 - Soufre (mg/kg) | S
Sulfur - ASTM D7039
Water Content - ASTM D6304, Procedure A | KF-04 |  |  |  | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
Dynamic Viscosity at 40°C - ASTM D7042
Kinematic Viscosity at 40°C - ASTM D7042 | MVV-03 |  |  |  | D7042 - Viscosité à 40°C (cSt) | VIS40
Density at 40°C - ASTM D7042
Density at 15°C - ASTM D7042 | MVV-03 |  |  |  | D4052 - Masse volumique (kg/L) | M.VOL
Kinematic Viscosity at 0°C, ASTM D7042
Biodiesel Content - ASTM D7371
Acid Number - ASTM D664 | TIT-02 |  |  |  | D664 - Acidité TAN (mg KOH/g) | TAN
Carbon Residue, Micro Conradson - ASTM D4530 | MCRT-02 |  |  |  | D4530 - Carbone 10% résiduel (% poids) | C
Nitrogen - ASTM D4629 | NSX-01 |  |  |  | D4629 - Azote total (mg/kg) | AZOTE TOTAL
Sulfur, Mercaptan - ASTM D3227 | RSH-02 |  |  |  | D3227_M - Mercaptans (% masse) | Mercaptans
Benzene - CAN/CGSB-3.0 No. 14.3 | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
Distillation, IBP Evaporated - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | P.I.E.
Distillation, 10% Evaporated - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | 10 PCT
Distillation, 50% Evaporated - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | 50 PCT
Distillation, 90% Evaporated - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | 90 PCT
Distillation, FBP Evaporated - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | P.F.E.
Distillation, % Evaporated @ 93°C - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | E200
Distillation, % Evaporated @ 149°C - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | E300
Recovery - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | VOLREC
Loss - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | VOLPER
Residue - ASTM D86 (Automatic) | DA-04 | DA-05 | DA-07 | DA-08 | D86 - Distillation automatique Évap (°C) | VOLRES
Olefins (corrected) - ASTM D1319 |  |  |  |  | D86 - Distillation automatique Évap (°C)
Aromatics - CAN/CGSB-3.0 No. 14.3 | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
Olefins - CAN/CGSB-3.0 No. 14.3 | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
Benzene Emissions Number, Summer - Regulation SOR/97-493 | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.)
Oxygen Content - All Methods | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
Ethanol - All Methods | DHA-01 |  |  |  | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
Driveablity Index (ethanol corrected) - CAN/CGSB-3.511
Gum, Unwashed Content - ASTM D381
Gum, Solvent Washed Content - ASTM D381
Vapor Pressure - ASTM D5191 | TVR-04 | TVR-05 | TVR-06 | TVR-07 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
Octane Number, Research - ASTM D2699 | IOR-01 |  |  |  | D2699 - Indice d'octane recherche | IOR
Octane Number, Motor - ASTM D2700 | IOM-01 |  |  |  | D2700 - Indice d'octane moteur | IOM
Aromatics (corrected) - ASTM D1319
Driveability Index - CAN/CGSB-3.5
Induction Period at 100°C, ASTM D525
Benzene Emissions Number, Winter - Regulation SOR/97-493
Composition - Ethane - All Methods
Composition - Ethane - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2
Composition - Propane - All Methods
Composition - C3 - All Methods
Composition - Isobutane - All Methods
Composition - n-Butane - All Methods
Density, LPG, at 15°C - All Methods | RGC-04 | RGC-05 | RGC-06 |  | D2598 - Propriétés physiques GPL (calculs) | M.VOL
Sulfur, LPG - All Methods | S-07 |  |  |  | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
Vapor Pressure LPG at 37.8°C - All Methods | RGC-04 | RGC-05 | RGC-06 |  | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
Butane and Heavier - All Methods | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
Composition - Propene - All Methods | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
Composition - Propene - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
Composition - Propane - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3
Composition - C3 - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
Composition - Isobutane - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4
Composition - n-Butane - ASTM D2163 | RGC-04 | RGC-05 | RGC-06 |  | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4
Acid Number - ASTM D3242
Total Aromatics - ASTM D6379
Existent Gum Content - IP540
Flash Point - ASTM D56
Freezing Point - ASTM D5972
Micro Separometer - ASTM D3948
Water Separation of Kerosine-type Aviation Turbine, MSEP rating, ASTM D7224 | MSEP-02 |  |  |  | D7224 - Indice MSEP | MSEP
Naphthalenes - ASTM D1840
Net Heat of Combustion - ASTM D3338
Smoke Point, corrected - ASTM D1322
Sulfur - ASTM D4294 | S-11 | S-12 |  |  | D4294 - Soufre (mg/kg) | S
Pressure Drop - ASTM D3241
Viscosity at -20°C - ASTM D445
Kinematic Viscosity at -20°C - ASTM D7042
Net Heat of Combustion - ASTM D4529
Composition | NE PAS CONSIDÉRER POUR LE MOMENT
Methane | NE PAS CONSIDÉRER POUR LE MOMENT
Ethane | NE PAS CONSIDÉRER POUR LE MOMENT
Propene | NE PAS CONSIDÉRER POUR LE MOMENT
Propane | NE PAS CONSIDÉRER POUR LE MOMENT
C3 | NE PAS CONSIDÉRER POUR LE MOMENT
Isobutane | NE PAS CONSIDÉRER POUR LE MOMENT
n-Butane | NE PAS CONSIDÉRER POUR LE MOMENT
Density at 15°C | NE PAS CONSIDÉRER POUR LE MOMENT
Density Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Sulfur | NE PAS CONSIDÉRER POUR LE MOMENT
Sulfur Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Vapor Pressure at 37.8°C | NE PAS CONSIDÉRER POUR LE MOMENT
VP Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Butane and Heavier | NE PAS CONSIDÉRER POUR LE MOMENT
Composition Method (eg. CGSB 14.3) | NE PAS CONSIDÉRER POUR LE MOMENT
Injection Method (eg. LSV, GSV) | NE PAS CONSIDÉRER POUR LE MOMENT
Column Type (eg. Al Plot or Hi Res Single Column) | NE PAS CONSIDÉRER POUR LE MOMENT
Column Length (eg. 100 m) | NE PAS CONSIDÉRER POUR LE MOMENT
Calibration Type | NE PAS CONSIDÉRER POUR LE MOMENT
Detector Type (eg. TCD, FID) | NE PAS CONSIDÉRER POUR LE MOMENT
Distillation Range | NE PAS CONSIDÉRER POUR LE MOMENT
(corrected, 760 mm Hg) | NE PAS CONSIDÉRER POUR LE MOMENT
IBP | NE PAS CONSIDÉRER POUR LE MOMENT
10% Evaporated | NE PAS CONSIDÉRER POUR LE MOMENT
50% Evaporated | NE PAS CONSIDÉRER POUR LE MOMENT
90% Evaporated | NE PAS CONSIDÉRER POUR LE MOMENT
End Point | NE PAS CONSIDÉRER POUR LE MOMENT
Recovery | NE PAS CONSIDÉRER POUR LE MOMENT
Residue | NE PAS CONSIDÉRER POUR LE MOMENT
Loss | NE PAS CONSIDÉRER POUR LE MOMENT
% Evaporated @ 93°C | NE PAS CONSIDÉRER POUR LE MOMENT
% Evaporated @ 149°C | NE PAS CONSIDÉRER POUR LE MOMENT
Aromatics (corrected) | NE PAS CONSIDÉRER POUR LE MOMENT
Olefins (corrected) | NE PAS CONSIDÉRER POUR LE MOMENT
Aromatics | NE PAS CONSIDÉRER POUR LE MOMENT
Olefins | NE PAS CONSIDÉRER POUR LE MOMENT
Benzene | NE PAS CONSIDÉRER POUR LE MOMENT
Benzene Emissions Number, Summer (Report to Nearest Whole Number) | NE PAS CONSIDÉRER POUR LE MOMENT
Oxygen Content | NE PAS CONSIDÉRER POUR LE MOMENT
Ethanol | NE PAS CONSIDÉRER POUR LE MOMENT
Methanol | NE PAS CONSIDÉRER POUR LE MOMENT
Oxygenates Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Copper Strip Corrosion, 3h at 50°C | NE PAS CONSIDÉRER POUR LE MOMENT
Corrosive Properties | NE PAS CONSIDÉRER POUR LE MOMENT
Rating | NE PAS CONSIDÉRER POUR LE MOMENT
% Corrosion | NE PAS CONSIDÉRER POUR LE MOMENT
Corrosion, Silver Wool | NE PAS CONSIDÉRER POUR LE MOMENT
Density, at 15°C | NE PAS CONSIDÉRER POUR LE MOMENT
Driveability Index | NE PAS CONSIDÉRER POUR LE MOMENT
Gum, Unwashed Content | NE PAS CONSIDÉRER POUR LE MOMENT
Gum, Solvent Washed Content | NE PAS CONSIDÉRER POUR LE MOMENT
Induction Period at 100°C | NE PAS CONSIDÉRER POUR LE MOMENT
Sulfur, Mercaptan | NE PAS CONSIDÉRER POUR LE MOMENT
Vapor Pressure (DVPE) | NE PAS CONSIDÉRER POUR LE MOMENT
Container Size | NE PAS CONSIDÉRER POUR LE MOMENT
Octane Number, Research | NE PAS CONSIDÉRER POUR LE MOMENT
Octane Number, Motor | NE PAS CONSIDÉRER POUR LE MOMENT
RON Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Barometric Pressure | NE PAS CONSIDÉRER POUR LE MOMENT
IAT, (ACTUAL) | NE PAS CONSIDÉRER POUR LE MOMENT
Cylinder Height Reading (Sample) | NE PAS CONSIDÉRER POUR LE MOMENT
TSF Rating (ACTUAL) | NE PAS CONSIDÉRER POUR LE MOMENT
MON Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Barometric Pressure (at time of TSF rating) | NE PAS CONSIDÉRER POUR LE MOMENT
IAT | NE PAS CONSIDÉRER POUR LE MOMENT
MAT | NE PAS CONSIDÉRER POUR LE MOMENT
Boiling Range Distribution | NE PAS CONSIDÉRER POUR LE MOMENT
0.05 | NE PAS CONSIDÉRER POUR LE MOMENT
0.1 | NE PAS CONSIDÉRER POUR LE MOMENT
0.2 | NE PAS CONSIDÉRER POUR LE MOMENT
0.3 | NE PAS CONSIDÉRER POUR LE MOMENT
0.4 | NE PAS CONSIDÉRER POUR LE MOMENT
0.5 | NE PAS CONSIDÉRER POUR LE MOMENT
0.6 | NE PAS CONSIDÉRER POUR LE MOMENT
0.7 | NE PAS CONSIDÉRER POUR LE MOMENT
0.8 | NE PAS CONSIDÉRER POUR LE MOMENT
0.9 | NE PAS CONSIDÉRER POUR LE MOMENT
0.95 | NE PAS CONSIDÉRER POUR LE MOMENT
Final Boiling Point | NE PAS CONSIDÉRER POUR LE MOMENT
10% Recovered | NE PAS CONSIDÉRER POUR LE MOMENT
50% Recovered | NE PAS CONSIDÉRER POUR LE MOMENT
90% Recovered | NE PAS CONSIDÉRER POUR LE MOMENT
Acid Number | NE PAS CONSIDÉRER POUR LE MOMENT
Carbon Residue, 10% Bottoms, Micro | NE PAS CONSIDÉRER POUR LE MOMENT
Cloud Point | NE PAS CONSIDÉRER POUR LE MOMENT
Cloud Point (Report to one decimal) | NE PAS CONSIDÉRER POUR LE MOMENT
Flash Point (corrected, 760 mm Hg) | NE PAS CONSIDÉRER POUR LE MOMENT
Nitrogen | NE PAS CONSIDÉRER POUR LE MOMENT
Pour Point (Must be divisible by 3) | NE PAS CONSIDÉRER POUR LE MOMENT
Pour Point (Report to nearest degree) | NE PAS CONSIDÉRER POUR LE MOMENT
Viscosity at 40°C | NE PAS CONSIDÉRER POUR LE MOMENT
Dynamic Viscosity and Density by Stabinger Viscometer | NE PAS CONSIDÉRER POUR LE MOMENT
Kinematic Viscosity at 40°C | NE PAS CONSIDÉRER POUR LE MOMENT
LTFT, Minimum Pass Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test Equipment | NE PAS CONSIDÉRER POUR LE MOMENT
Cetane/Ignition Delay Meter Model | NE PAS CONSIDÉRER POUR LE MOMENT
Nozzle Used | NE PAS CONSIDÉRER POUR LE MOMENT
Handwheel Reading by Thread Sighting | NE PAS CONSIDÉRER POUR LE MOMENT
Secondary Refernce Fuel Blends | NE PAS CONSIDÉRER POUR LE MOMENT
Upper Reference Cetane No. | NE PAS CONSIDÉRER POUR LE MOMENT
Lower Reference Cetane No. | NE PAS CONSIDÉRER POUR LE MOMENT
Cetane Number-Duplicate | NE PAS CONSIDÉRER POUR LE MOMENT
Test Result | NE PAS CONSIDÉRER POUR LE MOMENT
Handwheel Setting for Sample | NE PAS CONSIDÉRER POUR LE MOMENT
Compression Pressure at Sample Handwheel Setting | NE PAS CONSIDÉRER POUR LE MOMENT
Check Fuel | NE PAS CONSIDÉRER POUR LE MOMENT
Lot Number | NE PAS CONSIDÉRER POUR LE MOMENT
Check Fuel ARV | NE PAS CONSIDÉRER POUR LE MOMENT
Check Fuel Standard Deviation | NE PAS CONSIDÉRER POUR LE MOMENT
Check Fuel Cetane Number | NE PAS CONSIDÉRER POUR LE MOMENT
Reference Fuel Blends Used | NE PAS CONSIDÉRER POUR LE MOMENT
Engine RPM | NE PAS CONSIDÉRER POUR LE MOMENT
Test Conditions | NE PAS CONSIDÉRER POUR LE MOMENT
Setpoint Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Charge Air Test Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Charge Air Test Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test Results | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Ignition Delay | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Charge air pressure | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Injection actuator air pressure | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Ignition Delay | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Derived Cetane Number | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Charge air pressure | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Injection actuator air pressure | NE PAS CONSIDÉRER POUR LE MOMENT
Heptane Check Ignition Delay | NE PAS CONSIDÉRER POUR LE MOMENT
QC Sample | NE PAS CONSIDÉRER POUR LE MOMENT
ARV | NE PAS CONSIDÉRER POUR LE MOMENT
Derived Cetane Number | NE PAS CONSIDÉRER POUR LE MOMENT
TALM Precision Package | NE PAS CONSIDÉRER POUR LE MOMENT
Average MCH Ignition Delay at last calibration | NE PAS CONSIDÉRER POUR LE MOMENT
Date of last calibration | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Wall Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Wall Temperature | NE PAS CONSIDÉRER POUR LE MOMENT
Test 1 Combustion Delay | NE PAS CONSIDÉRER POUR LE MOMENT
Test 2 Combustion Delay | NE PAS CONSIDÉRER POUR LE MOMENT
CRM Ignition Delay | NE PAS CONSIDÉRER POUR LE MOMENT
CRM Combustion Delay | NE PAS CONSIDÉRER POUR LE MOMENT
QC Sample Lot Number | NE PAS CONSIDÉRER POUR LE MOMENT
QC Sample ARV | NE PAS CONSIDÉRER POUR LE MOMENT
QC Sample Derived Cetane Number | NE PAS CONSIDÉRER POUR LE MOMENT
Average MCH Combustion Delay at last calibration | NE PAS CONSIDÉRER POUR LE MOMENT
Gravity Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Total Aromatics Method Used | NE PAS CONSIDÉRER POUR LE MOMENT
Aniline Point | NE PAS CONSIDÉRER POUR LE MOMENT
Total Aromatics | NE PAS CONSIDÉRER POUR LE MOMENT
Saturates | NE PAS CONSIDÉRER POUR LE MOMENT
Note on B7: Eric Salvail: UTILISER LES NOMS EN JAUNE TELS QUELS, POURRAINET ÊTRE CHANGÉS DANS LE FUTUR.

## Nom PTP
NOM
D-1191
D-1192
D-1193
D-1194
D-1195
D-1196
D-1197
D-1198
D-1199
D-1200
D-1201
D-1202
D328
D329
D330
D331
D332
D333
D334
D335
D336
D337
D338
F305
F306
F307
F308
F309
F310
F311
F312
F313
F314
F315
M967
M968
M969
M970
M971
M972
M973
M974
M975
M976
P118
T330
T331
T332
T333
T334
T335
T336
T337
T338
T339
T340

## Mapping appareil
ILCPDESC | SAMPLE_DATE | SAMPLEPTID | TESTID | PROPERTYID
D-1190 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1190 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1190 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1191 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1191 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1191 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1192 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1192 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1192 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1193 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1193 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1193 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1194 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1194 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1194 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1195 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1195 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1195 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1196 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1196 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1196 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1197 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1197 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1197 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1198 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1198 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1198 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1199 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1199 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1199 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1200 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1200 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1200 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1201 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1201 | 2023-04-26 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1201 | 2024-04-24 10:33 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | 10 PCT
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | 50 PCT
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | 90 PCT
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.F.E.
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.I.E.
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLPER
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLREC
D-1202 | 2023-12-08 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLRES
D-1202 | 2023-12-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1202 | 2023-12-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1202 | 2023-12-08 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1202 | 2023-12-08 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1202 | 2023-12-08 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1202 | 2023-12-08 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1202 | 2023-11-24 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
D-1202 | 2023-12-08 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
D-1202 | 2023-12-08 | 5071-PEA-05 | D93A - Point d'éclair (°F) | P.ECL
D-1202 | 2023-12-08 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 10 PCT
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 50 PCT
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 90 PCT
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.F.E.
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.I.E.
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLPER
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLREC
D-1203 | 2024-01-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLRES
D-1203 | 2024-01-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1203 | 2024-01-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1203 | 2024-01-12 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1203 | 2024-01-12 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1203 | 2024-01-12 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1203 | 2024-01-12 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1203 | 2024-01-12 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
D-1203 | 2024-01-12 | 5071-PEA-06 | D93A - Point d'éclair (°F) | P.ECL
D-1203 | 2024-01-12 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | 10 PCT
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | 50 PCT
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | 90 PCT
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.F.E.
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.I.E.
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLPER
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLREC
D-1204 | 2024-02-09 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLRES
D-1204 | 2024-02-09 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1204 | 2024-02-09 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1204 | 2024-02-09 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1204 | 2024-02-09 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1204 | 2024-02-09 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1204 | 2024-02-09 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1204 | 2024-02-09 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
D-1204 | 2024-02-09 | 5071-PEA-07 | D93A - Point d'éclair (°F) | P.ECL
D-1204 | 2024-02-09 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 10 PCT
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 50 PCT
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 90 PCT
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.F.E.
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.I.E.
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLPER
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLREC
D-1205 | 2024-03-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLRES
D-1205 | 2024-03-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1205 | 2024-03-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1205 | 2024-03-08 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1205 | 2024-03-08 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1205 | 2024-03-08 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1205 | 2024-03-08 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1205 | 2024-03-08 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
D-1205 | 2024-03-08 | 5071-PEA-08 | D93A - Point d'éclair (°F) | P.ECL
D-1205 | 2024-03-08 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | 10 PCT
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | 50 PCT
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | 90 PCT
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.F.E.
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.I.E.
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLPER
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLREC
D-1206 | 2024-04-12 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLRES
D-1206 | 2024-04-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1206 | 2024-04-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1206 | 2024-04-12 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1206 | 2024-04-12 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1206 | 2024-04-12 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1206 | 2024-04-12 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1206 | 2024-07-26 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
D-1206 | 2024-04-12 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
D-1206 | 2024-04-12 | 5071-PEA-05 | D93A - Point d'éclair (°F) | P.ECL
D-1206 | 2024-04-12 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 10 PCT
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | 10 PCT
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | 50 PCT
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 50 PCT
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | 90 PCT
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | 90 PCT
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.F.E.
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.F.E.
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.I.E.
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.I.E.
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLPER
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLPER
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLREC
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLREC
D-1207 | 2024-05-10 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLRES
D-1207 | 2024-04-12 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLRES
D-1207 | 2024-05-10 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1207 | 2024-05-10 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1207 | 2024-05-10 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1207 | 2024-05-10 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1207 | 2024-05-10 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1207 | 2024-05-10 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1207 | 2024-05-10 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
D-1207 | 2024-05-10 | 5071-PEA-06 | D93A - Point d'éclair (°F) | P.ECL
D-1207 | 2024-05-10 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | 10 PCT
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | 50 PCT
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | 90 PCT
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.F.E.
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.I.E.
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLPER
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLREC
D-1208 | 2024-06-07 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLRES
D-1208 | 2024-06-07 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1208 | 2024-06-07 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1208 | 2024-06-07 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1208 | 2024-06-07 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1208 | 2024-06-07 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1208 | 2024-06-07 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1208 | 2024-06-07 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
D-1208 | 2024-06-07 | 5071-PEA-07 | D93A - Point d'éclair (°F) | P.ECL
D-1208 | 2024-06-07 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | 10 PCT
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | 50 PCT
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | 90 PCT
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.F.E.
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.I.E.
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLPER
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLREC
D-1209 | 2024-07-12 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLRES
D-1209 | 2024-07-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1209 | 2024-07-12 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1209 | 2024-07-12 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1209 | 2024-07-12 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1209 | 2024-07-12 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1209 | 2024-07-12 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1209 | 2024-07-12 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
D-1209 | 2024-07-12 | 5071-PEA-08 | D93A - Point d'éclair (°F) | P.ECL
D-1209 | 2024-07-12 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | 10 PCT
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | 50 PCT
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | 90 PCT
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.F.E.
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | P.I.E.
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLPER
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLREC
D-1210 | 2024-08-09 | 5071-DA-07 | D86 - Distillation automatique (°F) | VOLRES
D-1210 | 2024-08-09 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1210 | 2024-08-09 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1210 | 2024-08-09 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1210 | 2024-08-09 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1210 | 2024-08-09 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1210 | 2024-08-09 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1210 | 2024-08-09 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
D-1210 | 2024-08-09 | 5071-PEA-05 | D93A - Point d'éclair (°F) | P.ECL
D-1210 | 2024-08-09 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | 10 PCT
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | 50 PCT
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | 90 PCT
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.F.E.
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | P.I.E.
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLPER
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLREC
D-1211 | 2024-08-09 | 5071-DA-08 | D86 - Distillation automatique (°F) | VOLRES
D-1211 | 2024-09-06 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1211 | 2024-09-06 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1211 | 2024-09-06 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1211 | 2024-09-06 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1211 | 2024-09-06 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1211 | 2024-09-13 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1211 | 2024-09-09 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
D-1211 | 2024-09-09 | 5071-PEA-06 | D93A - Point d'éclair (°F) | P.ECL
D-1211 | 2024-09-06 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | 10 PCT
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | 50 PCT
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | 90 PCT
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.F.E.
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | P.I.E.
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLPER
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLREC
D-1212 | 2024-10-11 | 5071-DA-04 | D86 - Distillation automatique (°F) | VOLRES
D-1212 | 2024-10-11 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1212 | 2024-10-11 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1212 | 2024-10-11 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1212 | 2024-10-11 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1212 | 2024-10-11 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1212 | 2024-10-11 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1212 | 2024-10-11 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
D-1212 | 2024-10-11 | 5071-PEA-07 | D93A - Point d'éclair (°F) | P.ECL
D-1212 | 2024-10-11 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 10 PCT
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 50 PCT
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | 90 PCT
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.F.E.
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | P.I.E.
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLPER
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLREC
D-1213 | 2024-11-08 | 5071-DA-05 | D86 - Distillation automatique (°F) | VOLRES
D-1213 | 2024-11-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)
D-1213 | 2024-11-08 | 5071-IQT-01 | D6890 - IQT Indice cétane calculé (NEG) | Cétane (DCN)-2
D-1213 | 2024-11-08 | 5071-MV-01 | D4052 - Densité à 60°F (°API) | API
D-1213 | 2024-11-08 | 5071-MVV-03 | D4052 - Densité à 60°F (°API) | API
D-1213 | 2024-11-08 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D-1213 | 2024-11-08 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D-1213 | 2024-11-08 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
D-1213 | 2024-11-08 | 5071-PEA-08 | D93A - Point d'éclair (°F) | P.ECL
D-1213 | 2024-11-08 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D325 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D325 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D325 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D325 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D325 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D325 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D325 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D325 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D325 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D326 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D326 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D326 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D326 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D326 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D326 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D326 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D326 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D326 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D327 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D327 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D327 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D327 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D327 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D327 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D327 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D327 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D327 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D328 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D328 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D328 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D328 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D328 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D328 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D328 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D328 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D328 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D329 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D329 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D329 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D329 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D329 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D329 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D329 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D329 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D329 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D330 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D330 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D330 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D330 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D330 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D330 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D330 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D330 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D330 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D331 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D331 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D331 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D331 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D331 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D331 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D331 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D331 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D331 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D332 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D332 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D332 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D332 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D332 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D332 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D332 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D332 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D332 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D333 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D333 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D333 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D333 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D333 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D333 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D333 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D333 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D333 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D334 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D334 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D334 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D334 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D334 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D334 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D334 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D334 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D334 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D335 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D335 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D335 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D335 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D335 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D335 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D335 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D335 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D335 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D336 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D336 | 2023-04-24 10:56 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D336 | 2023-04-24 10:59 | 5071-PEA-04 | D93A - Point d'éclair (°C) | P.ECL
D336 | 2024-04-24 10:48 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D336 | 2023-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D336 | 2023-04-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D336 | 2024-04-24 10:46 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D336 | 2023-04-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D336 | 2023-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D337 | 2023-11-24 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D337 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D337 | 2023-11-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D337 | 2023-11-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D337 | 2023-11-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D337 | 2023-11-24 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D337 | 2023-11-24 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D337 | 2023-11-24 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D337 | 2023-11-24 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D337 | 2023-11-24 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D337 | 2023-11-24 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D337 | 2023-11-24 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E. D-2887
D337 | 2023-11-24 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E. D-2887
D337 | 2023-11-24 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D337 | 2023-11-24 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D338 | 2024-01-19 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
D338 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
D338 | 2024-01-19 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D338 | 2024-01-19 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D338 | 2024-01-19 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D338 | 2024-01-19 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D338 | 2024-01-19 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D338 | 2024-01-19 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D338 | 2024-01-19 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D338 | 2024-01-19 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
D338 | 2024-01-19 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
D338 | 2024-01-19 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
D338 | 2024-01-19 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
D338 | 2024-01-19 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
D338 | 2024-01-19 | 5071-TIT-KF-04 | D6304 - Eau Karl-Fischer - produits pétroliers (mg/kg) | EAU KF
D338 | 2024-01-19 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D339 | 2024-03-22 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D339 | 2024-03-22 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D339 | 2024-03-22 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D339 | 2024-03-22 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D339 | 2024-03-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D339 | 2024-03-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D339 | 2024-03-22 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D339 | 2024-03-22 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D339 | 2024-03-22 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
D339 | 2024-03-22 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D339 | 2024-03-22 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D339 | 2024-03-22 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D339 | 2024-05-10 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
D339 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
D339 | 2024-03-22 | 5071-TIT-KF-04 | D6304 - Eau Karl-Fischer - produits pétroliers (mg/kg) | EAU KF
D339 | 2024-03-22 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D340 | 2024-05-24 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | 10 PCT
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | 50 PCT
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | 90 PCT
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.F.E.
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.I.E.
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLPER
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLREC
D340 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLRES
D340 | 2024-05-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D340 | 2024-05-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D340 | 2024-05-24 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D340 | 2024-05-24 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D340 | 2024-05-24 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D340 | 2024-05-24 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D340 | 2024-05-24 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D340 | 2024-05-24 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
D340 | 2024-05-24 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
D340 | 2024-05-24 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
D340 | 2024-05-24 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
D340 | 2024-05-24 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
D340 | 2024-05-24 | 5071-TIT-KF-04 | D6304 - Eau Karl-Fischer - produits pétroliers (mg/kg) | EAU KF
D340 | 2024-05-24 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D341 | 2024-07-26 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
D341 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
D341 | 2024-07-26 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D341 | 2024-07-26 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D341 | 2024-07-26 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D341 | 2024-07-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D341 | 2024-07-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D341 | 2024-07-26 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D341 | 2024-07-26 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D341 | 2024-07-26 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D341 | 2024-07-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D341 | 2024-07-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D341 | 2024-07-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D341 | 2024-07-26 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D341 | 2024-07-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D342 | 2024-09-20 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
D342 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
D342 | 2024-09-20 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D342 | 2024-09-20 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D342 | 2024-09-20 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D342 | 2024-09-20 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D342 | 2024-09-20 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D342 | 2024-09-20 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D342 | 2024-09-20 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D342 | 2024-09-20 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
D342 | 2024-09-20 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
D342 | 2024-09-20 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
D342 | 2024-09-20 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
D342 | 2024-09-20 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
D342 | 2024-09-20 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D342 | 2024-09-20 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
D343 | 2024-10-25 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | 10 PCT
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | 50 PCT
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | 90 PCT
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.F.E.
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.I.E.
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLPER
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLREC
D343 | 2024-10-25 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLRES
D343 | 2024-10-25 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MAJOR AXIS
D343 | 2024-10-25 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | MINOR AXIS
D343 | 2024-10-25 | 5071-HFRR-01 | D7688 - Onctuosité (HFRR) | ONCTUOSITE (HFRR)
D343 | 2024-10-25 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
D343 | 2024-10-25 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
D343 | 2024-10-25 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
D343 | 2024-10-25 | 5071-P.ANIL-1 | D611 - Point d'aniline (°C) | P.ANIL
D343 | 2024-10-25 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
D343 | 2024-10-25 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
D343 | 2024-10-25 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
D343 | 2024-10-25 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
D343 | 2024-10-25 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
D343 | 2024-10-25 | 5071-TIT-KF-04 | D6304 - Eau par titration Karl-Fischer (mg/kg) | EAU KF
D343 | 2024-10-25 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F303 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F303 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F303 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F303 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F303 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F304 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F304 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F304 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F304 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F304 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F305 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F305 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F305 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F305 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F305 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F306 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F306 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F306 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F306 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F306 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F307 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F307 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F307 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F307 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F307 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F308 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F308 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F308 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F308 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F308 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F309 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F309 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F309 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F309 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F309 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F310 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F310 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F310 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F310 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F310 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F311 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F311 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F311 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F311 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F311 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F312 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F312 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F312 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F312 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F312 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F313 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F313 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F313 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F313 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F313 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F314 | 2023-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F314 | 2023-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F314 | 2023-04-24 10:50 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F314 | 2024-04-24 10:40 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F314 | 2023-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | 10 PCT
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | 50 PCT
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | 90 PCT
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.F.E.
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.I.E.
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLPER
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLREC
F315 | 2023-12-22 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLRES
F315 | 2023-11-24 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
F315 | 2023-12-22 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F315 | 2023-12-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F315 | 2023-12-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F315 | 2023-12-22 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F315 | 2023-12-22 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F315 | 2023-12-22 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
F315 | 2023-12-22 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
F315 | 2023-12-22 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
F315 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F315 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F315 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F315 | 2023-12-22 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
F315 | 2023-12-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
F315 | 2023-12-22 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F315 | 2023-12-22 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
F316 | 2024-02-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
F316 | 2024-02-23 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F316 | 2024-01-19 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F316 | 2024-02-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F316 | 2024-02-23 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F316 | 2024-02-23 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F316 | 2024-02-23 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
F316 | 2024-02-23 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F316 | 2024-02-23 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
F316 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F316 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F316 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F316 | 2024-02-23 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F316 | 2024-02-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F316 | 2024-02-23 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F316 | 2024-02-16 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
F317 | 2024-03-22 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
F317 | 2024-03-22 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F317 | 2024-03-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F317 | 2024-03-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F317 | 2024-03-22 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F317 | 2024-03-22 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F317 | 2024-03-22 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
F317 | 2024-03-22 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
F317 | 2024-03-22 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
F317 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F317 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F317 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F317 | 2024-03-22 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
F317 | 2024-03-22 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
F317 | 2024-03-22 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F317 | 2024-03-22 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | 10 PCT
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | 50 PCT
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | 90 PCT
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.F.E.
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.I.E.
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLPER
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLREC
F318 | 2024-04-26 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLRES
F318 | 2024-04-26 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F318 | 2024-04-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F318 | 2024-04-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F318 | 2024-04-26 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F318 | 2024-04-26 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F318 | 2024-04-26 | 5071-PEA-07 | D93A - Point d'éclair (°C) | P.ECL
F318 | 2024-04-26 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F318 | 2024-04-26 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
F318 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F318 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F318 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F318 | 2024-04-26 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F318 | 2024-04-26 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F318 | 2024-04-26 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F318 | 2024-04-26 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | 10 PCT
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | 50 PCT
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | 90 PCT
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.F.E.
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.I.E.
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLPER
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLREC
F319 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLRES
F319 | 2024-06-21 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F319 | 2024-06-21 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F319 | 2024-06-21 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F319 | 2024-06-21 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F319 | 2024-06-21 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F319 | 2024-06-21 | 5071-PEA-08 | D93A - Point d'éclair (°C) | P.ECL
F319 | 2024-06-21 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
F319 | 2024-06-21 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
F319 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F319 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F319 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F319 | 2024-06-21 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
F319 | 2024-06-21 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
F319 | 2024-06-21 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F319 | 2024-06-21 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
F320 | 2024-08-23 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
F320 | 2024-08-23 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F320 | 2024-08-23 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F320 | 2024-08-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F320 | 2024-08-23 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F320 | 2024-08-23 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F320 | 2024-08-23 | 5071-PEA-05 | D93A - Point d'éclair (°C) | P.ECL
F320 | 2024-08-23 | 5071-PEF-01 PSA | D5773 - Point de trouble (°C) | P.TRO
F320 | 2024-08-23 | 5071-PEF-01 PSA | D5949 - Point d'écoulement (°C) | P.ECO
F320 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F320 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F320 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F320 | 2024-08-23 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 10 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 20 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 30 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 40 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 5 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 50 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 60 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 70 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 80 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 90 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | 95 PCT
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.F.E.
F320 | 2024-08-23 | 5071-SIMDIS-02 | D2887 - Distillation simulée (°C) | P.I.E.
F320 | 2024-08-23 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F320 | 2024-08-23 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
F321 | 2024-09-20 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
F321 | 2024-09-20 | 5071-MCRT-02 | D4530 - Carbone 10% résiduel (% poids) | C
F321 | 2024-09-20 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
F321 | 2024-09-20 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
F321 | 2024-09-20 | 5071-MVV-03 | D7042 - Viscosité à 40°C (cSt) | VIS40
F321 | 2024-09-20 | 5071-NSX-01 | D4629 - Azote total (mg/kg) | AZOTE TOTAL
F321 | 2024-09-20 | 5071-PEA-06 | D93A - Point d'éclair (°C) | P.ECL
F321 | 2024-09-20 | 5071-PEF-02 JFA | D5773 - Point de trouble (°C) | P.TRO
F321 | 2024-09-20 | 5071-PEF-02 JFA | D5949 - Point d'écoulement (°C) | P.ECO
F321 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
F321 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
F321 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
F321 | 2024-09-20 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 10 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 20 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 30 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 40 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 5 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 50 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 60 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 70 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 80 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 90 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | 95 PCT
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.F.E.
F321 | 2024-09-20 13:05 | 5071-SIMDIS-03 | D2887 - Distillation simulée (°C) | P.I.E.
F321 | 2024-09-20 | 5071-TIT-02 | D664 - Acidité TAN (mg KOH/g) | TAN
F321 | 2024-09-20 | 5071-VIS-01 | D445 - Viscosité à 40°C (cSt) | VIS40
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M962 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M962 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M962 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M962 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M962 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M962 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M962 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M962 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M962 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M963 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M963 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M963 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M963 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M963 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M963 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M963 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M963 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M963 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M964 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M964 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M964 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M964 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M964 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M964 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M964 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M964 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M964 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M965 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M965 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M965 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M965 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M965 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M965 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M965 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M965 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M965 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M966 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M966 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M966 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M966 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M966 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M966 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M966 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M966 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M966 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M967 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M967 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M967 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M967 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M967 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M967 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M967 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M967 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M967 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M968 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M968 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M968 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M968 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M968 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M968 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M968 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M968 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M968 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M969 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M969 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M969 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M969 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M969 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M969 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M969 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M969 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M969 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M970 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M970 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M970 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M970 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M970 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M970 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M970 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M970 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M970 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M971 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M971 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M971 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M971 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M971 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M971 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M971 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M971 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M971 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M972 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M972 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M972 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M972 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M972 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M972 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M972 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M972 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M972 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M973 | 2023-04-24 10:59 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M973 | 2023-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M973 | 2023-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M973 | 2023-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M973 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M973 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M973 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M973 | 2024-04-26 11:38 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M973 | 2023-04-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M974 | 2023-11-24 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M974 | 2023-11-24 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M974 | 2023-11-24 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M974 | 2023-11-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M974 | 2023-11-24 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M974 | 2023-11-24 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M974 | 2023-11-24 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M974 | 2023-11-24 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M974 | 2023-11-24 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M975 | 2023-12-22 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M975 | 2023-12-22 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 10EVAP
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 50EVAP
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 90EVAP
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E200
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E300
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.F.E.
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.I.E.
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLPER
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLREC
M975 | 2023-12-22 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLRES
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M975 | 2023-12-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M975 | 2023-12-22 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M975 | 2023-12-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M975 | 2023-12-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M975 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M975 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M975 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M975 | 2023-12-22 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M975 | 2023-12-22 | 5071-TVR-05 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M976 | 2024-01-19 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M976 | 2024-01-19 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 10EVAP
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 50EVAP
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 90EVAP
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E200
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E300
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.F.E.
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.I.E.
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLPER
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLREC
M976 | 2024-02-02 10:45 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLRES
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M976 | 2024-01-19 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M976 | 2024-01-19 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M976 | 2024-01-19 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M976 | 2024-01-19 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M976 | 2024-01-19 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M976 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M976 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M976 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M976 | 2024-01-19 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M977 | 2024-02-23 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M977 | 2024-02-23 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 10EVAP
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 50EVAP
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 90EVAP
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E200
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E300
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.F.E.
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.I.E.
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLPER
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLREC
M977 | 2024-02-23 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLRES
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M977 | 2024-02-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M977 | 2024-02-23 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M977 | 2024-02-23 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M977 | 2024-02-23 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M977 | 2024-02-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M977 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M977 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M977 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M977 | 2024-02-23 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M977 | 2024-02-23 | 5071-TVR-07 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M978 | 2024-03-22 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M978 | 2024-03-22 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M978 | 2024-03-22 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M978 | 2024-03-22 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M978 | 2024-03-22 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M978 | 2024-03-22 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M978 | 2024-03-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M978 | 2024-03-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M978 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M978 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M978 | 2024-03-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M978 | 2024-03-22 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M978 | 2024-03-22 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M979 | 2024-04-26 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M979 | 2024-04-26 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 10EVAP
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 50EVAP
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 90EVAP
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E200
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E300
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.F.E.
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.I.E.
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLPER
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLREC
M979 | 2024-04-26 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLRES
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M979 | 2024-04-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M979 | 2024-04-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M979 | 2024-04-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M979 | 2024-04-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M979 | 2024-04-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M979 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M979 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M979 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M979 | 2024-04-26 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M979 | 2024-04-26 | 5071-TVR-05 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M980 | 2024-05-24 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M980 | 2024-05-24 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 10EVAP
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 50EVAP
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 90EVAP
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E200
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E300
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.F.E.
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.I.E.
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLPER
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLREC
M980 | 2024-05-24 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLRES
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M980 | 2024-05-24 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M980 | 2024-05-24 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M980 | 2024-05-24 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M980 | 2024-05-24 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M980 | 2024-05-24 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M980 | 2024-05-24 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M980 | 2024-05-24 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M980 | 2024-05-24 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M980 | 2024-05-24 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M980 | 2024-05-24 | 5071-TVR-06 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M981 | 2024-06-21 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M981 | 2024-06-21 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 10EVAP
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 50EVAP
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 90EVAP
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E200
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E300
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.F.E.
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.I.E.
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLPER
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLREC
M981 | 2024-06-21 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLRES
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M981 | 2024-06-21 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M981 | 2024-06-21 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M981 | 2024-06-21 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M981 | 2024-06-21 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M981 | 2024-06-21 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M981 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M981 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M981 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M981 | 2024-06-21 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M981 | 2024-06-21 | 5071-TVR-07 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M982 | 2024-07-26 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M982 | 2024-07-26 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 10EVAP
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 50EVAP
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | 90EVAP
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E200
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | E300
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.F.E.
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | P.I.E.
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLPER
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLREC
M982 | 2024-07-26 | 5071-DA-04 | D86 - Distillation automatique Évap (°C) | VOLRES
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M982 | 2024-07-26 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M982 | 2024-07-26 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M982 | 2024-07-26 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M982 | 2024-07-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M982 | 2024-07-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M982 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M982 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M982 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M982 | 2024-07-26 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M982 | 2024-07-26 | 5071-TVR-04 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M983 | 2024-08-23 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M983 | 2024-08-23 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 10EVAP
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 50EVAP
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | 90EVAP
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E200
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | E300
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.F.E.
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | P.I.E.
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLPER
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLREC
M983 | 2024-08-23 | 5071-DA-05 | D86 - Distillation automatique Évap (°C) | VOLRES
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M983 | 2024-08-23 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M983 | 2024-08-23 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M983 | 2024-08-23 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M983 | 2024-08-23 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M983 | 2024-08-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M983 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M983 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M983 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M983 | 2024-08-23 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M983 | 2024-08-23 | 5071-TVR-05 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M984 | 2024-09-20 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M984 | 2024-09-20 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 10EVAP
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 50EVAP
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | 90EVAP
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E200
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | E300
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.F.E.
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | P.I.E.
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLPER
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLREC
M984 | 2024-09-20 | 5071-DA-07 | D86 - Distillation automatique Évap (°C) | VOLRES
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M984 | 2024-09-20 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M984 | 2024-09-20 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M984 | 2024-09-20 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M984 | 2024-09-20 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M984 | 2024-09-20 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M984 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M984 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M984 | 2024-09-20 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M984 | 2024-09-20 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
M984 | 2024-09-20 | 5071-TVR-06 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
M985 | 2024-10-25 | 5071-BST-01 | D525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M985 | 2024-10-25 | 5071-BST-02 | D7525 - Stabilité à l'oxydation (min.) | STAB. OXYD.
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 10EVAP
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 50EVAP
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | 90EVAP
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E200
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | E300
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.F.E.
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | P.I.E.
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLPER
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLREC
M985 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique Évap (°C) | VOLRES
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | AROMAT
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Autres Oxygénés
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | BZ
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | ETHANOL
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | METHANOL
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | MTBE
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | OLEFIN
M985 | 2024-10-25 | 5071-DHA-01 | CGSB 3.0-14.3 - PIONA (% vol.) | Oxygénés
M985 | 2024-10-25 | 5071-L-IOM-01 | D2700 - Indice d'octane moteur | IOM
M985 | 2024-10-25 | 5071-L-IOR-01 | D2699 - Indice d'octane recherche | IOR
M985 | 2024-10-25 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
M985 | 2024-10-25 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
M985 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
M985 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
M985 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
M985 | 2024-10-25 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
M985 | 2024-10-25 | 5071-TVR-07 | D5191 - Tension de Vapeur TVR* (kPa) | TVR
P118 | 2024-01-26 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C1
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2?
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3==
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C6+
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C4
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C5
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C4
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C5
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | T-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | T-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | Total
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | BUTAD=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C1
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2?
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3==
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C6+
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C5
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C4=
P118 | 2024-01-26 | 5071-RGC-04 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C5=
P118 | 2024-01-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P118 | 2024-01-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P118 | 2024-01-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P118 | 2024-01-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
P119 | 2024-03-22 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C1
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2?
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3==
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C6+
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C5
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | Total
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | BUTAD=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C1
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2?
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3==
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C6+
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C5
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C4=
P119 | 2024-03-22 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C5=
P119 | 2024-03-22 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P119 | 2024-03-22 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P119 | 2024-03-22 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P119 | 2024-03-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
P120 | 2024-05-24 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C1
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2?
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3==
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C6+
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C4
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C5
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C4
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C5
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | T-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | T-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | Total
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | BUTAD=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C1
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2?
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3==
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C6+
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C5
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C4=
P120 | 2024-05-24 | 5071-RGC-06 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C5=
P120 | 2024-05-24 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P120 | 2024-05-24 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P120 | 2024-05-24 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P120 | 2024-05-24 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
P121 | 2024-07-26 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C1
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C2?
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C3==
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C6+
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | C-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C4
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | I-C5
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C4
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C5
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | N-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | T-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | T-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% mol) | Total
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | BUTAD=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C1
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C2
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C2=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C2?
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C3
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C3=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C3==
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C6+
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | C-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | I-C5
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | T-C4=
P121 | 2024-07-26 | 5071-RGC-04 | D2163 - Chromatographie GPL (% vol. liquide) | T-C5=
P121 | 2024-07-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P121 | 2024-07-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P121 | 2024-07-26 | 5071-RGC-04 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P121 | 2024-07-26 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
P122 | 2024-09-20 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C1
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2?
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3==
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C6+
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C5
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | Total
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | BUTAD=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C1
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C2?
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C3==
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C6+
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | C-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | FacteurMol_VolLiq
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | I-C5
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | N-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C4=
P122 | 2024-09-20 | 5071-RGC-05 | D2163 - D2421 - Chromatographie GPL (% vol. liquide) | T-C5=
P122 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P122 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P122 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P122 | 2024-09-20 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
P123 | 2024-11-22 | 5071-RESID-01 | D2158 - Résidu Évaporation GPL (mL) | RESIDU
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | BUTAD=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C1
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C2?
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C3==
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C6+
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | C-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C4
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | I-C5
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C4
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C5
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | N-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | T-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | T-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% mol) | Total
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | 2M2C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | BUTAD=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C1
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C2
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C2=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C2?
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C3
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C3=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C3==
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C6+
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | C-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | I-C5
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | T-C4=
P123 | 2024-11-22 | 5071-RGC-06 | D2163 - Chromatographie GPL (% vol. liquide) | T-C5=
P123 | 2024-11-22 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
P123 | 2024-11-22 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
P123 | 2024-11-22 | 5071-RGC-06 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | AUTRE COMPOSE
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | CARBONYL SULFIDE+SO2
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | DIMETHYL SULFIDE
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ETHYL MERCAPTAN
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | H2S
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | ISO-PROPYL MERCAPTAN
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | METHYL MERCAPTAN
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | N-PROPYL MERCAPTAN
P123 | 2024-11-22 | 5071-S-07 | D5504 - Chromatographie soufre (mg/kg de S) | SOUFRE TOTAL
T329 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T329 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T329 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T329 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T329 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T329 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T329 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T329 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T329 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T330 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T330 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T330 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T330 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T330 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T330 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T330 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T330 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T330 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T331 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T331 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T331 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T331 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T331 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T331 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T331 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T331 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T331 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T332 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T332 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T332 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T332 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T332 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T332 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T332 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T332 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T332 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T333 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T333 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T333 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T333 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T333 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T333 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T333 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T333 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T333 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T334 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T334 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T334 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T334 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T334 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T334 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T334 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T334 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T335 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T335 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T335 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T335 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T335 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T335 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T335 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T335 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T335 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T336 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T336 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T336 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T336 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T336 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T336 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T336 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T336 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T336 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T337 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T337 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T337 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T337 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T337 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T337 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T337 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T337 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T337 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T338 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T338 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T338 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T338 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T338 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T338 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T338 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T338 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T338 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T339 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T339 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T339 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T339 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T339 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T339 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T339 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T339 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T339 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T340 | 2023-04-24 10:59 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T340 | 2023-04-24 10:54 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T340 | 2023-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T340 | 2023-04-26 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T340 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T340 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T340 | 2023-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T340 | 2023-04-26 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T340 | 2023-04-26 | 5071-S-11 | D5453 - Soufre (mg/kg) | S
T341 | 2023-12-22 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T341 | 2023-12-22 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T341 | 2023-12-22 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T341 | 2023-12-22 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T341 | 2023-12-22 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T341 | 2023-12-22 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T341 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T341 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T341 | 2023-12-22 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T341 | 2023-12-22 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
T341 | 2023-12-22 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T342 | 2024-01-19 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
T342 | 2024-01-19 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
T342 | 2024-01-19 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T342 | 2024-01-19 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T342 | 2024-01-19 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T342 | 2024-01-19 | 5071-PEF-02 JFA | D5972 - Point de congélation (°C) | P.CONG
T342 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T342 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T342 | 2024-01-19 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T342 | 2024-01-19 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
T342 | 2024-01-19 | 5071-S-12 | D4294 - Soufre (mg/kg) | S
T343 | 2024-02-23 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 10 PCT
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 50 PCT
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 90 PCT
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.F.E.
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.I.E.
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLPER
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLREC
T343 | 2024-02-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLRES
T343 | 2024-02-23 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T343 | 2024-02-23 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T343 | 2024-02-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T343 | 2024-02-23 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T343 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T343 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T343 | 2024-02-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T343 | 2024-02-23 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
T343 | 2024-02-23 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T344 | 2024-04-26 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | 10 PCT
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | 50 PCT
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | 90 PCT
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.F.E.
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.I.E.
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLPER
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLREC
T344 | 2024-04-26 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLRES
T344 | 2024-04-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T344 | 2024-04-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T344 | 2024-04-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T344 | 2024-04-26 | 5071-PEF-02 JFA | D5972 - Point de congélation (°C) | P.CONG
T344 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T344 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T344 | 2024-04-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T344 | 2024-04-26 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
T344 | 2024-04-26 | 5071-S-12 | D4294 - Soufre (mg/kg) | S
T345 | 2024-06-21 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | 10 PCT
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | 50 PCT
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | 90 PCT
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.F.E.
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | P.I.E.
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLPER
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLREC
T345 | 2024-06-21 | 5071-DA-04 | D86 - Distillation automatique (°C) | VOLRES
T345 | 2024-06-21 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T345 | 2024-06-21 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T345 | 2024-06-21 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T345 | 2024-06-21 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T345 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T345 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T345 | 2024-06-21 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T345 | 2024-06-21 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
T345 | 2024-06-21 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T346 | 2024-07-26 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | 10 PCT
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | 50 PCT
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | 90 PCT
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.F.E.
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | P.I.E.
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLPER
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLREC
T346 | 2024-07-26 | 5071-DA-05 | D86 - Distillation automatique (°C) | VOLRES
T346 | 2024-07-26 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T346 | 2024-07-26 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T346 | 2024-07-26 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T346 | 2024-07-26 | 5071-PEF-02 JFA | D5972 - Point de congélation (°C) | P.CONG
T346 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T346 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T346 | 2024-07-26 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T346 | 2024-07-26 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
T346 | 2024-07-26 | 5071-S-12 | D4294 - Soufre (mg/kg) | S
T347 | 2024-08-23 | 5071-CONDM-5 | D2624 - Conductivité (pS/m) | COND.
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 10 PCT
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 50 PCT
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | 90 PCT
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.F.E.
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | P.I.E.
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLPER
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLREC
T347 | 2024-08-23 | 5071-DA-07 | D86 - Distillation automatique (°C) | VOLRES
T347 | 2024-08-23 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T347 | 2024-08-23 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T347 | 2024-08-23 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T347 | 2024-08-23 | 5071-PEF-01 PSA | D5972 - Point de congélation (°C) | P.CONG
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | 2M2C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | BUTAD=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C1
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C2?
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3 TOTAL
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C3==
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C6+
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | C-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | I-C5
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | N-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | T-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% mol) | Total
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | 2M2C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | BUTAD=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C1
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C2
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C2=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C2?
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C3
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C3 TOTAL
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C3=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C3==
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C4 TOTAL ET +
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C6+
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | C-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | I-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | I-C5
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | N-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | N-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | T-C4=
T347 | 2024-09-20 | 5071-RGC-05 | D2163 - Chromatographie GPL (% vol. liquide) | T-C5=
T347 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL
T347 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | M.VOL Calculée
T347 | 2024-09-20 | 5071-RGC-05 | D2598 - Propriétés physiques GPL (calculs) | TVR Calculée
T347 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T347 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T347 | 2024-08-23 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T347 | 2024-08-23 | 5071-S-08 | D5453 - Soufre (mg/kg) | S
T347 | 2024-08-23 | 5071-S-11 | D4294 - Soufre (mg/kg) | S
T348 | 2024-10-25 | 5071-CONDM-6 | D2624 - Conductivité (pS/m) | COND.
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | 10 PCT
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | 50 PCT
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | 90 PCT
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.F.E.
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | P.I.E.
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLPER
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLREC
T348 | 2024-10-25 | 5071-DA-08 | D86 - Distillation automatique (°C) | VOLRES
T348 | 2024-10-25 | 5071-MSEP-02 | D7224 - Indice MSEP | MSEP
T348 | 2024-10-25 | 5071-MV-01 | D4052 - Masse volumique (kg/L) | M.VOL
T348 | 2024-10-25 | 5071-MVV-03 | D4052 - Masse volumique (kg/L) | M.VOL
T348 | 2024-10-25 | 5071-PEF-02 JFA | D5972 - Point de congélation (°C) | P.CONG
T348 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | M.VOL
T348 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Mercaptans
T348 | 2024-10-25 | 5071-RSH-02 | D3227_M - Mercaptans (% masse) | Vol.AgNO3
T348 | 2024-10-25 | 5071-S-09 | D5453 - Soufre (mg/kg) | S
T348 | 2024-10-25 | 5071-S-12 | D4294 - Soufre (mg/kg) | S